## Finding
Findings | Status | Previous Severity | Current Severity | Remarks
8.1 DesktopApp: Lack of Input Validation | Open | info | Low | special chars should be restricted

## System Test Cases
Test Case SuiteTitle | Test Case System Id | Test Case System Title | Test Case System Description | Test Case Steps | Test Case ExpectedResult | Test Case Actual Result | Test Case Status Checks (Yes/No) | Test Case Run by | Test Case BugID | Test Case Analysis Entry
System | 65990 | Unprotected BIOS | Check If BIOS is protected with strong password. | Press either F2 or F12 or F8 After pressing power button. | BIOS should be protected by a strong password. | secured with BIOS password. | Yes | Harshal Kukade
System | 65993 | Philips Support Connect | Try Connecting to Philips connect support before starting of the application. | Press F8 or access from boot screen by clicking immediately during its loading process. | PSC should not be accessible before the application loads. | Not accessible | Yes | Harshal Kukade
System | 65994 | Philips Support Connect | Try If you can access the explorer from date button browse functionality. | once you get into explorer all you need to do is browse to "This PC" | File Explorer should not be accessible. | Not accessible | Yes | Harshal Kukade
System | 65995 | Unprotected BIOS | Try For default Password If it is prompting for password like Space, Simply press enter, #Welcome124. | Try For default Password If it is prompting for password like Space, Simply press enter, #Welcome124. | BIOS should not have default/weak passwords. | Default credentials are not accepted by the application. #Welcome124 #Welcome123 IntraSight@123 Admin Supervisor Admin password Adminpassword Supervisor password Supervisorpassword [Enter] Pw#12345 | Yes | Harshal Kukade |  |  | BIOS settings not working on 20 OCT2023
System | 65996 | Check for Windows hotkey/shortcut key combinations | Check for Windows hotkey/shortcut key combinations. | For Windows XP: https://support.microsoft.com/en-in/help/301583/list-of-the-keyboard-shortcuts-that-are-available-in-windows-xp For Win 7/8/10: https://support.microsoft.com/en-us/help/12445/windows-keyboard-shortcuts Windows App Shortcuts: https://support.microsoft.com/en-us/help/13805/windows-keyboard-shortcuts-in-apps Accessibility shortcuts: https://support.microsoft.com/en-us/help/13810/windows-keyboard-shortcuts-accessibility Surface Hub: https://support.microsoft.com/en-us/help/13813/windows-10-microsoft-surface-hub-keyboard-shortcuts Miscellaneous: https://support.microsoft.com/en-us/help/13781/windows-10-mobile-keyboard-shortcuts-continuum-for-phones | All windows shortcut key combinations should be blocked. | Shortcuts not working over current system. Not Applicable since there is no initial foothold got in target system. No shortcut identified which can lead to give foothold of the Operating System. For Windows XP: https://support.microsoft.com/en-in/help/301583/list-of-the-keyboard-shortcuts-that-are-available-in-windows-xp For Win 7/8/10: https://support.microsoft.com/en-us/help/12445/windows-keyboard-shortcuts Windows App Shortcuts: https://support.microsoft.com/en-us/help/13805/windows-keyboard-shortcuts-in-apps Accessibility shortcuts: https://support.microsoft.com/en-us/help/13810/windows-keyboard-shortcuts-accessibility Surface Hub: https://support.microsoft.com/en-us/help/13813/windows-10-microsoft-surface-hub-keyboard-shortcuts Miscellaneous: https://support.microsoft.com/en-us/help/13781/windows-10-mobile-keyboard-shortcuts-continuum-for-phones | Yes | Harshal Kukade
System | 65997 | Accessing Internet Explorer at login screen. | Accessing Internet Explorer at login screen. | After accessing access bar using sticky key In the address bar type google.com (make sure you are connected to internet or attached LAN cable) this will open up internet explorer failed to load page from there press CTRL +U or CTRL +E that will lead to page save or favorite save page from there you can access explorer MY PC. | Internet Explorer should not be accessible from the login screen. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 65998 | Internal File Path Disclosure | Check if you can find out internal application running path from the login screen itself. | You can try pressing ALT+T command to see if you can access the services running along with your application and it will display you the application internal path. | Application Internal path should not be revealed from the login page. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 65999 | Nmap Scan (-p 1=0-65535 ,sV, -0,-A) | Run nmap scan. | Run nmap scan after connecting the system with LAN cable .make sure your system ip address match that with the system ip pool or subnet to complete a nmap scan. | Nmap full TCP scan should help identify listening ports, associated services, its version information and OS fingerprinting. | Found 2994 port open in Nmap scan | Yes | Harshal Kukade
System | 66000 | Gaining Administrator access with unprotected BIOS | Gaining Administrator access with unprotected BIOS. | Having a Bootable Pen drive of KaliLinux turn ON machine and press Del / F10 / F12 to enter BIOS. Since BIOS was not password protected leave it blank and press enter. Go to boot preference and change Boot order with Pendrive first. Once other OS is booted, replace cmd.exe with Windows/system32/Utilman.exe or sethc.exe. (these can be invoked from Windows even when user is not logged in, Press Shift 5 times to invoke sethc.exe or Screen Narrator to invoke Utilman.exe from bottom-left on windows logon screen | BIOS should be password protected. | BIOS is secured with password. Tester is not able to make changes to boot sequece to bypass Wondows OS authentication via rubber ducky | Yes | Harshal Kukade
System | 66001 | Accessing Powershell, taskmgr, notepad, cmd etc. | Accessing Powershell, taskmgr, notepad, cmd etc. | Once you have access to system explore go ti C:\ Drive and search for powershell or taskmgr or cmd.exe you can try if you have access to any of them. You can also access notepad of any local then use same notepad instance to open any notepad openable file. If you have access to powershell ISE then you can try running some of the command like net user, cmd.exe, shutdown -s. | Powershell, taskmgr, notepad, cmd etc. should not be accessible. | Not Applicable since there is no initial foothold got in target system. Tester is not able to open mentioned applications without login | Yes | Harshal Kukade
System | 66002 | Checking for unencrypted Drive | Checking for unencrypted Drive. | Browse to Bit locker from control panel explorer and see if bitlocker drive encryption is enabled. Also browse to different section of the local system and check if there exist any unencrypted or sensitive file. | Drives should be encrypted. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 66003 | Usb Port and read/write Permission | Plugin any usb drive and see if you are able to access usb drive via explorer. | Plugin any usb drive and see if you are able to access usb drive via explorer. Try copy/pasting files from local storage to usb and vice-versa. | USB should not be accessible. | Not Applicable since there is no initial foothold got in target system. read/Write permission is not restricted. Tester was able to copy paste the files via USB storage drive. access restricted to install applications. Able to get the Files explorer post login as a FSE user. No threat identified as this access is with authenticated user. User is able to copy paste the files to usb and viceversa. but exeutable files are restricted for execution. right click is restricted under application | Yes | Harshal Kukade
System | 66004 | Sensitive files | Check For all local path from where exactly software is installed and where is the installer. It is important to check for the all the files and directory of the installed application. Do check for configuration file or exe to retrieve sensitive information. | Check For all local path from where exactly software is installed and where is the installer. It is important to check for the all the files and directory of the installed application. Do check for configuration file or exe to retrieve sensitive information. | Sensitive files should be configured with appropriate permissions. Low privilege user should not be able to view those files. | Not Applicable since there is no initial foothold got in target system. Tester was able to get the installation path but after application login. With FSE user we are not able to copy paste the cmd.exe file. unable to create and maintain access via backdoor. | Yes | Harshal Kukade
System | 66005 | Check for application default credential | Check for application default credential. | Try For default Password while login into the application like admin/admin,root/root,eleva/eleva,eleca/system,eleva/,user/user.etc | Default credentials should not be used. | Default credentials are not accepted by the application admin/admin: invalid credential root/root:invalid credential eleva/eleva:invalid credential eleca/system:invalid credential eleva/:invalid credential user/user: invalid credential Tester is not able to guess whether username is incorrect or password is incorrect. once the incorrect username and password identified, login page goes to sleep mode for 5sec due to which fuzzing attack is not possible if got foothold and tried attacks. Based on the warning notification for incorrect login : "your username and/or password has been entered incorrectly. Please wait and try again." | Yes | Harshal Kukade
System | 66006 | Internet explorer exploit | See which version of internet explorer they are running and then try some available exploit. | See which version of internet explorer they are running and then try some available exploit. | Internet explorer version should not be disclosed. | Not Applicable since there is no initial foothold got in target system. Not able to access Internet Explorer without user login. | Yes | Harshal Kukade
System | 66007 | Check for bypass of McAfee Application Control (MAC/Solidcore) - Get ability to start applications. | Check for bypass of McAfee Application Control (MAC/Solidcore) - Get ability to start applications. | Get the ability to start applications: using .hta files: using .js files: var objShell = new ActiveXObject("WScript.shell"); objShell.Run("calc.exe"); File shortcuts: Create a shortcut to the required application(eg; Powershell). Pass arguments inside shortcut Abuse of unchecked file type - *.chm Abuse of unchecked filetype - *.gadget Abuse of unchecked filetype - *.jse (same as *.js) Using Java applets If powershell is blocked, preinstalled applications may be abused: Rundll32.exe (make function calls to achieve tasks) Script interpreters (perl,php,jsp,python etc) Debuggers Microsoft Office Macros | System should block execution of *.hta,*.chm,*.gadget,*.jse file types and Java Applets. System should not allow use of "WScript.shell". System should block execution of attacker introduced scripts and binaries. System should not allow execution of office macros. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 66008 | Check for bypass of McAfee Application Control (MAC/Solidcore) - Complete solidcore bypass | Check for bypass of McAfee Application Control (MAC/Solidcore) - Complete solidcore bypass. | Check if you have read/modify permissions to "C:\Program Files\McAfee\Solidcore". This can be raised as an issue if you have permissions. Issue command "sadmin disable" within solidcore folder. If solidcore was not configured with a password, this will instantly disable solidcore protection and you'll have to restart machine for changes to take effect. Issue command "sadmin scripts list" within solidcore folder and it will list all whitelisted scripts. Issue command "sadmin updaters list" within solidcore folder. Any application, binaries, scripts, users etc. added to updaters list are allowed to overwrite any file on the system and therefore can be used to completely bypass the write protection offered by McAfee Application Control. "sadmin aef list" will list all items in exclusion list. If there are custom entries, check if they can be abused to bypass solid core. If solidcore is configured with a password and you already managed to get admin privilege on the system, you could inject into "scsrvc.exe" process to delete the password file at "C:\Program Files\McAfee\Solidcore\passwd". Once deleted, just issue "sadmin disable" to completely disable solidcore. | Attacker should not have read/modify permissions in "C:\Program Files\McAfee\Solidcore". "sadmin" should be configured to use a password. Binaries introduced by attackers into the system should not be allowed to execute. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 66009 | Check for windows update enabled | Check if windows update or windows firewall is enable/disabled. | Check if windows update or windows firewall is enable/disabled. | Windows update and firewall should be enabled. | Not Applicable since there is no initial foothold got in target system. With Authenticated App login : Firewall is enabled (FW details visible after login) new FW network ACL policies rules can be added under inbound and outbound rules. Ex. Any to any. | Yes | Harshal Kukade
System | 66010 | Remote exploit | Remote exploit. | Try connecting LAN cable and running for Remote shell via HTA access exploit using metasploit. | Execution of HTA files should not be allowed. | Performed remote exploit using port 2994 - not exploitable different Auxiliaries, Payloads Exploits were executed over the system, remotely via LAN nw connection. Port/protocol service(veritas) is not executable. | Yes | Harshal Kukade
System | 66011 | TCP Socket using java as backdoor | Try if you can plant a java program, and make it like a backdoor via which attacker can read internal files present within the File system. Using same TCP socket channel if program is written which can give terminal access then, remotely attacker can control the device also. | Try if you can plant a java program, and make it like a backdoor via which attacker can read internal files present within the File system. Using same TCP socket channel if program is written which can give terminal access then, remotely attacker can control the device also. | System should not allow execution of attacker introduced binaries. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 66012 | Remote Account Lockout | Remote Account Lockout | Try running the attached script against the target system after connecting to network and see if remote account lock out occurs. for ($1=0; $i -le 15 ; $i++) { #Stop-Computer -ComputerName 192.168.1.183 -Credential scoe Qazwsxedc@12345 $password = Get-Content password.txt ¦ ConvertTo-SecureString $cred = New-Object System.Management.Automation.PsCredential "PMWRestrictedUser",$password Stop-Computer -ComputerName 169.254.102.219 -Credential $cred } Write-Host "Device User Locked" You can add any random text in passwords.txt | System should have a strong policy that lockout's accounts after 5 invalid attempts. | Account lockout enabled in the system | Yes | Harshal Kukade
System | 66013 | File read/write in restricted directory | File read/write in restricted directory. | Try if you can copy file from or to all directory within explorer. | File should not have read/write permissions. | Not Applicable since there is no initial foothold got in target system after login to application, it is found that copying files is restricted. Computer to computer and computer to usb. | Yes | Harshal Kukade
System | 66014 | Nessus/Nmap | Try running normal nmap/nessus (authenticated/unauthenticated) scan after connecting the LAN to system (type ipconfig to know the ip also you need to bring your attacker machine IP address to same pool as of testing machine). | Try running normal nmap/nessus (authenticated/unauthenticated) scan after connecting the LAN to system (type ipconfig to know the ip also you need to bring your attacker machine IP address to same pool as of testing machine). | Nmap/Nessus scans should help identify listening ports and services running on the system. | Performed nessus and Nmap scan not major critical, high or medium issues found if dhcp enabled, then webapp will be available on (https://192.168.1.2/) TLS1.2 with weak CBC cipher visible. | Yes | Harshal Kukade
System | 66015 | Regedit | Access regedit and try making changes to the registry and importing it. | See if you can access registry edit from powershell/cmd. Try making changes to registry. You can try import registry and see how it behave. | Regedit should not be accessible. | Not Applicable since there is no initial foothold got in target system. after login to application, it is found that we are able to open cmd>>regedit Able to make updations under Registry editor(to unlock app installation restriction tried to revert back changes via regedit/gpedit/applock).but didnt work. unable to bypass restrictions. | Yes | Harshal Kukade
System | 66016 | Guest Account Checking | Check if guest account is enabled or not and also see if you can access it. | Check if guest account is enabled or not and also see if you can access it. | Guest account should not be enabled. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 66017 | Hook and Hide Tool | Hook and hide tool is mostly used in DXR machines. | see if you access it. Mostly used shortcut ALT+SHIFT+F12. Try default password #Welcome123, #Welcome@123. Also you can check for default installed directory of hook and hide tool and see if you can launch it. | Hook and Hide Tool should not be accessible. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 66018 | Accessing cmd from help | Try accessing windows help button either from application or by pressing shortcut key SHIFT +F12 and there type cmd.exe and check if you can access cmd.exe prompt. | Try accessing windows help button either from application or by pressing shortcut key SHIFT +F12 and there type cmd.exe and check if you can access cmd.exe prompt. | cmd.exe should not be accessible. | sticky keys (pressing shift key 5times), Cmd not accessible by the tester via login screen. Once tester is logged in, tester was able to access cmd, fileexplorer, task manager. It’s a functionality provided by the application itself for business requirement. | Yes | Harshal Kukade
System | 66019 | Accessing admin running keyboard | Accessing admin running keyboard. | See if you can access keyboard shortcut ALT+SHIFT+T if that open keyboard panel then you can search for help and then run cmd with admin privilege. | Attacker should not be able to access the admin running keyboard. | Tester was unable to access the admin running keyboard. What all applications executed to open under windows OS , were hidden behind the thick client. Running services were visible under task manager. Not able to see any files properties. (rightclick>>properties) | Yes | Harshal Kukade
System | 66020 | DLL Injection | Test to see if an application can inject into the running process, and make changes to process address space. | Test to see if an application can inject into the running process, and make changes to process address space. Reference: https://blogs.msdn.microsoft.com/cobold/2010/10/22/dll-injection/ | The application should not be able to inject into the running process, and make changes to process address space. | Not Applicable since there is no initial foothold got in target system | Yes | Harshal Kukade
System | 66021 | Integrity Check | Check for application binaries like .exe, .dll etc are they signed with valid developer certificate. Also check if they are verifying the CRC of such files, to identify tampering of file | Check for application binaries like .exe, .dll etc are they signed with valid developer certificate. Also check if they are verifying the CRC of such files, to identify tampering of file. | Application binaries like .exe, .dll etc should be signed with valid developer certificate and CRC of such files has to be verified to identify tampering. | Not Applicable since there is no initial foothold got in target system Application is signed or unsigned not tested. thick client application testing application not tested as application installation was restricted. And without getting foothold, its not feasible to test thick client. | Yes | Harshal Kukade
System | 66022 | PXE Network Boot | Even when the system BIOS is protected with a password, the system can use network boot without the need for a password if &quot;PXE Network Boot&quot; is enabled. | Even when the system BIOS is protected with a password, the system can use network boot without the need for a password if "PXE Network Boot" is enabled. | "PXE Network Boot" should not be enabled | Not Applicable since there is no initial foothold got in target system BIOS was password protected and was not able to check the PXE boot (BIOS>>security>>NetworkBoot) Boot via LAN was not successfull. | Yes | Harshal Kukade |  |  | https://www.youtube.com/watch?v=tmRZt1q7fO4&ab_channel=ITAssist https://www.youtube.com/watch?v=sb3yU2L5U7Y&ab_channel=NETVN82
System | 88321 | Abuse of Token Privileges | Checking for Abuse of Token Privileges | After getting initial access(basic shell), try the following command to list token privileges: | The process/thread should not have any of the following privileges: | Not Applicable since there is no initial foothold got in target system. | Yes | Harshal Kukade
|  |  |  | whoami /priv | 1. SeImpersonatePrivilege
2. SeBackupPrivilege
|  |  |  | Check if the process/thread has any of the following privileges: | 3. SeRestorePrivilege
4. SeLoadDriverPrivilege
|  |  |  | 1. SeImpersonatePrivilege | 5. SeTakeOwnershipPrivilege
|  |  |  | 2. SeBackupPrivilege | 6. SeDebugPrivilege
|  |  |  | 3. SeRestorePrivilege | 7. SeAssignPrimaryPrivilege
|  |  |  | 4. SeLoadDriverPrivilege | 8. SeTcbPrivilege
5. SeTakeOwnershipPrivilege
6. SeDebugPrivilege
7. SeAssignPrimaryPrivilege
8. SeTcbPrivilege
Refer the following document for possible abuses:-
https://share-intra.philips.com/:w:/s/STS20150410114500/ESMRBooa8MtDoswm3RlJVe0BGGQtbhnYHFjHJXpDS_k46A?e=Wwohog
|  | input validation in thick client application |  | enter malicious javascript payload or input values with special characters ex. <script>alert(document.cookie)</script> <script>alert(1)</script> asd1234asdf!@#$ |  | Tester is able to invalid input entries. | no (failed) | Harshal Kukade

## NMAP
SN | Port | Protocol service | Status | Remarks
1 | 2994 | Veritas | open | Identified Port is not identified as a risky port. Tester was not able to enter into the system remotely. No exploitable vulnerabilities identified with respect to veritas protocol service with is used for the backup of the patients scanned images.
2 | 49669 | ssl | open | There is no specific exploit available for this port. MicrosoftRPC. Port 49669 is not associated with a well-known or standardized service. Ports in the range 49152 to 65535 are typically used for dynamic or private purposes. These ports are often chosen by applications or services for temporary or specific uses, and their exact purpose may vary. To determine the exact purpose of this port in your specific case, we may need to investigate the configuration of the device or software using it.

## Nessus
Severity | Vulnerabilities | CVSS v3.0 Base Score | CVE ID | resolution
Critical | 179487 - KB5029247: Windows 10 version 1809 / Windows Server 2019 Security Update (August 2023) | 9.8 (CVSS:3.0/AV:N/AC:L/PR:N/UI:N/S:U/C:H/I:H/A:H) | CVE CVE-2023-29348 CVE CVE-2023-35349 CVE CVE-2023-36431 CVE CVE-2023-36434 CVE CVE-2023-36436 CVE CVE-2023-36438 CVE CVE-2023-36557 CVE CVE-2023-36563 CVE CVE-2023-36564 CVE CVE-2023-36567 CVE CVE-2023-36570 CVE CVE-2023-36571 CVE CVE-2023-36572 CVE CVE-2023-36573 CVE CVE-2023-36574 CVE CVE-2023-36575 CVE CVE-2023-36576 CVE CVE-2023-36577 CVE CVE-2023-36578 CVE CVE-2023-36579 CVE CVE-2023-36581 CVE CVE-2023-36582 CVE CVE-2023-36583 CVE CVE-2023-36584 CVE CVE-2023-36585 CVE CVE-2023-36589 CVE CVE-2023-36590 CVE CVE-2023-36591 CVE CVE-2023-36592 CVE CVE-2023-36593 CVE CVE-2023-36594 CVE CVE-2023-36596 CVE CVE-2023-36598 CVE CVE-2023-36602 CVE CVE-2023-36603 CVE CVE-2023-36605 CVE CVE-2023-36606 CVE CVE-2023-36697 CVE CVE-2023-36698 CVE CVE-2023-36701 CVE CVE-2023-36702 CVE CVE-2023-36703 CVE CVE-2023-36704 CVE CVE-2023-36706 CVE CVE-2023-36707 CVE CVE-2023-36709 CVE CVE-2023-36710 CVE CVE-2023-36711 CVE CVE-2023-36712 CVE CVE-2023-36713 CVE CVE-2023-36717 CVE CVE-2023-36718 CVE CVE-2023-36720 CVE CVE-2023-36721 CVE CVE-2023-36722 CVE CVE-2023-36723 CVE CVE-2023-36724 CVE CVE-2023-36725 CVE CVE-2023-36726 CVE CVE-2023-36729 CVE CVE-2023-36731 CVE CVE-2023-36732 CVE CVE-2023-36743 CVE CVE-2023-36776 CVE CVE-2023-36902 CVE CVE-2023-38159 CVE CVE-2023-38166 CVE CVE-2023-41765 CVE CVE-2023-41766 CVE CVE-2023-41767 CVE CVE-2023-41768 CVE CVE-2023-41769 CVE CVE-2023-41770 CVE CVE-2023-41771 CVE CVE-2023-41772 CVE CVE-2023-41773 CVE CVE-2023-41774 CVE CVE-2023-44487 MSKB 5031361 XREF MSFT:MS23-5031361 XREF CISA-KNOWN-EXPLOITED:2023/10/31 | Apply Security Update 5031361 The remote Windows host is missing security update 5031361. It is, therefore, affected by multiple vulnerabilities https://support.microsoft.com/help/5031361
Critical | 182865 - KB5031361: Windows 10 version 1809 / Windows Server 2019 Security Update (October 2023) | 9.8 (CVSS:3.0/AV:N/AC:L/PR:N/UI:N/S:U/C:H/I:H/A:H) | CVE CVE-2023-29348 CVE CVE-2023-35349 CVE CVE-2023-36431 CVE CVE-2023-36434 CVE CVE-2023-36436 CVE CVE-2023-36438 CVE CVE-2023-36557 CVE CVE-2023-36563 CVE CVE-2023-36564 CVE CVE-2023-36567 CVE CVE-2023-36570 CVE CVE-2023-36571 CVE CVE-2023-36572 CVE CVE-2023-36573 CVE CVE-2023-36574 CVE CVE-2023-36575 CVE CVE-2023-36576 CVE CVE-2023-36577 CVE CVE-2023-36578 CVE CVE-2023-36579 CVE CVE-2023-36581 CVE CVE-2023-36582 CVE CVE-2023-36583 CVE CVE-2023-36584 CVE CVE-2023-36585 CVE CVE-2023-36589 CVE CVE-2023-36590 CVE CVE-2023-36591 CVE CVE-2023-36592 CVE CVE-2023-36593 CVE CVE-2023-36594 CVE CVE-2023-36596 CVE CVE-2023-36598 CVE CVE-2023-36602 CVE CVE-2023-36603 CVE CVE-2023-36605 CVE CVE-2023-36606 CVE CVE-2023-36697 CVE CVE-2023-36698 CVE CVE-2023-36701 CVE CVE-2023-36702 CVE CVE-2023-36703 CVE CVE-2023-36704 CVE CVE-2023-36706 CVE CVE-2023-36707 CVE CVE-2023-36709 CVE CVE-2023-36710 CVE CVE-2023-36711 CVE CVE-2023-36712 CVE CVE-2023-36713 CVE CVE-2023-36717 CVE CVE-2023-36718 CVE CVE-2023-36720 CVE CVE-2023-36721 CVE CVE-2023-36722 CVE CVE-2023-36723 CVE CVE-2023-36724 CVE CVE-2023-36725 CVE CVE-2023-36726 CVE CVE-2023-36729 CVE CVE-2023-36731 CVE CVE-2023-36732 CVE CVE-2023-36743 CVE CVE-2023-36776 CVE CVE-2023-36902 CVE CVE-2023-38159 CVE CVE-2023-38166 CVE CVE-2023-41765 CVE CVE-2023-41766 CVE CVE-2023-41767 CVE CVE-2023-41768 CVE CVE-2023-41769 CVE CVE-2023-41770 CVE CVE-2023-41771 CVE CVE-2023-41772 CVE CVE-2023-41773 CVE CVE-2023-41774 CVE CVE-2023-44487 MSKB 5031361 XREF MSFT:MS23-5031361 XREF CISA-KNOWN-EXPLOITED:2023/10/31 | Apply Security Update 5031361 The remote Windows host is missing security update 5031361. It is, therefore, affected by multiple vulnerabilities https://support.microsoft.com/help/5031361
High | 65057 - Insecure Windows Service Permissions | 8.4 (CVSS:3.0/AV:L/AC:L/PR:N/UI:N/S:U/C:H/I:H/A:H) | NA | Ensure the groups listed in nessus report do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. Path : d:\philipsserviceagent\istbroker\bin\ibsservice.exe
High | 181303 - KB5030214: Windows 10 version 1809 / Windows Server 2019 Security Update (September 2023) | 7.5 (CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P) | BID 8026 CVE CVE-1999-0519 CVE CVE-1999-0520 | To restrict access under Windows, open Explorer, right click on each share, go to the 'Sharing' tab, and click on 'Permissions'.
High | 42411 - Microsoft Windows SMB Shares Unprivileged Access | 7.5 (CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P) | BID 8026 CVE CVE-1999-0519 CVE CVE-1999-0520 | To restrict access under Windows, open Explorer, right click on each share, go to the 'Sharing' tab, and click on 'Permissions'.
High | 42873 - SSL Medium Strength Cipher Suites Supported (SWEET32) | 7.5 (CVSS:3.0/AV:N/AC:L/PR:N/UI:N/S:U/C:H/I:N/A:N) | CVE CVE-2016-2183 | The Microsoft .NET Framework installation on the remote host is missing a security update.
High | 179664 - Security Updates for Microsoft .NET Framework (August 2023) | 8.8 (CVSS:3.0/AV:N/AC:L/PR:L/UI:N/S:U/C:H/I:H/A:H) | CVE CVE-2023-36873 CVE CVE-2023-36899 MSKB 5028946 MSKB 5028947 MSKB 5028948 MSKB 5028950 MSKB 5028951 MSKB 5028952 MSKB 5028953 MSKB 5028954 MSKB 5028955 MSKB 5028956 MSKB 5028957 MSKB 5028958 MSKB 5028960 MSKB 5028961 MSKB 5028962 MSKB 5028963 MSKB 5028967 MSKB 5028968 MSKB 5028969 MSKB 5028970 MSKB 5028973 MSKB 5028974 MSKB 5028975 MSKB 5028976 MSKB 5028977 MSKB 5028978 MSKB 5028979 MSKB 5028980 MSKB 5028981 MSKB 5028982 XREF MSFT:MS23-5028946 XREF MSFT:MS23-5028947 XREF MSFT:MS23-5028948 XREF MSFT:MS23-5028950 XREF MSFT:MS23-5028951 XREF MSFT:MS23-5028952 XREF MSFT:MS23-5028953 XREF MSFT:MS23-5028954 XREF MSFT:MS23-5028955 XREF MSFT:MS23-5028956 XREF MSFT:MS23-5028957 XREF MSFT:MS23-5028958 XREF MSFT:MS23-5028960 XREF MSFT:MS23-5028961 XREF MSFT:MS23-5028962 XREF MSFT:MS23-5028963 XREF MSFT:MS23-5028967 XREF MSFT:MS23-5028968 XREF MSFT:MS23-5028969 XREF MSFT:MS23-5028970 XREF MSFT:MS23-5028973 XREF MSFT:MS23-5028974 XREF MSFT:MS23-5028975 XREF MSFT:MS23-5028976 XREF MSFT:MS23-5028977 XREF MSFT:MS23-5028978 XREF MSFT:MS23-5028979 XREF MSFT:MS23-5028980 XREF MSFT:MS23-5028981 XREF MSFT:MS23-5028982 XREF IAVA:2023-A-0406-S | The Microsoft .NET Framework installation on the remote host is missing a security update.
High | 181375 - Security Updates for Microsoft .NET Framework (September 2023) | 7.8 (CVSS:3.0/AV:L/AC:L/PR:N/UI:R/S:U/C:H/I:H/A:H) | CVE CVE-2023-36788 CVE CVE-2023-36792 CVE CVE-2023-36793 CVE CVE-2023-36794 CVE CVE-2023-36796 MSKB 5029915 MSKB 5029916 MSKB 5029917 MSKB 5029919 MSKB 5029920 MSKB 5029921 MSKB 5029922 MSKB 5029923 MSKB 5029924 MSKB 5029925 MSKB 5029926 MSKB 5029927 MSKB 5029928 MSKB 5029929 MSKB 5029931 MSKB 5029932 MSKB 5029933 MSKB 5029937 MSKB 5029938 MSKB 5029940 MSKB 5029941 MSKB 5029942 MSKB 5029943 MSKB 5029944 MSKB 5029945 MSKB 5029946 MSKB 5029947 MSKB 5029948 MSKB 5030030 MSKB 5030160 XREF MSFT:MS23-5029916 XREF MSFT:MS23-5029917 XREF MSFT:MS23-5029919 XREF MSFT:MS23-5029920 XREF MSFT:MS23-5029921 XREF MSFT:MS23-5029922 XREF MSFT:MS23-5029923 XREF MSFT:MS23-5029924 XREF MSFT:MS23-5029925 XREF MSFT:MS23-5029926 XREF MSFT:MS23-5029927 XREF MSFT:MS23-5029928 XREF MSFT:MS23-5029929 XREF MSFT:MS23-5029931 XREF MSFT:MS23-5029932 XREF MSFT:MS23-5029933 XREF MSFT:MS23-5029937 XREF MSFT:MS23-5029938 XREF MSFT:MS23-5029940 XREF MSFT:MS23-5029941 XREF MSFT:MS23-5029942 XREF MSFT:MS23-5029943 XREF MSFT:MS23-5029944 XREF MSFT:MS23-5029945 XREF MSFT:MS23-5029946 XREF MSFT:MS23-5029947 XREF MSFT:MS23-5029948 XREF MSFT:MS23-5030030 XREF MSFT:MS23-5030160 XREF IAVA:2023-A-0470 | Microsoft has released security updates for Microsoft .NET Framework.
High | 166555 - WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | 7.4 (CVSS:3.0/AV:N/AC:L/PR:N/UI:R/S:C/C:N/I:H/A:N) | CVE CVE-2013-3900 XREF CISA-KNOWN-EXPLOITED:2022/07/10 XREF IAVA:2013-A-0227 | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck

## metatsploit
Open Ports | Protocol Service | Exploit Name | Test Case Status | Severity
2994 | veritas | exploit/multi/misc/veritas_netbackup_cmdexec | passed | info
|  | exploit/multi/veritas/beagent_sha_auth_rce | passed | info
|  | exploit/windows/backupexec/remote_agent | passed | info
49669 | MicrosoftRPC | exploit/windows/dcerpc/ms03_026_dcom | passed | info
|  | exploit/windows/dcerpc/ms07_029_msdns_zonename | passed | info
|  | exploit/windows/smb/ms08_067_netapi | passed | info
|  | exploit/windows/smb/ms10_061_spoolss | passed | info

## Thick
Binary/DesktopApp | Signature Verification (signed/ unsigned) | Authenticode (TRUE/ FALSE) | SafeSEH | CFG (TRUE/ FALSE) | Remarks
No System Foothold