## NIST
Product Security Standard Assessment - NIST SP 800-53 Controls
INSTRUCTIONS FOR USE:
1. Based upon identified User Needs and System Architecture, determine which, if any, of the 19 Security Capabilities are applicalbe to this product design. (Ref IEC/TR 80001-2-2)
A. If it is determined at a capability is not in scope for this product, provide a rationale of why it is not in scope.
2. Determine the System Security level for the product under development and provide rationale for this classification. (Ref FIPS 199 Table 1 if utilzing NIST SP 800-53 ) (Note: Currently Medical Devices are considered High System Security Level)
3. For those Security Capabilities identified as in scope in step 1, and the System Security levels determined in step 2, evaluate and identify which Security Controls will be utilized in the design. Include additional controls/requirements based upon the determined System Security Level. (Ref NIST SP 800-53 Table D-2,)
A. If a particular Security code that is specified per step 3 will not be utilized, provide a rationale.
4. If PHI and/or PII will be part of the product system design, then additional Security Controls need to be evaluated and added if applicable (Ref ID 200 & 210)
|  | Pick your target Security Level for System: | [8 empty] | Low
|  |  | SL | SSL | [6 empty] | Moderate
|  | Confidentiality | High | High | [6 empty] | High
|  | Integrity | Moderate
|  | Availability | Low
No
| IEC 80001-2-2 |  |  | IEC 80001-2-8
|  |  |  | System Security level: SP 800-53 Low Mod High | [5 empty] | System Security level Rationale
ID | Security Capability (Applicable = Design Inputs) | In scope of the product? | Rationale (if not utilizing capability) | Security Controls | System Security Level Baseline Low Moderate High |  |  | Additional Control/Requirement Enhancements based upon System Security Level | Rationale for not utilizing control if identified per Security Level | Traceability reference (SRS,SDD, etc)
10 | ALOF - System shall provide automatic logoff after a period of inactivity | Yes |  | AC-1 Access control policy and management | Yes | Yes | Yes
|  |  |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-7 Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-11 Session lock | No | Yes | Yes
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AC-24 Access control decisions | No | No | No
|  |  |  | CM-4 Security impact analysis | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IA-4 Identifier management | Yes | Yes | Yes
|  |  |  | IA-11 Re-authentication | No | No | No
20 | AUDT - System shall provide audit controls documenting who is doing what with HEALTH DATA | Yes |  | AC-21 Information sharing | No | No | No
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AU-1 Audit and accountability policy and procedures | Yes | Yes | Yes
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | AU-3 Content of audit records | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | AU-4 Audit storage capacity | Yes | Yes | Yes
|  |  |  | AU-5 Response to audit processing failures | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | AU-6 Audit review, analysis and reporting | Yes | Yes | Yes
|  |  |  | AU-7 Audit reduction and report generation | No | Yes | Yes
|  |  |  | AU-8 Time stamps | Yes | Yes | Yes
|  |  |  | AU-9 Protection of audit information | Yes | Yes | Yes
|  |  |  | AU-10 Non-repudiation | No | No | Yes
|  |  |  | AU-11 Audit record retention | Yes | Yes | Yes
|  |  |  | AU-12 Audit generation | Yes | Yes | Yes
|  |  |  | AU-13 Monitoring for information disclosure | No | Yes | Yes
|  |  |  | AU-14 Session audit | No | Yes | Yes
|  |  |  | AU-15 Alternate audit capacity | No | No | No
|  |  |  | AU-16 Cross-organizational auditing | No | No | No
30 | AUTH - The system shall provide role based controlled access to HEALTH DATA and functions - access to be provided only as necessary to perform the tasks required consistent with intended use | Yes |  | AC-1: Access Control Policy and Management | Yes | Yes | Yes
|  |  |  | AC-2: Account Management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-3: Access enforcement | Yes | Yes | Yes
|  |  |  | AC-5: Separation of duties | No | Yes | Yes
|  |  |  | AC-6: Least Privilege | No | Yes + (1) (2) | Yes + (1) (2) | Mod (1) (2), High (1) (2)
|  |  |  | AC-7: Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-17: Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | AC-18 Wireless access | Yes | Yes + (1) | Yes + (1) (2) (4) (5) | Mod (1) , High (1) (2) (4) (5)
|  |  |  | AC-19 Access control for mobile devices | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | AC-21 Information sharing | No | No | No
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AC-24 Access control decisions | No | No | No
|  |  |  | PL-4 Rules of behavior | Yes | Yes | Yes
40 | CNFS - System shall provide the Hospital system administrator the ability to configure the products security capabilities | Yes |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-5 Separation of duties | No | Yes | Yes
|  |  |  | AC-6 Least privilege | No | Yes + (1) (2) | Yes + (1) (2) | Mod (1) (2), High (1) (2)
|  |  |  | CM-1 Configuration management policy and procedures | Yes | Yes | Yes
|  |  |  | CM-2 Baseline configuration | Yes | Yes + (1) (3) (4) | Yes + (1) (2) (3) (5) (6) | Mod (1) (3) (4), High (1) (2) (3) (5) (6)
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-4 Security impact analysis | Yes | Yes | Yes + (1) | High (1)
|  |  |  | CM-5 Access restrictions for change | Yes | Yes | Yes + (1) (2) (3) | High (1) (2) (3)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | CM-7 Least functionality | Yes | Yes | Yes
|  |  |  | CM-9 Configuration management plan | No | Yes | Yes
|  |  |  | SA-10 Developer configuration management | No | Yes | Yes
50 | CSUP - The system shall provide the ability of on-site service staff, remote service staff, or authorized customer staff to install/upgrade device's security patches. | Yes |  | AC-17 Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | CM-2 Baseline configuration | Yes | Yes + (1) (3) (4) | Yes + (1) (2) (3) (5) (6) | Mod (1) (3) (4), High (1) (2) (3) (5) (6)
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-4 Security impact analysis | Yes | Yes | Yes + (1) | High (1)
|  |  |  | CM-5 Access restrictions for change | Yes | Yes | Yes + (1) (2) (3) | High (1) (2) (3)
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
|  |  |  | IA-9 Service identification and authentication | No | No | No
|  |  |  | MA-1 System maintenance policy and procedures | Yes | Yes | Yes
|  |  |  | MA-2 Controlled maintenance | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | MA-3 Maintenance tools | No | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MA-4 Nonlocal maintenance | Yes | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MA-5 Maintenance personnel | Yes | Yes | Yes
|  |  |  | MA-6 Timely maintenance | No | Yes | Yes
|  |  |  | MP-1 Media protection policy and procedures | Yes | Yes | Yes
|  |  |  | SA-8 Security engineering principles | No | Yes | Yes
|  |  |  | SA-11 Developer security testing and evaluation | No | Yes | Yes
|  |  |  | SA-14 Criticality analysis | No | No | No
|  |  |  | SI-11 Error handling | No | Yes | Yes
60 | DIDT - The system shall provide the ability to directly remove information that allows identification of a person. | Yes |  | AC-8 System use notification | Yes | Yes | Yes
|  |  |  | AC-21 Information sharing | No | No | No
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AR-7 Privacy-enhanced system design and development | No | No | No
|  |  |  | AT-1 Security assurance and training policy and protection | Yes | Yes | Yes
|  |  |  | AU-3 Content of audit records | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | AU-9 Protection of audit information | Yes | Yes | Yes
|  |  |  | AU-11 Audit record retention | Yes | Yes | Yes
|  |  |  | DM-1 Minimization of personally identifiable information | No | No | No
|  |  |  | DM-2 Data retention and disposal | No | No | No
70 | DTBK - The system shall have an integral data backup capability to recover after damage or destruction of device data, hardware or software. | Yes |  | AU-9 Protection of audit information | Yes | Yes | Yes
|  |  |  | CM-1 Configuration management policy and procedure | Yes | Yes | Yes
|  |  |  | CM-2 Baseline configuration | Yes | Yes + (1) (3) (4) | Yes + (1) (2) (3) (5) (6) | Mod (1) (3) (4), High (1) (2) (3) (5) (6)
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-5 Access restrictions for changes | Yes | Yes | Yes + (1) (2) (3) | High (1) (2) (3)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | CP-1 Contingency planning policy and procedures | Yes | Yes | Yes
|  |  |  | CP-2 Contingency plan | Yes | Yes + (1) | Yes + (1) (2) (3) | Mod (1), High (1) (2) (3)
|  |  |  | CP-3 Contingency training | Yes | Yes | Yes + (1) | High (1)
|  |  |  | CP-4 Contingency plan testing | Yes | Yes + (1) | Yes + (1) (2) (4) | Mod (1), High (1) (2) (4)
|  |  |  | CP-6 Alternate storage site | No | Yes + (1) (3) | Yes + (1) (2) (3) | Mod (1) (3), High (1) (2) (3)
|  |  |  | CP-7 Alternate processing site | No | Yes + (1) (2) (3) (5) | Yes + (1) (2) (3) (4) (5) | Mod (1) (2) (3) (5), High (1) (2) (3) (4) (5)
|  |  |  | CP-8 Telecommunications services | No | Yes + (1) (2) | Yes + (1) (2) (3) (4) | Mod (1) (2), High (1) (2) (3) (4)
|  |  |  | CP-9 Information system backup | Yes | Yes + (1) | Yes + (1) (2) (3) | Mod (1), High (1) (2) (3)
|  |  |  | CP-10 Information system recovery and reconstitution | Yes | Yes + (2) (3) | Yes + (2) (3) (4) | Mod (2) (3), High (2) (3) (4)
|  |  |  | CP-13 Alternative security mechanisms | No | No | No
|  |  |  | IR-1 Incident response policy and procedures | Yes | Yes | Yes
|  |  |  | IR-2 Incident response training | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | IR-3 Incident response testing | No | Yes | Yes + (1) | High (1)
|  |  |  | IR-4 Incident handling | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-5 Incident monitoring | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IR-6 Incident reporting | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-7 Incident response assistance | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-8 Incident response plan | Yes | Yes | Yes
|  |  |  | IR-9 Information spillage response | No | No | No
|  |  |  | IR-10 Integrated information security analysis team | No | No | No
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | PM-9 RISK MANAGEMENT strategy | Yes | Yes | Yes
80 | EMRG - The system shall provide users the ability to access private data in case of an emergency that requires immediate access to private data. | Yes |  | AC-1 Access control policy and management | Yes | Yes | Yes
|  |  |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-14 Permitted actions without identification or authentication | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
|  |  |  | RA-5 Vulnerability scanning | Yes | Yes + (1) | Yes + (1) (2) (3) (4) (5) (7) | Mod (1), High (1) (2) (3) (4) (5) (7)
90 | IGAU - The device shall ensure the integrity of stored data with implicit/explicit error detection/correction technology | Yes |  | SA-13 Trustworthiness | No | No | Yes
|  |  |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SC-13 Cryptographic protection | Yes | Yes | Yes
|  |  |  | SC-17 Public key infrastructure certificates | No | Yes | Yes
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | SI-7 Software and information integrity | No | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | SI-10 Information input validation | No | Yes | Yes
100 | MLDP - The system shall support the use of anti-malware mechanism | Yes |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | IR-1 Incident response policy and procedures | Yes | Yes | Yes
|  |  |  | IR-2 Incident response training | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | IR-3 Incident response testing | No | Yes | Yes + (1) | High (1)
|  |  |  | IR-4 Incident handling | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-5 Incident monitoring | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IR-6 Incident reporting | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-7 Incident response assistance | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-8 Incident response plan | Yes | Yes | Yes
|  |  |  | MA-3 Maintenance tools | No | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MP-2 Media access | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | RA-5 Vulnerability scanning | Yes | Yes + (1) | Yes + (1) (2) (3) (4) (5) (7) | Mod (1), High (1) (2) (3) (4) (5) (7)
|  |  |  | SA-4 Acquisition PROCESS | Yes | Yes + (1) (4) | Yes + (1) (2) (4) | Mod (1) (4), High (1) (2) (4)
|  |  |  | SA-8 Security engineering principles | No | Yes | Yes
|  |  |  | SA-12 Supply chain protection | No | No | Yes
|  |  |  | SA-13 Trustworthiness | No | No | Yes
|  |  |  | SC-7 Boundary protection | Yes | Yes + (1) (2) (3) (4) (5) (7) | Yes + (1) (2) (3) (4) (5) (6) (7) (8) | Mod (1) (2) (3) (4) (5) (7), High (1) (2) (3) (4) (5) (6) (7) (8)
|  |  |  | SC-26 Honeypots | No | No | No
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
|  |  |  | SC-30 Concealment and misdirection | No | No | No
|  |  |  | SC-34 Non-modifiable executable programs | No | No | No
|  |  |  | SC-35 Honey clients | No | No | No
|  |  |  | SC-37 Out-of-band channels | No | No | No
|  |  |  | SC-44 Detonation chambers | No | No | No
|  |  |  | SI-2 Flaw remediation | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | SI-4 Information system monitoring | No | Yes + (2) (4) (5) (6) | Yes + (2) (4) (5) (6) | Mod (2) (4) (5) (6), High (2) (4) (5) (6)
|  |  |  | SI-7 Software and information integrity | No | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | SI-15 Information output filtering | No | No | No
110 | NAUT - The system shall provide node authentication to ensure that only known devices can connect with each other and share data | Yes |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-7 Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-14 Permitted actions without identification or authentication | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | AC-17 Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | AC-18 Wireless access | Yes | Yes + (1) | Yes + (1) (2) (4) (5) | Mod (1) , High (1) (2) (4) (5)
|  |  |  | AC-19 Access control for mobile devices | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | AU-10 Non-repudiation | No | No | Yes
|  |  |  | CM-1 Configuration management policy and procedures | Yes | Yes | Yes
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
120 | PAUT - The system shall provide the ability to authenticate users | Yes |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-7 Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-14 Permitted actions without identification or authentication | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | AC-17 Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | AC-18 Wireless access | Yes | Yes + (1) | Yes + (1) (2) (4) (5) | Mod (1) , High (1) (2) (4) (5)
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | AU-10 Non-repudiation | No | No | Yes
|  |  |  | CM-1 Configuration management policy and procedures | Yes | Yes | Yes
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
|  |  |  | IA-2 Identification and authentication (organizational users) | Yes + (1) | Yes + (1) (2) (3) (8) | Yes + (1) (2) (3) (4) (8) (9) | Low (1), Mod (1) (2) (3) (8), High (1) (2) (3) (4) (8) (9)
|  |  |  | IA-4 Identifier management | Yes | Yes | Yes
|  |  |  | IA-5 Authenticator management | Yes + (1) | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Low (1), Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | IA-7 Cryptographic module authentication | Yes | Yes | Yes
|  |  |  | IA-8 Identification and authentication (non-organizational users) | Yes | Yes | Yes
|  |  |  | IA-10 Adaptive identification and authentication | No | No | No
|  |  |  | IA-11 Re-authentication | No | No | No
|  |  |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
130 | PLOK - The system components that maintain private data shall be physically secure | Yes |  | AC-1: Access Control Policy and Management | Yes | Yes | Yes
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | CA-7 Continuous monitoring | Yes | Yes | Yes
CP-6 Alternate storage site
|  |  |  | MP-2 Media access | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | MP-4 Media | No | Yes | Yes
|  |  |  | MP-7 Media use | No | No | No
|  |  |  | PE-1 Physical and environmental protection policy and procedures | Yes | Yes | Yes
|  |  |  | PE-2 Physical access authorizations | Yes | Yes | Yes
|  |  |  | PE-3 Physical access control | Yes | Yes | Yes + (1) | High (1)
|  |  |  | PE-4 Access control for transmission medium | No | Yes | Yes
|  |  |  | PE-5 Access control for output devices | No | Yes | Yes
|  |  |  | PE-6 Monitoring physical access | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | PE-9 Power equipment and power cabling | No | Yes | Yes
|  |  |  | PE-18 Location of information system components | No | Yes | Yes + (1) | High (1)
|  |  |  | PL-2 System security plan | Yes | Yes | Yes
|  |  |  | RA-5 Vulnerability scanning | Yes | Yes + (1) | Yes + (1) (2) (3) (4) (5) (7) | Mod (1), High (1) (2) (3) (4) (5) (7)
|  |  |  | SC-8 Transmission confidentiality and integrity | No | Yes + (1) | Yes + (1) | Mod (1), High (1)
140 | RDMP - The system software development plan shall address security support of 3rd party components throughout system life cycle | Yes |  | MA-1 System maintenance policy and procedures | Yes | Yes | Yes
|  |  |  | MA-2 Controlled maintenance | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | MA-3 Maintenance tools | No | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MA-6 Timely maintenance | No | Yes | Yes
|  |  |  | MP-1 Media protection policy and procedures | Yes | Yes | Yes
|  |  |  | MP-8 Media downgrading | No | No | No
|  |  |  | SA-1 System and services acquisition policy and procedures | Yes | Yes | Yes
|  |  |  | SA-3 System development life cycle | Yes | Yes | Yes
|  |  |  | SA-4 Acquisition PROCESS | Yes | Yes + (1) (4) | Yes + (1) (2) (4) | Mod (1) (4), High (1) (2) (4)
|  |  |  | SA-5 Information system documentation | Yes | Yes + (1) (3) | Yes + (1) (2) (3) | Mod (1) (3), High (1) (2) (3)
|  |  |  | SA-8 Security engineering principles | No | Yes | Yes
|  |  |  | SA-9 External information system services | Yes | Yes | Yes
|  |  |  | SA-10 Developer configuration management | No | Yes | Yes
|  |  |  | SA-11 Developer security testing and evaluation | No | Yes | Yes
|  |  |  | SA-12 Supply chain protection | No | No | Yes
|  |  |  | SA-15 Development PROCESS, standards and tools | No | No | No
|  |  |  | SA-16 Developer-provided training | No | No | No
|  |  |  | SA-17 Developer security architecture and design | No | No | No
|  |  |  | SA-21 Developer screening | No | No | No
150 | SAHD - System shall provide hardening features to guard against cyber attacks and malware | Yes |  | AC-19 Access control for mobile devices | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | CM-7 Least functionality | Yes | Yes | Yes
|  |  |  | SA-14 Criticality analysis | No | No | No
|  |  |  | SA-17 Developer security architecture and design | No | No | No
|  |  |  | SA-18 Tamper resistance and detection | No | No | No
|  |  |  | SC-25 Thin nodes | No | No | No
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
|  |  |  | SC-29 Heterogeneity | No | No | No
|  |  |  | SC-30 Concealment and misdirection | No | No | No
|  |  |  | SC-31 Covert channel analysis | No | No | No
|  |  |  | SC-35 Honey clients | No | No | No
|  |  |  | SC-40 Wireless link protection | No | No | No
|  |  |  | SC-41 Port and I/O device access | No | No | No
|  |  |  | SC-42 Sensor capability and data | No | No | No
|  |  |  | SC-43 Usage restrictions | No | No | No
|  |  |  | SI-11 Error handling | No | Yes | Yes
160 | SGUD - The system service/user manuals shall provide security guidance for operators and administrators | Yes |  | AC-1 Access control policy and management | Yes | Yes | Yes
|  |  |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AT-1 Security awareness and training policy and procedures | Yes | Yes | Yes
|  |  |  | AT-2 Security awareness training | Yes | Yes | Yes
|  |  |  | AT-3 Security training | Yes | Yes | Yes
|  |  |  | CP-1 Contingency planning policy and procedures | Yes | Yes | Yes
|  |  |  | CP-2 Contingency plan | Yes | Yes + (1) | Yes + (1) (2) (3) | Mod (1), High (1) (2) (3)
|  |  |  | CP-3 Contingency training | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IR-1 Incident response policy and procedures | Yes | Yes | Yes
|  |  |  | IR-2 Incident response training | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | IR-7 Incident response assistance | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-8 Incident response plan | Yes | Yes | Yes
|  |  |  | PL-1 Security planning policy and procedures | Yes | Yes | Yes
|  |  |  | PL-2 System security plan | Yes | Yes | Yes
|  |  |  | PL-4 Rules of behavior | Yes | Yes | Yes
|  |  |  | PL-7 Security concept of operations | No | No | No
|  |  |  | PL-8 Information security architecture | No | No | No
|  |  |  | PS-1 Personnel security policy and procedures | Yes | Yes | Yes
|  |  |  | SA-4 Acquisition PROCESS | Yes | Yes + (1) (4) | Yes + (1) (2) (4) | Mod (1) (4), High (1) (2) (4)
|  |  |  | SA-5 Information system documentation | Yes | Yes + (1) (3) | Yes + (1) (2) (3) | Mod (1) (3), High (1) (2) (3)
|  |  |  | SA-16 Developer-provided training | No | No | No
|  |  |  | SC-1 System and communications protection policy and procedures | Yes | Yes | Yes
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | SI-2 Flaw remediation | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | SI-4 Information system monitoring | No | Yes + (2) (4) (5) (6) | Yes + (2) (4) (5) (6) | Mod (2) (4) (5) (6), High (2) (4) (5) (6)
|  |  |  | SI-5 Security alerts, advisories, and directives | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SI-6 Security functionality VERIFICATION | No | No | Yes
|  |  |  | SI-7 Software and information integrity | No | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | SI-8 Spam protection | No | Yes | Yes + (1) | High (1)
|  |  |  | SI-10 Information input validation | No | Yes | Yes
|  |  |  | SI-11 Error handling | No | Yes | Yes
|  |  |  | SI-12 Information handling and retention | Yes | Yes | Yes
|  |  |  | SI-17 Fail-safe procedures | No | No | No
|  |  |  | PM-1 Information security program plan | Yes | Yes | Yes
|  |  |  | PM-9 RISK MANAGEMENT strategy | Yes | Yes | Yes
|  |  |  | PM-12 Insider threat program | Yes | Yes | Yes
|  |  |  | PM-14 Testing, training and monitoring | Yes | Yes | Yes
|  |  |  | PM-15 Contacts with security groups and associations | Yes | Yes | Yes
|  |  |  | PM-16 Threat awareness program | Yes | Yes | Yes
170 | STCF - The system shall establish technical controls to mitigate the potential for compromise to the integrity and confidentiality of health data stored on product or removable media | Yes |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SC-13 Cryptographic protection | Yes | Yes | Yes
|  |  |  | SC-17 Public key infrastructure certificates | No | Yes | Yes
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
180 | TXCF - The system shall ensure confidentiality of transmitted Health Data | Yes |  | PE-4 Access control for transmission medium | No | Yes | Yes
|  |  |  | SC-1 System and communications protection policy and procedures | Yes | Yes | Yes
|  |  |  | SC-8 Transmission confidentiality and integrity | No | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SC-13 Cryptographic protection | Yes | Yes | Yes
190 | TXIG - The system shall ensure integrity of transmitted Health Data | Yes |  | PE-4 Access control for transmission medium | No | Yes | Yes
|  |  |  | SC-1 System and communications protection policy and procedures | Yes | Yes | Yes
|  |  |  | SC-8 Transmission confidentiality and integrity | No | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
200 | PII | Yes |  | AP-1 Authority to Collect
AP-2 Purpose Specification
AR-1 Governance and Privacy Program
AR-2 Privacy Impact and Risk Assessment
AR-3 Privacy Requirements for Contractors and Service Providers
AR-4 Privacy Monitoring and Auditing
AR-5 Privacy Awareness and Training
AR-6 Privacy Reporting
AR-7 Privacy-enhanced system design and development
AR-8 Accounting of Disclosures
DI-1 Data Quality
DI-2 Data Integrity and Data Integrity Board
DM-1 Minimization of personally identifiable information
DM-2 Data retention and disposal
DM-3 Minimization of PII Used in Testing, Training, and IP
IP-1 Consent
IP-2 Individual Access
IP-3 Redress
IP-4 Complaint Management
SE-1 Inventory of Personally Idntifiable Information
SE-2 Privacy Incident Response
TR-1 Privacy Notice
TR-2 System of Records Notices and Privacy Act
TR-3 Dissemination of Privacy Program Information
UL-1 Internal Use
UL-2 Information Sharing with Third Parties
210 | PHI | Yes |  | 164.312(a)(2)(i) Unique User Identification (R): Assign a unique name and/or number for identifying and tracking user identity.
AC-2 Account management
AC-3: Access enforcement
IA-2 Identification and authentication (organizational users)
IA-3 Device Identification and Authentication
IA-4 Identifier management
164.312(a)(2)(ii) Emergency Access Procedure (R): Establish (and implement as needed) procedures for obtaining necessary electronic protected health information during an emergency.
AC-2 Account management
AC-3: Access enforcement
CP-2 Contingency Plan
164.312(a)(2)(iii) Automatic Logoff (A): Implement electronic procedures that terminate an electronic session after a predetermined time of inactivity.
AC-11 Session lock
AC-12 Session Termination
164.312(a)(2)(iv) Encryption and Decryption (A): Implement a mechanism to encrypt and decrypt electronic protected health information.
AC-3: Access enforcement
SC-13 Cryptographic protection
164.312(b) Audit Controls: Implement hardware, software, and/or procedural mechanisms that record and examine activity in information systems that contain or use electronic protected health information.
AU-1 Audit and accountability policy and procedures
AU-2 Audit events
AU-3 Content of audit records
AU-4 Audit storage capacity
AU-6 Audit review, analysis and reporting
AU-7 Audit reduction and report generation
164.312(c)(1) Integrity: Implement policies and procedures to protect electronic protected health information from improper alteration or destruction.
CP-9 Information system backup
MP-2 Media access
MP-5 Media Transport
SC-8 Transmission confidentiality and integrity
SI-1 System and information integrity policy and procedures
SI-7 Software and information integrity
164.312(c)(2) Mechanism to Authenticate Electronic Protected Health Information (A): Implement electronic mechanisms to corroborate that electronic protected health information has not been altered or destroyed in an unauthorized manner.
SC-8 Transmission confidentiality and integrity
SI-7 Software and information integrity
164.312(d) Person or Entity Authentication: Implement procedures to verify that a person or entity seeking access to electronic protected health information is the one claimed.
IA-2 Identification and authentication (organizational users)
IA-3 Device Identification and Authentication
IA-4 Identifier management
164.312 (e)(1) Transmission Security: Implement technical security measures to guard against unauthorized access to electronic protected health information that is being transmitted over an electronic communications network.
SC-9 Transmission Confidentiality
164.312(e)(2)(i) Integrity Controls (A): Implement security measures to ensure that electronically transmitted electronic protected health information is not improperly modified without detection until disposed of.
SC-8 Transmission confidentiality and integrity
SI-7 Software and information integrity
164.312(e)(2)(ii) Encryption (A): Implement a mechanism to encrypt electronic protected health information whenever deemed appropriate.
SC-9 Transmission Confidentiality
SC-12 Cryptographic key establishment and management
SC-13 Cryptographic protection
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## IEC
Product Security Standard Assessment - ISO IEC 62443-3-3 Controls
INSTRUCTIONS FOR USE:
1. Based upon identified User Needs and System Architecture, determine which, if any, of the 19 Security Capabilities are applicalbe to this product design. (Ref IEC/TR 80001-2-2)
A. If it is determined at a capability is not in scope for this product, provide a rationale of why it is not in scope.
2. Determine the System Security level for the product under development and provide rationale for this classification. (Ref IEC 62443-3-3 Annex A.)
3. For those Security Capabilities identified as in scope in step 1, and the System Security levels determined in step 2, evaluate and identify which Security Controls will be utilized in the design. Iinclude additional controls/requirements based upon the determined System Security Level. (Ref IEC 62443-3-3 Table B.1)
A. If a particular Security code that is specified per step 3 will not be utilized, provide a rationale.
4. If PHI and/or PII will be part of the product system design, then additional Security Controls need to be evaluated and added if applicable (Ref ID 200 & 210)
Pick your target Security Level for each chapter here:
|  | Chapter | Target Security Level | Security Level Rationale:
|  | 1 | SL-3
|  | 2 | SL-3
|  | 3 | SL-2
|  | 4 | SL-2
|  | 5 | SL-1
|  | 6 | SL-2
|  | 7 | SL-1
| IEC 80001-2-2 |  |  | IEC 80001-2-8
|  |  |  | System Security level: ISO IEC 62443-3-3 SL-1 SL-2 SL-3 SL-4 | [8 empty] | System Security level Rationale
ID | Security Capability (Applicable = Design Inputs) | In scope of the product? | Rationale (if not utilizing capability) | Security Controls | Chapter | Target Level | System Security Level Baseline SL-1 SL-2 SL-3 SL-4 |  |  |  | Additional Control/Requirement Enhancements based upon System Security Level | Rationale for not utilizing control if identified per Security Level | Traceability reference (SRS,SDD, etc)
10 | ALOF - System shall provide automatic logoff after a period of inactivity | Yes |  | SR 2.5 Session Lock | 2 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 2.6 Remote session termination | 2 | SL-3 | No | Yes | Yes | Yes
20 | AUDT - System shall provide audit controls documenting who is doing what with HEALTH DATA | Yes |  | SR 2.8 Auditable events | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.9 Audit storage capacity | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.10 Response to audit processing failures | 2 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 2.11 Timestamps | 2 | SL-3 | No | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 2.12 Non-repudiation | 2 | SL-3 | No | No | Yes | Yes + (1) | SL-4 (1)
|  |  |  | SR 3.9 Protection of audit information | 3 | SL-2 | No | Yes | Yes | Yes + (1) | SL-4 (1)
|  |  |  | SR 6.1 Audit reduction and report generation | 6 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 6.2 Continuous monitoring | 6 | SL-2 | No | Yes | Yes | Yes
30 | AUTH - The system shall provide role based controlled access to HEALTH DATA and functions - access to be provided only as necessary to perform the tasks required consistent with intended use | Yes |  | SR 1.3 Account management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.1 Authorization enforcement | 2 | SL-3 | Yes | Yes + (1) | Yes + (1) (2) (3) | Yes + (1) (2) (3) (4) | SL-2 (1) (2), SL-3 (1) (2) (3), SL-4 (1) (2) (3) (4)
40 | CNFS - System shall provide the Hospital system administrator the ability to configure the products security capabilities | Yes |  | SR 1.3 Account management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 7.6 Network and security configuration settings | 7 | SL-1 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
50 | CSUP - The system shall provide the ability of on-site service staff, remote service staff, or authorized customer staff to install/upgrade device's security patches. | Yes |  | No applicable SECURITY CONTROLS | 0 | #N/A | N/A | N/A | N/A | N/A
60 | DIDT - The system shall provide the ability to directly remove information that allows identification of a person. | Yes |  | SR 4.2 Information persistence | 4 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
70 | DTBK - The system shall have an integral data backup capability to recover after damage or destruction of device data, hardware or software. | Yes |  | SR 2.8 Auditable events | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 3.6 Deterministic output | 3 | SL-2 | Yes | Yes | Yes | Yes
|  |  |  | SR 7.3 Control system backup | 7 | SL-1 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2)
|  |  |  | SR 7.4 Control system recovery and reconstitution | 7 | SL-1 | Yes | Yes | Yes | Yes
80 | EMRG - The system shall provide users the ability to access private data in case of an emergency that requires immediate access to private data. | Yes |  | SR 1.4 Identifier Management | 1 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.8 Auditable events | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
90 | IGAU - The device shall ensure the integrity of stored data with implicit/explicit error detection/correction technology | Yes |  | SR 3.1 Communication Integrity | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 3.3 Security functionality VERIFICATION | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 3.4 Software and information integrity | 3 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 3.5 Input validation | 3 | SL-2 | Yes | Yes | Yes | Yes
|  |  |  | SR 3.7 Error handling | 3 | SL-2 | No | Yes | Yes | Yes
100 | MLDP - The system shall support the use of anti-malware mechanism | Yes |  | SR 1.2 Software PROCESS and device identification and authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.3 Use Control for portable and mobile devices | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 3.2 Malicious Code protection | 3 | SL-2 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2)
|  |  |  | SR 3.3 Security functionality VERIFICATION | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 5.3 General purpose person-to-person communication restrictions | 5 | SL-1 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 6.2 Continuous monitoring | 6 | SL-2 | No | Yes | Yes | Yes
110 | NAUT - The system shall provide node authentication to ensure that only known devices can connect with each other and share data | Yes |  | SR 1.2 Software PROCESS and device identification and authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.3 Account Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.4 Identifier Management | 1 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.6 Wireless access management | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1)
|  |  |  | SR 1.8 Public Key Infrastructure (PKI) certificates | 1 | SL-3 | No | Yes | Yes | Yes
|  |  |  | SR 1.9 Strength of Public Key Authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.10 Authenticator Feedback | 1 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 1.11 Unsuccessful Login Attempts | 1 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 1.13 Access via Untrusted Networks | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1)
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes
120 | PAUT - The system shall provide the ability to authenticate users | Yes |  | SR 1.1 Human user identification and authentication | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3)
|  |  |  | SR 1.3 Account Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.4 Identifier Management | 1 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.6 Wireless access management | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1)
|  |  |  | SR 1.7 Strength of Password-Based Authentication | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 1.8 Public Key Infrastructure (PKI) certificates | 1 | SL-3 | No | Yes | Yes | Yes
|  |  |  | SR 1.9 Strength of Public Key Authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.10 Authenticator Feedback | 1 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 1.11 Unsuccessful Login Attempts | 1 | SL-3 | Yes | Yes | Yes | Yes
|  |  |  | SR 1.13 Access via Untrusted Networks | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1)
|  |  |  | SR 2.3 Use Control for portable and mobile devices | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.8 Auditable events | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.11 Timestamps | 2 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 2.12 Non-repudiation | 2 | SL-3 | No | No | Yes | Yes | SL-4 (1)
|  |  |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes
|  |  |  | SR 6.2 Continuous monitoring | 6 | SL-2 | No | Yes | Yes | Yes
|  |  |  | SR 1.12 – System use notification | 1 | SL-3 | Yes | Yes | Yes | Yes
130 | PLOK - The system components that maintain private data shall be physically secure | Yes |  | SR 1.1 Human user identification and authentication | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3)
|  |  |  | SR 1.3 Account Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 7.7 Least functionality | 7 | SL-1 | Yes | Yes | Yes | Yes
140 | RDMP - The system software development plan shall address security support of 3rd party components throughout system life cycle | Yes |  | SR 4.2 Information persistence | 4 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
150 | SAHD - System shall provide hardening features to guard against cyber attacks and malware | Yes |  | SR 2.1 Authorization Enforcement | 2 | SL-3 | Yes | Yes + (1) (2) | Yes + (1) (2) (3) | Yes + (1) (2) (3) (4) | SL-2 (1) (2), SL-3 (1) (2) (3), SL-4 (1) (2) (3) (4)
|  |  |  | SR 2.2 Wireless use control | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.3 Use Control for portable and mobile devices | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 3.4 Software and information integrity | 3 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 5.1 Network Segmentation | 5 | SL-1 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3)
|  |  |  | SR 5.2 Zone boundary protection | 5 | SL-1 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3)
|  |  |  | SR 5.3 General purpose person-to-person communication restrictions | 5 | SL-1 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.4 – Mobile code | 2 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 5.4 Application partitioning | 5 | SL-1 | Yes | Yes | Yes | Yes
|  |  |  | SR 7.1 Denial of service protection | 7 | SL-1 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2)
|  |  |  | SR 7.2 Resource management | 7 | SL-1 | Yes | Yes | Yes | Yes
|  |  |  | SR 7.5 Emergency power | 7 | SL-1 | Yes | Yes | Yes | Yes
|  |  |  | SR 7.7 Least functionality | 7 | SL-1 | Yes | Yes | Yes | Yes
160 | SGUD - The system service/user manuals shall provide security guidance for operators and administrators | Yes |  | No applicable SECURITY CONTROLS | 0 | #N/A | N/A | N/A | N/A | N/A
170 | STCF - The system shall establish technical controls to mitigate the potential for compromise to the integrity and confidentiality of health data stored on product or removable media | No |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes
180 | TXCF - The system shall ensure confidentiality of transmitted Health Data | Yes |  | SR 1.8 Public Key Infrastructure (PKI) certificates | 1 | SL-3 | No | Yes | Yes | Yes
|  |  |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes
190 | TXIG - The system shall ensure integrity of transmitted Health Data | Yes |  | SR 3.1 Communication Integrity | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 3.8 Session Integrity | 3 | SL-2 | No | Yes | Yes + (1) (2) | Yes + (1) (2) (3) | SL-3 (1) (2), SL-4 (1) (2) (3)
200 | PII | Yes |  | AP-1 Authority to Collect
AP-2 Purpose Specification
AR-1 Governance and Privacy Program
AR-2 Privacy Impact and Risk Assessment
AR-3 Privacy Requirements for Contractors and Service Providers
AR-4 Privacy Monitoring and Auditing
AR-5 Privacy Awareness and Training
AR-6 Privacy Reporting
AR-7 Privacy-enhanced system design and development
AR-8 Accounting of Disclosures
DI-1 Data Quality
DI-2 Data Integrity and Data Integrity Board
DM-1 Minimization of personally identifiable information
DM-2 Data retention and disposal
DM-3 Minimization of PII Used in Testing, Training, and IP
IP-1 Consent
IP-2 Individual Access
IP-3 Redress
IP-4 Complaint Management
SE-1 Inventory of Personally Idntifiable Information
SE-2 Privacy Incident Response
TR-1 Privacy Notice
TR-2 System of Records Notices and Privacy Act
TR-3 Dissemination of Privacy Program Information
UL-1 Internal Use
UL-2 Information Sharing with Third Parties
210 | PHI | Yes |  | 164.312(a)(2)(i) Unique User Identification (R): Assign a unique name and/or number for identifying and tracking user identity.
AC-2 Account management
AC-3: Access enforcement
IA-2 Identification and authentication (organizational users)
IA-3 Device Identification and Authentication
IA-4 Identifier management
164.312(a)(2)(ii) Emergency Access Procedure (R): Establish (and implement as needed) procedures for obtaining necessary electronic protected health information during an emergency.
AC-2 Account management
AC-3: Access enforcement
CP-2 Contingency Plan
164.312(a)(2)(iii) Automatic Logoff (A): Implement electronic procedures that terminate an electronic session after a predetermined time of inactivity.
AC-11 Session lock
AC-12 Session Termination
164.312(a)(2)(iv) Encryption and Decryption (A): Implement a mechanism to encrypt and decrypt electronic protected health information.
AC-3: Access enforcement
SC-13 Cryptographic protection
164.312(b) Audit Controls: Implement hardware, software, and/or procedural mechanisms that record and examine activity in information systems that contain or use electronic protected health information.
AU-1 Audit and accountability policy and procedures
AU-2 Audit events
AU-3 Content of audit records
AU-4 Audit storage capacity
AU-6 Audit review, analysis and reporting
AU-7 Audit reduction and report generation
164.312(c)(1) Integrity: Implement policies and procedures to protect electronic protected health information from improper alteration or destruction.
CP-9 Information system backup
MP-2 Media access
MP-5 Media Transport
SC-8 Transmission confidentiality and integrity
SI-1 System and information integrity policy and procedures
SI-7 Software and information integrity
164.312(c)(2) Mechanism to Authenticate Electronic Protected Health Information (A): Implement electronic mechanisms to corroborate that electronic protected health information has not been altered or destroyed in an unauthorized manner.
SC-8 Transmission confidentiality and integrity
SI-7 Software and information integrity
164.312(d) Person or Entity Authentication: Implement procedures to verify that a person or entity seeking access to electronic protected health information is the one claimed.
IA-2 Identification and authentication (organizational users)
IA-3 Device Identification and Authentication
IA-4 Identifier management
164.312 (e)(1) Transmission Security: Implement technical security measures to guard against unauthorized access to electronic protected health information that is being transmitted over an electronic communications network.
SC-9 Transmission Confidentiality
164.312(e)(2)(i) Integrity Controls (A): Implement security measures to ensure that electronically transmitted electronic protected health information is not improperly modified without detection until disposed of.
SC-8 Transmission confidentiality and integrity
SI-7 Software and information integrity
164.312(e)(2)(ii) Encryption (A): Implement a mechanism to encrypt electronic protected health information whenever deemed appropriate.
SC-9 Transmission Confidentiality
SC-12 Cryptographic key establishment and management
SC-13 Cryptographic protection
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## System Security Level
REFERENCE FOR SYSTEM SECURITY LEVEL
NIST SP 800-53:
SC information type = {(confidentiality, impact), (integrity, impact), (availability, impact)},
The potential impact is LOW if—
− The loss of confidentiality, integrity, or availability could be expected to have a limited adverse effect on organizational operations, organizational assets, or individuals.2
AMPLIFICATION: A limited adverse effect means that, for example, the loss of confidentiality, integrity, or availability might: (i) cause a degradation in mission capability to an extent and duration that the organization is able to perform its primary functions, but the effectiveness of the functions is noticeably reduced; (ii) result in minor damage to organizational assets; (iii) result in minor financial loss; or (iv) result in minor harm to individuals.
The potential impact is MODERATE if—
− The loss of confidentiality, integrity, or availability could be expected to have a serious adverse effect on organizational operations, organizational assets, or individuals.
AMPLIFICATION: A serious adverse effect means that, for example, the loss of confidentiality, integrity, or availability might: (i) cause a significant degradation in mission capability to an extent and duration that the organization is able to perform its primary functions, but the effectiveness of the functions is significantly reduced; (ii) result in significant damage to organizational assets; (iii) result in significant financial loss; or (iv) result in significant harm to individuals that does not involve loss of life or serious life threatening injuries.
The potential impact is HIGH if—
− The loss of confidentiality, integrity, or availability could be expected to have a severe or catastrophic adverse effect on organizational operations, organizational assets, or individuals.
AMPLIFICATION: A severe or catastrophic adverse effect means that, for example, the loss of confidentiality, integrity, or availability might: (i) cause a severe degradation in or loss of mission capability to an extent and duration that the organization is not able to perform one or more of its primary functions; (ii) result in major damage to organizational assets; (iii) result in major financial loss; or (iv) result in severe or catastrophic harm to individuals involving loss of life or serious life threatening injuries.
A low-impact system is an information system in which all three of the security objectives are low.
A moderate-impact system is an information system in which at least one of the security objectives is moderate and no security objective is greater than moderate.
A high-impact system is an information system in which at least one security objective is high.
IEC 62443-3-3
Security Level Vector: SL-?([FR,]domain) = { IAC UC SI DC RDF TRE RA } Example: SL-C ("Product") = {3 2 2 1 1 3 2}
SL= Security Level
FR = Foundational Requirement
IAC = Identification and authentication control
UC = Use control
SI = System integrity
DC = Data confidentiality
RDF = Restricted data flow
TRE = Timely response to events
RA = Resource availability
RE = Requirement Enhancement
SL-? = (Required) The SL type. SL-C = Capabilities SL [FR,] = (Optional) Field indicating the FR that the SL value applies. The FRs are written out in abbreviated form instead of numerical form to aid in readability. domain = (Required) The applicable domain that the SL applies. Domains can refer to zones, control systems, subsystems or components.In this particular standard, all requirements refer to a control system, so the domain term is not used as it would be by other documents in the IEC 62443 series.
General description and distinctions between Security Levels
SL 0: No specific requirements or security protection necessary SL 0 has multiple meanings depending on the situation in which it is applied. In defining SL-C it would mean that the component or system fails to meet some of the SL 1 requirements for that particular FR. This would most likely be for components or systems that would be part of a larger zone where other components or systems would provide compensating countermeasures. In defining SL-T for a particular zone it means that the asset owner has determined that the results of their risk analysis indicate that less than the full SL 1 specific requirements are necessary for that particular FR on that component or system. This would more likely happen for individual components within a system or zone that do not contribute in any way to the FR-specific requirements. In defining SL-A it would mean that the particular zone fails to meet some of the SL 1 requirements for that particular FR.
SL 1: Protection against casual or coincidental violation Casual or coincidental violations of security are usually through the lax application of security policies. These can be caused by well-meaning employees just as easily as they can be by an outsider threat. Many of these violations will be security program related and will be handled by enforcing policies and procedures.A simple example would be an operator able to change a set point to a value outside certain conditions determined by the engineering staff. The system did not enforce the proper authentication and use control restrictions to disallow the change by the operator. Another example would be a password being sent in clear text, allowing a network engineer to view the password while troubleshooting the system. The system did not enforce proper data confidentiality to protect the password.
SL 2: Protection against intentional violation using simple means with low resources, generic skills and low motivation Simple means do not require much knowledge on the part of the attacker. The attacker does not need detailed knowledge of security, the domain or the particular system under attack. These attack vectors are well known and there may be automated tools for aiding the attacker. They are also designed to attack a wide range of systems instead of targeting a specific system, so an attacker does not need a significant level of motivation or resources at hand. An example would be a virus that infects a workstation spreading to another engineering workstation since they both use the same general purpose operating system. Another example would be an attacker compromising a web server in the enterprise network by an exploit downloaded from the Internet for a publicly known vulnerability in the general purpose operating system of the web server. The attacker uses the web server as a pivot point in an attack against other systems in the enterprise network as well as the industrial network. A third example would be an operator that views a website which downloads a Trojan that opens a hole in the routers and firewalls to the Internet.
SL 3: Protection against intentional violation using sophisticated means with moderate resources, IACS specific skills and moderate motivation Sophisticated means require advanced security knowledge, advanced domain knowledge, advanced knowledge of the target system or any combination of these. An attacker going after a SL 3 system will likely be using attack vectors that have been customized for the specific target system. The attacker may use exploits in operating systems that are not well known, weaknesses in industrial protocols, specific information about a particular target to violate the security of the system or other means that require a greater motivation as well as skill and knowledge set than are required for SL 1 or 2. An example of sophisticated means could be password or key cracking tools based on hash tables. These tools are available for download, but applying them takes knowledge of the system (such as the hash of a password to crack). Using Figure A.1, another example would be an attacker that gains access to the FS-PLC through the serial conduit after gaining access to the control PLC through a vulnerability in the Ethernet controller. Using Figure A.2, a third example would be an attacker that gains access to the data historian by using a brute-force attack through the industrial/enterprise DMZ firewall initiated from the enterprise wireless network.
SL 4: Protection against intentional violation using sophisticated means with extended resources, IACS specific skills and high motivation SL 3 and SL 4 are very similar in that they both involve sophisticated means used to violate the security requirements of the system. The difference comes from the attacker being even more motivated and having extended resources at their disposal. These may involve high-performance computing resources, large numbers of computers or extended periods of time. An example of sophisticated means with extended resources would be using super computers or computer clusters to conduct brute-force password cracking using large hash tables. Another example would be a botnet used to attack a system using multiple attack vectors at once. A third example would be an organized crime organization that has the motivation and resources to spend weeks attempting to analyze a system and develop custom “zero-day” exploits.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.