## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Formula 3D Knee Planning Software
Scope: |  | Formula 3D Knee Planning is indicated for pre-operative planning for a knee replacement surgery where Stryker’s Triathlon knee implant is used. The Formula 3D Knee Planning is intended to provide a surgeon facing, easy to use CT Knee planning software, that uses the patient’s CT scans to visualize the disease condition of Knee in three dimension and enable effective decision making for the surgeons before they even go into the operating room on the day of the surgery. Functionality Includes: • Auto-segmentation and landmark identification (manual modification possible) • Effortless implant planning • Saves the planned report for quick reference, The surgeon has the control and can choose how to interpret and use the results from the pre-operative planning.
Description |  | Product Security Risk Table
Document Number and Revision |  | D005010014 Rev 0
Date: |  | 2021-11-30
Conducted by: |  | Deepak Sharma
ID # | Asset Type (Information/Physical) | Asset | Asset Description | Comments
A01 | Physical Asset | User Laptop | Utilizing computer resources and computing power by adversary, allows various general purpose attacks, such as incl. denial of service attacks (Ransomware deployment, Bitcoin Mining, abuse of peripheral devices such as WebCam, Microphones, etc., ). | Outside of Trust Boundary. May include guidance in SOM for this asset, if appriopriate
A02 | Physical Asset | USB Drive | Utilizing usb drive by adversary, allows various general purpose attacks, such malware , virus attacks etc | Outside of Trust Boundary. May include guidance in SOM for this asset, if appriopriate
A03 | Information asset | User Password / Credentials | User Passwords/Credentials provide access to ePHI, Software Binaries, read access to config files.
A04 | Information asset | THOR Knee Planning Application | Contains Knee Planning software which controls the overall features for segmentation, landmark planning, implant size calculation etc.
A05 | Information asset | Patient Personal and Health data | DICOM data stored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A06 | Information asset | Laptop Operating System | Information about internals of the system (Device identification, software versions, supported protocols, etc.) | Outside of Trust Boundary. May include guidance in SOM for this asset, if appriopriate
A07 | Information asset | Implant Database | Database of Stryker implants
A08 | Physical Asset | Network Location | Network location where the CT Dicom data is stored | Outside of Trust Boundary. May include guidance in SOM for this asset, if appriopriate
A09 | Information Asset | Application Settings | Application specific configuration settings
A10 | Information Asset | User Settings | User specific configuration settings
A11 | Information Asset | Log Files | Log files containing application errors and debug logs
A12 | Information Asset | Statistical Model | Statistical Model Data is utilized by auto-segmentation and bone referencing algorithms.
A13 | Information asset | Translation data | GUI specific japanese translation files
A14 | Information asset | User License data | Application instance specific license data
A15 | Information asset | Security keys, tokens, certificates | Encryption keys and tokens
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
D05788-1, Ver 1

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user credentials | Yes | Customers are required to manage the Laptop/Windows credentials
V02 | Improper Authentication | Yes
V03 | Improper Input Validation | Yes
V04 | Unprotected Storage of Credential | Yes
SBOM
V11 | Unpatched COTS operating system | Yes | Sofware is Installed on the end customer laptop so Operating System is provided by the customer for Software Installation
V12 | UnProtected Application Binaries | Yes | n/a
V13 | Ineffective patch management | Yes | Security Operations Manual
V14 | Lack of software vulnerability management | Yes
V15 | 3rd Party Component Dependency | Yes
V16 | Embed Virus into Binary or Embedded Malicious Code | Yes
V17 | Hidden Functionality | Yes
Access points
V21 | Unprotected network port | Yes | n/a
V22 | Unprotected external USB Port | Yes | n/a
V23 | Unprotected external CD/DVD Drive | Yes | n/a
V24 | Unprotected Physical storage of Laptop | Yes | n/a
V25 | Unprotected public network connections | Yes
V26 | Lack of monitoring mechanisms | Yes
Data
V31 | Insecure Storage of Sensitive Information | Yes | n/a
V32 | Unencrypted ePHI in flight | No | No network related transmission of data by the application
V33 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength | Yes
V34 | Exposure of Private Information | Yes
V35 | Sensitive Data Storage in Improperly Locked Memory | Yes
V36 | Lack of fault reports recorded in administrator and operator logs | Yes
V37 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging | Yes
V38 | Information Exposure Through Log Files | Yes
InSecure Configurations of Resources
V40 | Lack of efficient configuration change control | Yes
V41 | Malicious Software Download | Yes
V42 | Malicious Automated Software Update | Yes
V43 | External Control of System or Configuration Setting | Yes
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification | Column1
1 | T01 | Deliver undirected malware | V11 | Unpatched COTS operating system | A06 | Laptop Operating System | Access to Thor knee Planning application impacted due to system being down or functional degradation |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1.Application Whitelisting 2.Firewall 3.Virus Scan 4.User authentication 5.Proper crash logs of application should be maintained | [1.1.] SOM- 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [2.1] SOM- 13.3 Firewall Implementation [3.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [3.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [4.1.] User authentication is Managed through Windows login SADD: ES05, ES06 [4.2] SOM - 16. PERSON AUTHENTICATION (PAUT) [4.3] SOM - 16.2 User Account Management [4.4] SOM- 7.1 Access Prevention (AUTH-1) [5.1.] Debug logs, Audit logs. Audit logs are encrypted. SRS: 2.17.10 SADD: DS13 SDD: ML11 "Debug logs, Audit logs. Audit logs are encrypted. T0-3598, T0-1830 [5.2] SOM - 6. AUDIT CONTROLS (AUDT) [5.3] SOM- 6.3 Audit Log Protection (AUDT-7)" | [1.1.1] SOM- 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [3.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [4.1.1] PR- PY-SYK-CC-001 User Authentication [4.2.1] SOM - 16. PERSON AUTHENTICATION (PAUT) [4.3.1] SOM - 16.2 User Account Management [4.4.1] SOM- 7.1 Access Prevention (AUTH-1) [5.1.1] PR- PY-SYK-CC-002 Proper Crash Logs are Maintained Locally in Encrypted Form [5.2.1] SOM - 6. AUDIT CONTROLS (AUDT) [5.3.1] SOM- 6.3 Audit Log Protection (AUDT-7) | None | None | Low | Network | Low | None | None | Unchanged | 3.887 | 0.22 | 1.4124 | 5.3 | 2.2 | LOW
2 | T01 | Deliver undirected malware | V13 | Ineffective patch management | A06 | Laptop Operating System | Access to Thor knee Planning application impacted due to system being down or functional degradation |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1. Virus Scan 2. Firewall 3. Windows security update install 4. Application should be updated with the latest version of dependencies. | [1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [1.3] Access management - Windows user MSI - SRS-2.22.1, 2.22.2, 2.10.3 SADD- PKG1, PKG2, PL01,ES05, ES06 [2.1] SOM- 13.3 Firewall Implementation [3.1] SOM - 8.3 Operating System Updates (CSUP-2) [4.1] SDP - 5.1.4 | [1.1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [1.3.1] PR - PY-SYK-CC-003 Application Dependency Check [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM - 8.3 Operating System Updates (CSUP-2) | None | None | Low | Network | Low | None | None | Unchanged | 3.887 | 0.22 | 1.4124 | 5.3 | 2.2 | LOW
4 | T01 | Deliver undirected malware | V14 | Lack of software vulnerability management | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1. Update drivers 2. Virus Scan 3. Input validation and proper access management in application 4. Application should not be allowed to get installed in shared directory/path. | [1.1] SOM - 8.3 8.3 Operating System Updates (CSUP-2) [2.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [2.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [3.1] Input data validation: SRS:2.2.1, 2.2.8, 2.5.27, 2.5.26 SADD: PL01, ML09 SDD: C231,C279, C278 [3.2] Access management - Windows user SADD: ES05, ES06 [4.1] Installation: SRS: 2.22.1, 2.22.2, 2.10.3; SADD: PKG1, PKG2 | [1.1.1] PR-PY-SYK-CC-003 Application Dependency Check [1.1.2] SOM - 8.3 8.3 Operating System Updates (CSUP-2) [2.1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [2.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [3.1.1] PR-PY-SYK-CC-005 Input Validation Check [3.2.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management [4.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management | None | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.22 | 1.4124 | 3.7 | 1.9 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
5 | T01 | Deliver undirected malware | V15 | 3rd Party Component Dependency | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1. The 3rd party modules and dependencies should be updated with latest versions. 2. Application upgrade should verify if all required dependencies are updated in system. 3. Application upgrade should detect any upgrade or change in dependecies on user system and show dependency issue message to application user. | [1.1] SDP - D005010020 - 5.1.4 [2.1] SDP - D005010020 - 5.1.4 [3.1] SDP - D005010020 - 5.1.4 | [1.1.1] Pentest Report [2.1.1] Pentest Report [3.1.1] Pentest Report | None | None | Low | Network | Low | None | None | Unchanged | 3.887 | 0.22 | 1.4124 | 5.3 | 2.2 | LOW
| T01 | Deliver undirected malware | V25 | Unprotected public network connections | A06 | Laptop Operating System | Access to Thor knee Planning application impacted due to system being down or functional degradation |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1. Firewall 2. Virus Scan 3. The internet communication of Thor application should be secured using SSL or other secure protocols. 4. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. | [1.1] SOM- 13.3 Firewall Implementation [2.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [2.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [3.1] Not applicable as no internet connection is getting used. [4.1] Not applicable as no internet connection is getting used. | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [2.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [3.1.1] Not applicable for v1.0 as no internet connection is getting used. [4.1.1] Not applicable for v1.0 as no internet connection is getting used. | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | 4.2 | MEDIUM
| T01 | Deliver undirected malware | V12 | UnProtected Application Binaries | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1. The application binaries should be obfuscated and should not take any runtime inputs from user console. 2. The permission to binaries should be appropriate so that only authorised system user should only have access. | [1.1] Obfuscation of code. [2.1] SAD - PKG1 | [1.1.1] PR - PY-SYK-CC-004 Obfuscation of Code [2.1.1] PR - PY-SYK-CC-006 Secured Application Path and Proper Access Management | None | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.22 | 1.4124 | 3.7 | 1.9 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T01 | Deliver undirected malware | V22 | Unprotected external USB Port | A06 | Laptop Operating System | Access to Thor knee Planning application impacted due to system being down or functional degradation |  | Low | None | None | Local | Low | None | None | Unchanged | 2.5151 | 0.22 | 1.4124 | 4 | Low | 0.2 | 2 | LOW | 1. Virus Scan 2. Firewall 3. Windows security update install 4. Application should be updated with the latest version of dependencies. | [1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [1.3] Access management - Windows user MSI - SRS-2.22.1, 2.22.2, 2.10.3 SADD- PKG1, PKG2, PL01,ES05, ES06 [2.1] SOM- 13.3 Firewall Implementation [3.1] SOM - 8.3 Operating System Updates (CSUP-2) [3.2] SOM - 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) | [1.1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [1.3.1] Access management - Windows user MSI - SRS-2.22.1, 2.22.2, 2.10.3 SADD- PKG1, PKG2, PL01,ES05, ES06 [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM - 8.3 Operating System Updates (CSUP-2) [3.2.1] SOM - 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [3.2.2] PR - PY-SYK-CC-003 Application Dependency Check | Low | None | None | Local | Low | None | None | Unchanged | 2.5151 | 0.22 | 1.4124 | 4 | 2 | LOW
2 | T02 | Deliver directed malware | V21 | Unprotected network port | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Network | High | None | Required | Unchanged | 1.6201 | 0.52545 | 3.3734 | 5 | Low | 0.2 | 3.7 | LOW | 1. Application Whitelisting 2. Firewall 3. Virus Scan 4. User authentication | [1.1] SOM- 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [2.1] SOM-13.3 Support for Firewall [3.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [3.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [4.1.] User authentication is Managed through Windows login SADD: ES05, ES06 [4.2] SOM - 16. PERSON AUTHENTICATION (PAUT) [4.3] SOM - 16.2 User Account Management [4.4] SOM- 7.1 Access Prevention (AUTH-1) | [1.1.1] SOM- 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [3.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [4.1.1] PR- PY-SYK-CC-001 User Authentication [4.2.1] SOM - 16. PERSON AUTHENTICATION (PAUT) [4.3.1] SOM - 16.2 User Account Management [4.4.1] SOM- 7.1 Access Prevention (AUTH-1) | Low | Low | Low | Network | High | None | Required | Unchanged | 1.6201 | 0.52545 | 3.3734 | 5 | 3.7 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T02 | Deliver directed malware | V22 | Unprotected external USB Port | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | High | None | Required | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | 1. Virus Scan 2. Firewall 3. Windows security update install 4. Application should be updated with the latest version of dependencies. | [1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1] SOM- 13.3 Firewall Implementation [3.1] SOM - 8.3 Operating System Updates (CSUP-2) | [1.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM - 8.3 Operating System Updates (CSUP-2) | Low | Low | Low | Local | High | None | Required | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | 3.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T02 | Deliver directed malware | V23 | Unprotected external CD/DVD Drive | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | High | None | Required | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | 1. Virus Scan | [1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) | [1.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) | Low | Low | Low | Local | High | None | Required | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | 3.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T02 | Deliver directed malware | V24 | Unprotected Physical storage of Laptop | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | High | None | Required | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | 1. Virus Scan 2. Firewall 3. Windows security update install. 4. Application should be updated with the latest version of dependencies. | [1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1] SOM- 13.3 Firewall Implementation [3.1] SOM - 8.3 Operating System Updates (CSUP-2) | [1.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM - 8.3 Operating System Updates (CSUP-2) | Low | Low | Low | Local | High | None | Required | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | 3.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T02 | Deliver directed malware | V25 | Unprotected public network connections | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Network | High | None | Required | Unchanged | 1.6201 | 0.52545 | 3.3734 | 5 | Low | 0.2 | 3.7 | LOW | 1. Application Whitelisting 2. Firewall 3. Virus Scan 4. User authentication | [1.1.] SOM- 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [2.1] SOM- 13.3 Firewall Implementation [3.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [3.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [4.1.] User authentication is Managed through Windows login SADD: ES05, ES06 [4.2] SOM - 16. PERSON AUTHENTICATION (PAUT) [4.3] SOM - 16.2 User Account Management [4.4] SOM- 7.1 Access Prevention (AUTH-1) | [1.1.1] SOM- 13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [3.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [4.1.1] PR- PY-SYK-CC-001 User Authentication [4.2.1] SOM - 16. PERSON AUTHENTICATION (PAUT) [4.3.1] SOM - 16.2 User Account Management [4.4.1] SOM- 7.1 Access Prevention (AUTH-1) | Low | Low | Low | Network | High | None | Required | Unchanged | 1.6201 | 0.52545 | 3.3734 | 5 | 3.7 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T02 | Deliver directed malware | V12 | UnProtected Application Binaries | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | High | Low | Required | Unchanged | 0.76466 | 0.52545 | 3.3734 | 4.2 | Low | 0.2 | 3.6 | LOW | 1. The application binaries should be obfuscated and should not take any runtime inputs from user console. 2. The permission to binaries should be appropriate so that only authorised system user should have access. | [1.1] SOM- 6. AUDIT CONTROLS (AUDT) [1.2] SOM- 6.3 Audit Log Protection (AUDT-7) [1.3] SOM- 6.4 Log Export, Use, and Notification (AUDT-5, 6, 8) [1.4]25.3 Security Testing [2.1] Access management - Windows user SADD: ES05, ES06 | [1.1.1] SOM- 6. AUDIT CONTROLS (AUDT) [1.2.1] SOM- 6.3 Audit Log Protection (AUDT-7) [1.3.1] SOM- 6.4 Log Export, Use, and Notification (AUDT-5, 6, 8) [1.4.1] SOM- 25.3 Security Testing [2.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management | Low | Low | Low | Local | High | Low | Required | Unchanged | 0.76466 | 0.52545 | 3.3734 | 4.2 | 3.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
3 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Implant Database | Loss of Stryker IP |  | Low | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. Firewall 2. The internet communication of Thor application should be secured using SSL or other secure protocols. 3. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. | [1.1] SOM- 13.3 Firewall Implementation [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable as no internet connection is getting used. | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] Not applicable as no internet connection is getting used. | Low | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.52545 | 3.3734 | 3.9 | 3.5 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A14 | User License data | Loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.7323 | 4.7014 | 5.3 | Low | 0.2 | 4.9 | MEDIUM | 1. Firewall 2. The internet communication of Thor application should be secured using SSL or other secure protocols. 3. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. 4. The License should be encrypted for communication over internet. | [1.1] SOM- 13.3 Firewall Implementation [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable as no internet connection is getting used. [4.1]License file encryption is done using AES SRS:2.17.13, 2.17.4 SADD: ES04, DS10, C277 | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] Not applicable as no internet connection is getting used. [4.1.1] License file encryption is done using AES [4.1.2] PR - PY-SYK-CC-010 Use of Secure Crypto | Low | Low | None | Network | High | High | Required | Unchanged | 0.51463 | 0.3916 | 2.5141 | 3.1 | 2.7 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A15 | Security keys, tokens, certificates | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | High | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.7323 | 4.7014 | 5.3 | Low | 0.2 | 4.9 | MEDIUM | 1. The internet communication of Thor application should be secured using SSL or other secure protocols. 2. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. 3. The keys and token should be stored using windows credential manager. | [1.1] Not applicable as no internet connection is getting used. [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable | [1.1.1] Not applicable as no internet connection is getting used. [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] PR - PY-SYK-CC-001 User Authentication [3.1.2] PR - PY-SYK-CC-010 Use of Secure Crypto | Low | Low | None | Network | High | High | Required | Unchanged | 0.51463 | 0.3916 | 2.5141 | 3.1 | 2.7 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T05 | Unauthorized modifications to Knee Planning Application | V16 | Embed Virus into Binary or Embedded Malicious Code | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Physical | Low | High | Required | Unchanged | 0.21191 | 0.52545 | 3.3734 | 3.6 | Low | 0.2 | 3.5 | LOW | 1. Virus scan 2. Application should verify the checksum of all the DLL files and executable inside application folder before starting main thread. 3. Proper Code review of applciation before releasing. 4. Anti-dynamic instrumentation techniques should be implimented in application. | [1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1] Checksum of DLL: TBD [3.1] Code review: process followed. [4.1] Anti dynamic instrumentation: TBD | [1.1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1.1] Checksum of DLL: TBD [3.1.1] Code review: process followed. [4.1.1] PR - PY-SYK-CC-007 Dynamic Instrumentation | Low | Low | Low | Physical | Low | High | Required | Unchanged | 0.21191 | 0.52545 | 3.3734 | 3.6 | 3.5 | LOW
| T05 | Unauthorized modifications to Knee Planning Application | V17 | Hidden Functionality | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. Application should not have direct access to hidden functionalities. It is recommended to delete all unused funtionalities from application that could be there while development of application. 2. Code Review. 3. Application Testing | [1.1] Hidden functionalities: Not applicable [2.1] Code review: process followed. [3.1] Application testing : Done | [1.1.1] Hidden functionalities: Not applicable [2.1.1] Code review: process followed. [3.1.1] Application testing : Done | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T05 | Unauthorized modifications to Knee Planning Application | V31 | Insecure Storage of Sensitive Information | A13 | Translation data | Potential system integrity loss. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The keys, password and token should be stored in encrypted form. Few of the sensitive data like keys should be stored using windows credential manager. | [1.1] Keys and tokens are kept in application binary, not stored on drive or configs. | [1.1.1] PR-PY-SYK-CC-008 Hard Coded Sensitive Information [1.1.2] PR-PY-SYK-CC-010 Use of Secure Crypto | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T05 | Unauthorized modifications to Knee Planning Application | V37 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging | A11 | Log Files | Potential system integrity loss. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. All log files should be stored encrypted . 2. The log files should not contain senstive user information and license information. | [1.1][2.1] Audit logs are encrypted SRS: 2.17.10, 2.17.4 SADD: DS13 SDD: ML11 | [1.1.1][2.1.1] PR-PY-SYK-CC-002 Proper Crash Logs are Maintained Locally in Encrypted Form [1.1.2][2.1.2] PR-PY-SYK-CC-010 Use of Secure Crypto | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T06 | Unauthorized access to system administration | V01 | Ineffective management of user credentials | A01 | User Laptop | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The keys, password and token should be stored in encrypted form.Few of the sensitive data like keys should be stored using windows credential manager. | [1.1] Keys and tokens are kept in application binary, not stored on drive or configs. | [1.1.1] PR-PY-SYK-CC-008 Hard Coded Sensitive Information [1.1.2] PR - PY-SYK-CC-010 Use of Secure Crypto | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T06 | Unauthorized access to system administration | V02 | Improper Authentication | A01 | User Laptop | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Moderate | 0.5 | 4.1 | MEDIUM | 1. Proper Access control should be implimented in application directory so that only authorised users can access application and its files. | [1.1] Handled through user id linked license, windows user directory. SRS: 2.17.13, 2.17.4 SADD: DS02, DS04, ES05, ES06 SDD: ML09 | [1.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management | Low | Low | High | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | 5.4 | MEDIUM
| T06 | Unauthorized access to system administration | V03 | Improper Input Validation | A01 | User Laptop | Application behaves in an unpredictable manner. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | 1. The application should accept inputs from application UI interface only. 2. The application should verify input type before accepting it. Example, Date type of input should be in a specific format and should not conatain any other symbol or any alphabetic letter inside it. | [1.1] Handled in application input widgets. [2.1] SRS:2.5.4 SADD: EL06, EL09, ML02, UI02 SDD: C115, C202, C261, C262 | [1.1.1] PR - PY-SYK-CC-009 Command Line Arguments [2.1.1] PR - PY-SYK-CC-005 Input Validation Check | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | 3.7 | LOW
| T06 | Unauthorized access to system administration | V25 | Unprotected public network connections | A01 | User Laptop | Application behaves in an unpredictable manner. |  | Low | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. Firewall 2. The internet communication of Thor application should be secured using SSL or other secure protocols. 3. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. 4. The License should be encrypted for communication over internet. | [1.1] SOM- 13.3 Firewall Implementation [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable as no internet connection is getting used. [4.1] Not applicable as no internet connection is getting used. | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] Not applicable as no internet connection is getting used. [4.1.1] Not applicable as no internet connection is getting used. | Low | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.52545 | 3.3734 | 3.9 | 3.5 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T07 | Unauthorized access to Treatment plan | V02 | Improper Authentication | A04 | THOR Knee Planning Application | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Moderate | 0.5 | 5 | MEDIUM | 1. Proper Access control should be implimented in application directory so that only authorised users can access application and its files. 2. The license should be verified everytime the application is started. 3. The license time period and authority should not get manipulated at client side by manupulating registery data or saved license data. 4. The license should be encrypted . 5. The treatment plan directory should have proper permissions only to authorized user . | [1.1] SOM-6.3 Audit Log Protection [1.2] SOM-7.2 Privilege and Access (AUTH-2, 3, 4, 5, SGUD-3, 3.1 [2.1] SOM-13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [3.1] SOM-16.1 Password/ID Assignments (PAUT-1) [4.1] SOM-13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [4.2] Handled through user id linked license, windows user directory. License data is encrypted. SRS: 2.17.13, 2.17.4 SADD: DS02, DS04, ES05, ES06 SDD: ML09, ES04, DS10, C277 [5.1] SOM-7.2 Privilege and Access (AUTH-2, 3, 4, 5, SGUD-3, 3.1) | [1.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management [2.1.1] PY-SYK-CC-014 Licence Check [3.1.1] PR-PY-SYK-CC-013 Secure Configuration File [4.1.1] PR-PY-SYK-CC-010 Use of Secure Crypto [5.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management [5.1.2] PY-SYK-CC-013 Secure Configuration File | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | 5 | MEDIUM
| T07 | Unauthorized access to Treatment plan | V01 | Ineffective management of user credentials | A04 | THOR Knee Planning Application | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Low | 0.2 | 4.8 | MEDIUM | 1. The user credentials should not be stored unencrypted. 2. The user credentials /license should not be stored in path where treatment plan data is getting stored. | [1.1] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) [2.1] SDD: DS10, DS07 | [1.1.1] PR-PY-SYK-CC-001 User Authentication [2.1.1] PR-PY-SYK-CC-001 User Authentication | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | 4.8 | MEDIUM
| T07 | Unauthorized access to Treatment plan | V31 | Insecure Storage of Sensitive Information | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Low | 0.2 | 4.8 | MEDIUM | 1. The Treatment plan should be stored encrypted. 2. The treatment plan directory should only be accessible to system user using the application. | [1.1] - SOM-4 MANAGEMENT OF PII and PHI (MPII-1) [2.1] - SOM-4 MANAGEMENT OF PII and PHI (MPII-1) Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 | [1.1.1] PR-PY-SYK-CC-012 Secure PHI Data [2.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | 4.8 | MEDIUM
| T07 | Unauthorized access to Treatment plan | V33 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | High | Low | Required | Unchanged | 0.27806 | 0.7323 | 4.7014 | 5 | Low | 0.2 | 4.8 | MEDIUM | 1. The application should use secure cryptographic algorithms to encrypt user data files. | [1.1] - SOM-6.4 Log Export, Use, and Notification (AUDT-5, 6, 8) [1.2]Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 D005010050 INSTALLATION_MANUAL_AND_TROUBLESHOOT | [1.1.1] PR-PY-SYK-CC-010 Use of Secure Crypto | Low | Low | Low | Physical | High | Low | Required | Unchanged | 0.27806 | 0.52545 | 3.3734 | 3.7 | 3.5 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T07 | Unauthorized access to Treatment plan | V34 | Exposure of Private Information | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Low | 0.2 | 4.8 | MEDIUM | 1. The application should use secure cryptographic algorithms to encrypt user data files. 2. No sensitive user data should be stored in windows registry. | [1.1] - SOM-4 MANAGEMENT OF PII and PHI (MPII-1) [1.2] Data encryption handler SRS: 2.17.3 SADD:UI15, ES05, ES06, DS02, DS04 SDD: C276, C277 [2.1] User data not stored in windows registry. Access management - Windows user, | [1.1.1] PR - PY-SYK-CC-010 Use of Secure Crypto [2.1.1] PR - PY-SYK-CC-011 Sensitive Data in Registry | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | 4.8 | MEDIUM
| T07 | Unauthorized access to Treatment plan | V35 | Sensitive Data Storage in Improperly Locked Memory | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | High | Low | Required | Unchanged | 0.27806 | 0.7323 | 4.7014 | 5 | Low | 0.2 | 4.8 | MEDIUM | 1. Identify data that needs to be protected from swapping and choose platform-appropriate protection mechanisms. 2. Check return values of method calls and method API calls to ensure locking operations are successful. | [1.1][2.1]Windows 10 is minimum OS version | [1.1.1][2.1.1] Windows 10 is minimum OS version | Low | Low | Low | Physical | High | Low | Required | Unchanged | 0.27806 | 0.52545 | 3.3734 | 3.7 | 3.5 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T07 | Unauthorized access to Treatment plan | V31 | Insecure Storage of Sensitive Information | A02 | USB Drive | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Low | 0.2 | 4.8 | MEDIUM | 1. The Treatment plan should be stored encrypted. | [1.1] SOM-4 MANAGEMENT OF PII and PHI (MPII-1) Encryption of treatment plan file : TBD | [1.1.1] PR-PY-SYK-CC-010 Use of Secure Crypto | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | 4.8 | MEDIUM
| T07 | Unauthorized access to Treatment plan | V31 | Insecure Storage of Sensitive Information | A08 | Network Location | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Low | 0.2 | 4.8 | MEDIUM | 1. The Treatment plan should be stored encrypted. 2. The treatment plan directory should only be accessible to system user using the application. | [1.1] SOM- 12 - HEALTH DATA INTEGRITY AND AUTHENTICITY (IGAU-1, 2) [2.1]Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 | [1.1.1] PR- PY-SYK-CC-012 Secure PHI Data [1.1.2] PR-PY-SYK-CC-010 Use of Secure Crypto [1.1.3] SOM- 12 -HEALTH DATA INTEGRITY AND AUTHENTICITY (IGAU-1, 2) [2.1.1] PY-SYK-CC-006 Secured Application Path and Proper Access Management | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | 4.8 | MEDIUM
| T08 | Unauthorized access to PHI | V31 | Insecure Storage of Sensitive Information | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | None | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.6568 | 4.2167 | 4.8 | Low | 0.2 | 4.4 | MEDIUM | 1. The PHI data should be stored encrypted. 2. The PHI data directory should only be accessible to system user using the application. 3. The PHI data should not be stored in local SQLIte databases in plain text. 4. The PHI data should not be stored in Log files. | [1.1]PHI data stored in audit log file and case list is encrypted. [2.1]Directory acces is managed using windows user credential. [3.1]PHI data is not stored in local DB. [4.1]PHI data stored in audit log file and case list is encrypted. SRS: 2.17.10, 2.17.13, 2.17.4 SADD: DS13, DS02, DS04, ES05, ES06 SDD:ML11, C276, C277 | [1.1.1]PR - PY-SYK-CC-012 Secure PHI Data [2.1.1] PR - PY-SYK-CC-006 Secured Application Path and Proper Access Management [3.1.1] PR - PY-SYK-CC-012 Secure PHI Data [4.1.1] PR - PY-SYK-CC-012 Secure PHI Data | High | Low | None | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.6568 | 4.2167 | 4.8 | 4.4 | MEDIUM
| T08 | Unauthorized access to PHI | V33 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.6568 | 4.2167 | 4.5 | Low | 0.2 | 4.3 | MEDIUM | 1. The application should use secure cryptographic algorithms to encrypt PHI data/ files. | [1.1]Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 D005010050 INSTALLATION_MANUAL_AND_TROUBLESHOOT | [1.1.1] PR - PY-SYK-CC-012 Secure PHI Data | Low | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.3916 | 2.5141 | 2.8 | 2.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T08 | Unauthorized access to PHI | V25 | Unprotected public network connections | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue |  | High | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.6568 | 4.2167 | 4.5 | Low | 0.2 | 4.3 | MEDIUM | 1. Firewall 2. The internet communication of Thor application should be secured using SSL or other secure protocols. 3. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. | [1.1] SOM- 13.3 Firewall Implementation [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable as no internet connection is getting used. | [1.1.1] SOM- 13.3 Firewall Implementation. [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] Not applicable as no internet connection is getting used. | Low | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.3916 | 2.5141 | 2.8 | 2.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T08 | Unauthorized access to PHI | V35 | Sensitive Data Storage in Improperly Locked Memory | A05 | Patient Personal and Health data | Financial penalties, loss of Stryker Reputation, potential loss of revenue | n/a | High | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.6568 | 4.2167 | 4.5 | Low | 0.2 | 4.3 | MEDIUM | 1. Identify data that needs to be protected from swapping and choose platform-appropriate protection mechanisms. 2. Check return values to ensure locking operations are successful. | [1.1] Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 [1.3] D005010050 INSTALLATION_MANUAL_AND_TROUBLESHOOT [2.1] Not applicable for Windows 10 | [1.1.1][2.1.1] Not applicable for Windows 10 | Low | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.3916 | 2.5141 | 2.8 | 2.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T08 | Unauthorized access to PHI | V35 | Sensitive Data Storage in Improperly Locked Memory | A02 | USB Drive | Potential loss of PHI |  | High | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.6568 | 4.2167 | 4.5 | Moderate | 0.5 | 4.4 | MEDIUM | 1. Identify data that needs to be protected from swapping and choose platform-appropriate protection mechanisms. 2. Check return values of method calls and method API calls to ensure locking operations are successful. | [1.1][2.1]TBD: Not applicable for Windows 10 | [1.1.1][2.1.1] Not applicable for Windows 10 | Low | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.3916 | 2.5141 | 2.8 | 2.7 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T09 | Unauthorized users may steal credentials/device's keys/certificates | V31 | Insecure Storage of Sensitive Information | A15 | Security keys, tokens, certificates | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | Low | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.3916 | 2.5141 | 2.8 | Low | 0.2 | 2.6 | LOW | 1. The keys, certificates , token etc. should stored using windows credential manager. 2. The certificates and token should be stored encrypted . | [1.1][2.1]Keys and tokens are kept in application binary, not stored on drive or configs. [1.2][2.2] Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 | [1.1.1] PR- PY-SYK-CC-001 User Authentication [2.1.] PR-PY-SYK-CC-010 Use of Secure Crypto | Low | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.3916 | 2.5141 | 2.8 | 2.6 | LOW
| T09 | Unauthorized users may steal credentials/device's keys/certificates | V33 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength | A15 | Security keys, tokens, certificates | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | Low | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.3916 | 2.5141 | 2.8 | Low | 0.2 | 2.6 | LOW | 1. The application should use secure cryptographic algorithms to encrypt security keys, tokens and certificates. | [1.1]Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 | [1.1.1]PR - PY-SYK-CC-010 Use of Secure Crypto | Low | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.3916 | 2.5141 | 2.8 | 2.6 | LOW
| T09 | Unauthorized users may steal credentials/device's keys/certificates | V04 | Unprotected Storage of Credential | A15 | Security keys, tokens, certificates | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | Low | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.3916 | 2.5141 | 2.8 | Low | 0.2 | 2.6 | LOW | 1. The keys, certificates , token etc. should stored using windows credential manager. 2. The certificates and token should be stored encrypted . | [1.1][2.1]Keys and tokens are kept in application binary, not stored on drive or configs. Access management - Windows user, Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 | [1.1.1] PR- PY-SYK-CC-001 User Authentication [2.1.1] PR-PY-SYK-CC-010 Use of Secure Crypto | Low | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.3916 | 2.5141 | 2.8 | 2.6 | LOW
| T09 | Unauthorized users may steal credentials/device's keys/certificates | V31 | Insecure Storage of Sensitive Information | A14 | User License data | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | High | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.6568 | 4.2167 | 4.5 | Moderate | 0.5 | 4.4 | MEDIUM | 1. The user encrypted license should be stored using windows credential manager. 2. The License data should be encrypted and should not be modified. | [1.1] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) [1.2] Access management - Windows user, AES Data encryption handler SRS: 2.17.3 SADD: 2.2.8.1, ES05, ES06, DS02, DS04 SDD: C276, C277 [2.1] SOM-3.5 System Security Context and Intended Environment (SGUD-4) [2.2] SAD: UI15, ML09, ES05, ES06, DS02, DS04 | [1.1.1] PR-PY-SYK-CC-001 User Authentication | High | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.6568 | 4.2167 | 4.5 | 4.4 | MEDIUM
| T09 | Unauthorized users may steal credentials/device's keys/certificates | V33 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength | A14 | User License data | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | High | Low | None | Physical | High | High | Required | Unchanged | 0.12109 | 0.6568 | 4.2167 | 4.4 | Moderate | 0.5 | 4.3 | MEDIUM | 1. The application should use secure cryptographic algorithms to encrypt user Licence data. | [1.1] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) [1.2]Used AES 256 Data encryption Algorithm SRS: 2.17.3 SADD: 2.2.8.5.1, 2.2.8.5.2. UI15 | [1.1.1] PR-PY-SYK-CC-010 Use of Secure Crypto | Low | Low | None | Physical | High | High | Required | Unchanged | 0.12109 | 0.3916 | 2.5141 | 2.7 | 2.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T09 | Unauthorized users may steal credentials/device's keys/certificates | V31 | Insecure Storage of Sensitive Information | A03 | User Password / Credentials | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | High | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.6568 | 4.2167 | 4.5 | Low | 0.2 | 4.3 | MEDIUM | 1. The Password credentials should be encrypted and should be stored using windows credential manager. | [1.1]Not Applicable. Using windows credentials to authenticate | [1.1.1] PR-PY-SYK-CC-001 User Authentication | High | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.6568 | 4.2167 | 4.5 | 4.3 | MEDIUM
| T10 | Brute force login attempts/password guessing attacks | V02 | Improper Authentication | A04 | THOR Knee Planning Application | Potential loss of PHI |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. The password to login should be complex and contain combination of Capital and small letters , symbols/ special chracters and numbers. 2. There should be server side rate limiting implimented on password field so that after certain wrong attempts user should not be able to try again. 3. There should be server side captcha verification when login to restrict brute force attack. 4. The application binary should not take password as a run time argument. | [1.1] SOM- 16. PERSON AUTHENTICATION (PAUT) [1.2] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) [2.1] Not Applicable as there is no server [2.2] SOM-3.5 System Security Context and Intended Environment (SGUD-4) [3.1] Not Applicable as there is no server [3.2] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) [4.1] | [1.1.1] PR-PY-SYK-CC-001 User Authentication [1.1.2] SOM- 16. PERSON AUTHENTICATION (PAUT) [1.2.1] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) [2.1.1] NA [3.1.1] NA [4.1.1]PR-PY-SYK-CC-009 Command Line Arguments | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | 3.5 | LOW
| T10 | Brute force login attempts/password guessing attacks | V02 | Improper Authentication | A06 | Laptop Operating System | Potential loss of PHI |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. Firewall 2. The System should be locked for some time after multiple wrong password attempts. 3. Application should not be having same password as system user password. | [1.1] SOM - 13.3 Firewall Implementation [2.1] SOM - 16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) [3.1] SOM - 16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) | [1.1.1] SOM - 13.3 Firewall Implementation [2.1.1] SOM - 16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) [3.1.1] SOM - 16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | 3.5 | LOW
| T11 | Compromised Config Files, causing the system to become inaccurate | V31 | Insecure Storage of Sensitive Information | A09 | Application Settings | Potential system integrity loss. |  | Low | Low | Low | Physical | Low | High | Required | Unchanged | 0.21191 | 0.52545 | 3.3734 | 3.6 | Low | 0.2 | 3.5 | LOW | 1. The application setting file should not have permission to get edited by any system user. 2. The application setting file should not contain any credentials , keys, tokens and license data. | [1.1]Config file is kept in user directory and not accesible to other user. [1.2] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) [2.1]SRS: 2.17.2 SADD- DS02, DS04 [2.2] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) | [1.1.1][2.1.1] PR - PY-SYK-CC-006 Secured Application Path and Proper Access Management [1.2.1][2.2.1] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) | Low | Low | Low | Physical | Low | High | Required | Unchanged | 0.21191 | 0.52545 | 3.3734 | 3.6 | 3.5 | LOW
| T11 | Compromised Config Files, causing the system to become inaccurate | V33 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength | A09 | Application Settings | Potential system integrity loss. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The application should use secure cryptographic algorithms to encrypt user any sensitive information in application settings. | [1.1]Application setting does not contain any PHI or PPI. [1.2] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) | [1.1.1] PR - PY-SYK-CC-012 Secure PHI Data [1.2.1] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T11 | Compromised Config Files, causing the system to become inaccurate | V31 | Insecure Storage of Sensitive Information | A10 | User Settings | Potential system integrity loss. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The User setting file should not have permission to get edited by any system user. 2. The user setting file should not contain any credentials , keys, tokens and license data. 3. The user settings should not contains PHI data. | [1.1]Config file is kept in user directory and not accesible to other user. [2.1]SRS: 2.17.2, 2.17.3 SADD- DS02, DS04, ES05, ES06 [3.1] SOM- 3.5 System Security Context and Intended Environment (SGUD-4) | [1.1.1] PR - PY-SYK-CC-013 Secure Configuration File [2.1.1] PR - PY-SYK-CC-006 Secured Application Path and Proper Access Management [3.1.1][3.1.2] SOM- 3.5 System Security Context and Intended Environment (SGUD-4) | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T11 | Compromised Config Files, causing the system to become inaccurate | V36 | Lack of fault reports recorded in administrator and operator logs | A11 | Log Files | Potential system integrity loss. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The logs should be properly maintained for administrator and operator operations. 2. The logs should not disclose any sensitive data. | [1.1] SOM- 7.1 Access Prevention (AUTH-1) [1.2] SOM- 16. PERSON AUTHENTICATION (PAUT) [1.3] SOM- 6. AUDIT CONTROLS (AUDT) [1.4] SOM- 6.3 Audit Log Protection (AUDT-7) [2.1]Debug logs, Audit logs are encrypted. SRS: 2.17.10 SADD: DS13 SDD:ML11 | [1.1.1] SOM- 7.1 Access Prevention (AUTH-1) [1.2.1] SOM- 16. PERSON AUTHENTICATION (PAUT) [1.3.1] SOM- 6. AUDIT CONTROLS (AUDT) [1.4.1] SOM- 6.3 Audit Log Protection (AUDT-7) [1.4.2][2.1.1]PR-PY-SYK-CC-002 Proper Crash Logs are Maintained Locally in Encrypted Form | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T11 | Compromised Config Files, causing the system to become inaccurate | V33 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength | A10 | User Settings | Potential system integrity loss. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The application should use secure cryptographic algorithms to encrypt sensitive user information. | [1.1]Access management - Windows user, AES 256 Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 [1.2] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) | [1.1.1] PR - PY-SYK-CC-010 Use of Secure Crypto [1.2.1] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T12 | Theft of Stryker's IP | V16 | Embed Virus into Binary or Embedded Malicious Code | A04 | THOR Knee Planning Application | Unauthorized agent acquires and modifies Stryker's source. Could offer Stryker's IP to competitor to accelerate time to market of competing product or for directed malware attack. |  | Low | Low | Low | Physical | Low | High | Required | Unchanged | 0.21191 | 0.52545 | 3.3734 | 3.6 | Low | 0.2 | 3.5 | LOW | 1. Virus scan 2. Application should verify the checksum of all the DLL files and executable inside application folder before starting main thread. 3. Anti hooking techniques should be implimented in application. | [1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1] [3.1] | [1.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [3.1.1] PR-PY-SYK-CC-007 Dynamic Instrumentation | Low | Low | Low | Physical | Low | High | Required | Unchanged | 0.21191 | 0.52545 | 3.3734 | 3.6 | 3.5 | LOW
| T12 | Theft of Stryker's IP | V12 | UnProtected Application Binaries | A04 | THOR Knee Planning Application | Unauthorized agent acquires and modifies Stryker's source. Could offer Stryker's IP to competitor to accelerate time to market of competing product or for directed malware attack. |  | Low | Low | Low | Local | High | High | Required | Unchanged | 0.333 | 0.52545 | 3.3734 | 3.8 | Low | 0.2 | 3.5 | LOW | 1. The application binaries should be obfuscated and should not take any runtime inputs from user console. 2. The permission to binaries should be appropriate so that only authorised system user should only have access. | [1.1] SOM- 3.5 System Security Context and Intended Environment (SGUD-4) [2.1] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) | [1.1.1] PR-PY-SYK-CC-004 Obfuscation of Code [2.1.1] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) | Low | Low | Low | Local | High | High | Required | Unchanged | 0.333 | 0.52545 | 3.3734 | 3.8 | 3.5 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T12 | Theft of Stryker's IP | V17 | Hidden Functionality | A04 | THOR Knee Planning Application | Unauthorized agent acquires and modifies Stryker's source. Could offer Stryker's IP to competitor to accelerate time to market of competing product or for directed malware attack. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. Application should not have direct access to hidden functionalities. It is recommended to delete all unused funtionalities from application that could be there while development of appplication. 2. Code review 3. Application Testing | [2.1] Code Review Done [3.1] Application Testing Done | [2.1.1] Code Review Done [3.1.1] Application Testing done | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T12 | Theft of Stryker's IP | V15 | 3rd Party Component Dependency | A04 | THOR Knee Planning Application | Unauthorized agent acquires and modifies Stryker's source code. Could offer Stryker's IP to competitor to accelerate time to market of competing product. |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The application should be updated using 3rd party modules and dependencies to latest versions. 2. Application upgrade should verify if all required dependencies are updated in system. 3. Application upgrade should detect any upgrade or change in dependecies on user system and show dependency issue message to application user. | [1.1] SOM-8.1 Secure Servicing and Security Upgrades Overview (CSUP-1) [1.2] SOM-28.1 List of 3rd party components [2.1] SOM-8.1 Secure Servicing and Security Upgrades Overview (CSUP-1) [3.1] SOM-8.1 Secure Servicing and Security Upgrades Overview (CSUP-1) | [1.1.1] PR-PY-SYK-CC-003 Application Dependency Check [2.1.1] PR-PY-SYK-CC-003 Application Dependency Check [3.1.1] PR-PY-SYK-CC-003 Application Dependency Check | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | 3.5 | LOW
| T12 | Theft of Stryker's IP | V38 | Information Exposure Through Log Files | A04 | THOR Knee Planning Application | Unauthorized agent acquires and modifies Stryker's source. Could offer Stryker's IP to competitor to accelerate time to market of competing product or for directed malware attack. |  | Low | Low | Low | Local | High | High | Required | Unchanged | 0.333 | 0.52545 | 3.3734 | 3.8 | Low | 0.2 | 3.5 | LOW | 1. The log files should not contain any sensitive data like tokens, stack trace, license data, user data, passwords etc. | [1.1] SOM- 6.3 Audit Log Protection (AUDT-7) [1.2] Debug logs, Audit logs. Audit logs are encrypted. SRS: 2.17.10 SADD: DS13 SDD:ML11 | [1.1.1]PR-PY-SYK-CC-002 Proper Crash Logs are Maintained Locally in Encrypted Form [1.1.2] SOM- 6.3 Audit Log Protection (AUDT-7) [1.2.1] PR-PY-SYK-CC-002 Proper Crash Logs are Maintained Locally in Encrypted Form | Low | Low | Low | Local | High | High | Required | Unchanged | 0.333 | 0.52545 | 3.3734 | 3.8 | 3.5 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T13 | Distributed Denial of Service ('DDoS') attack | V26 | Lack of monitoring mechanisms | A04 | THOR Knee Planning Application | Access to Thor knee planning application impacted due to system being down or functional degradation |  | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | 1. Firewall 2. Virus Scan | [1.1] SOM- 13.3 Firewall Implementation [2.1] SOM-6.4 Log Export, Use, and Notification (AUDT-5, 6, 8) [2.2] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [2.3] SOM-13. MALWARE DETECTION/PROTECTION (MLDP-1) | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] SOM- 8.5 Anti-Malware Software Updates (CSUP-4) [2.2.1][2.3.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | 3.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
| T13 | Distributed Denial of Service ('DDoS') attack | V26 | Lack of monitoring mechanisms | A01 | User Laptop | Access to Thor knee planning application impacted due to system being down or functional degradation |  | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | 1. Firewall 2. The proper dependency check and system specification should be checked whenever application starts. If there is anything not suitable for Thor , the application shsould show error message and should not load. | [1.1] SOM- 13.3 Firewall Implementation System health check. [2.1] SOM-3.5 System Security Context and Intended Environment (SGUD-4) SRS: 2.2.7 | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] PR-PY-SYK-CC-003 Application Dependency Check | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | 3.6 | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
|  | #N/A |  | #N/A |  | #N/A | [10 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[6 empty] | Document references: | [18 empty] | SRS - CT knee SRS: D005010021
SAD - THOR: CT Knee Planning SOFTWARE_ARCHITECTURAL_DESIGN: D005010022
SDD - THOR: CT Knee Planning SOFTWARE_DETAILED_DESIGN: D005
PR- Pentest report
Note on AA6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav for 4th point "Application should be updated with the latest version of dependencies" add a refrence. Reply: Updated with 4.1
Note on AA8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: # Gaurav add the refrences properly. we have added the verififcation. Just go through it and create a cross reference. These points are technical so kindly see to it. Reply: Updated with correct section
Note on AB8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav give a number to this point
Note on AA11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav fill the reference for 4th point
Note on AA13: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav for 4th point "Application should be updated with the latest version of dependencies" add a refrence.
Note on AA15: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav for 4th point "Application should be updated with the latest version of dependencies" add a refrence.
Note on AA20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: # Gaurav add reference for 3.1 Reply: Updated
Note on AA21: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav Please add reference to point 2.1 and 4.1
Note on AA30: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav add reference for point 2 Reply: Updated
Note on AA33: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav please add reference here point no. 2 Reply: Updated
Note on AA40: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav there are no reference for point 1 Reply: Updated
Note on AA45: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav please add proper points to document references Reply: Updated
Note on AA46: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav Please provide reference for the point 1. The one given here is taking about the access management but it is related to license data. So please check if its applicable here Reply: updated
Note on AA47: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav please provide reference for [1.1] if applicable. Reply: updated
Note on AA48: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav please add a reference for point 4
Note on AA55: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav kindly fill refrence for 2.1 and 3.1
Note on AA57: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav please provide reference for point 1

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware | ` | TSA-3 - Skript Kiddies | Yes | n/a
T02 | Deliver directed malware | Thread source delivers malware on a removable media which was designed to exploit a known vulnerability of the Software. Directed attack on the Application using knowledge about the Application. | TSA-2 Organization | Yes | n/a
T03 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies | Yes | n/a
T04 | Intercept network communication | Adversary uses commercial or free software to manipulate network communication (man-in-the middle attack) | TSA-2 Organization | No | Application donot share any data over the network
T05 | Unauthorized modifications to Knee Planning Application | Adversary gains unauthorized access to the system and removes or modifies application binaries. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T06 | Unauthorized access to system administration | Adversary gains unauthorized access to the system and removes or modifies application configuration or binaries. | TSA-3 - Skript Kiddies | Yes
T07 | Unauthorized access to Treatment plan | Adversary gains unauthorized access to the treatment design plan of the patient. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes
T08 | Unauthorized access to PHI | Adversary gains unauthorized access to the system and steals ePHI information | TSA-3 - Skript Kiddies- | Yes
T09 | Unauthorized users may steal credentials/device's keys/certificates | Adversary gains unauthorized access to the encryption keys and tokens | TSA-3 - Skript Kiddies | Yes
T10 | Brute force login attempts/password guessing attacks | Adversary gains unauthorized access to the system | TSA-3 - Skript Kiddies | Yes
T11 | Compromised Config Files, causing the system to become inaccurate | Adversary gains unauthorized access to the system and make modifications to the software configuration files | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes
T12 | Theft of Stryker's IP | Unauthorized agent acquires and modifies Stryker's source code. Could offer Stryker's IP to competitor to accelerate time to market of competing product. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes
T13 | Distributed Denial of Service ('DDoS') attack | Access to Knee Planning application impacted due to system being down or functional degradation | TSA-3 - Skript Kiddies | Yes
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V11 | Unpatched COTS operating system | A06 | Laptop Operating System | Access to Thor knee Planning application impacted due to system being down or functional degradation |  | MEDIUM | 1.Application Whitelisting 2.Firewall 3.Virus Scan 4.User authentication 5.Proper crash logs of application should be maintained | LOW
2 | T01 | Deliver undirected malware | V13 | Ineffective patch management | A06 | Laptop Operating System | Access to Thor knee Planning application impacted due to system being down or functional degradation |  | MEDIUM | 1. Virus Scan 2. Firewall 3. Windows security update install 4. Application should be updated with the latest version of dependencies. | LOW
4 | T01 | Deliver undirected malware | V14 | Lack of software vulnerability management | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | MEDIUM | 1. Update drivers 2. Virus Scan 3. Input validation and proper access management in application 4. Application should not be allowed to get installed in shared directory/path. | LOW | Post the implementation of Security Control measures the security risk level is low, so residual risk is acceptable.
5 | T01 | Deliver undirected malware | V15 | 3rd Party Component Dependency | A04 | THOR Knee Planning Application | Application behaves in an unpredictable manner. |  | MEDIUM | 1. The 3rd party modules and dependencies should be updated with latest versions. 2. Application upgrade should verify if all required dependencies are updated in system. 3. Application upgrade should detect any upgrade or change in dependecies on user system and show dependency issue message to application user. | LOW
0 | T01 | Deliver undirected malware | V25 | Unprotected public network connections | A06 | Laptop Operating System | Access to Thor knee Planning application impacted due to system being down or functional degradation |  | MEDIUM | 1. Firewall 2. Virus Scan 3. The internet communication of Thor application should be secured using SSL or other secure protocols. 4. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. | MEDIUM
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls