## Header (Optional)
Product Security Risk Table
Ortho Slice 3D Knee Navigation Software
Document Control
| Author: | Siva Kumar
| Document No.: | D005020007
| Revision No.: | 01
| Project: | CT Knee Intra Op- SGTC-NPD-005-02
| Project Lead: | Akhil Gupta
Change Records
| Rev. No. | Comment | Date | Author
| 0 | Initial revision | Mar-11-2022 | Siva Kumar
| 1 | Update for DR3-DR4 | 8 July 2022 | Siva Kumar
Note on A4: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: General Comments - If we are planning to have this sheet into OnePLM and final release document will be PDF, then formating must be done to according to correct print configguration so that data will not mess up.
Note on B7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Please upate name
Note on B19: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Stryker Flex Cart( Nav3i/Q Cart/Ortho Q ) Navigation System
Scope: |  | CT Knee Intra Op- SGTC-NPD-005-02
Date: |  | <2022-07-14>
Conducted by: |  | <Siva Kumar / Security / L&T> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description | Remarks (architecture)
A01 | Physical Asset | Nav3,3i cart/ System running with windows 8.1 | Contains software module which controls the overall functionality of Knee System i.e. intra-operative planning and guidance.
A02 | Information Asset | Admin Password / Credentials / System Configuration / Certificates | Admin Passwords/Credentials allow modifying config files, accessing ePHI, software binaries, computer configuration files, log files, and computer resources as well as installing applications, security certificates, OS updates.
A03 | Information Asset | User Password / Credentials / System Configuration | User Passwords/Credentials provide access to ePHI, Software Binaries, read access to config files, log files.
A05 | Information asset | Patient health information at rest (PHI) | Data stored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A07 | Information asset | Computer/OS network identification | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A08 | Information Asset | Intellectual Property (Stryker IP) | Algorithms and other core-elements of interest to competitors.
A09 | Information Asset | Navigation Accuracy | Extended arm of Nav3i cart contains multiple sensors & navigation cameras. The data collected & provided to the Intra-OP application has to be accurate
A10 | Information Asset | SSH credentials for system admin remote access | Only using an 2048bit SSH key allows full access to all system functions. That SSH key is not provided to the customer and can be used by Stryker internals only.
A11 | Information asset | THOR Knee Intra-op Application | Contains Knee intra-op software which controls the overall features for segmentation, landmark planning, implant size calculation etc.
A12 | Information asset | Patient related details | Contains patient personal details such as patient identity, DOB, Age/Gender, image data
A13 | Physical Asset | Input device to Intra-OP | Any kind of removable device which is connected to knee intra-op from knee planning application
A14 | Information Asset | Patient Health Information in transit (PHI) | Any kind of input data from removable device to knee intra-op from knee planning application
A15 | Physical Asset | Nav3,3i device configurable information (for eg: bios, etc.,) | Knee Intra-Op running Cart/System having HW resources and Configurable Bios
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user credentials | Yes | n/a
V02 | Ineffective management of admin credentials | Yes | n/a
V03 | Untrained/Malicious User | Yes | n/a
V04 | Unauthorized modification to system bios | Yes | n/a
V05 | No proper validation for the removable devices | Yes | n/a
V06 | Lack of multi-factor authentication | Yes | n/a
V07 | Application allowing unlimited login attempts | Yes | n/a
V08 | User account creation without providing password | Yes | n/a
SBOM/Software Management
V11 | Unpatched OS | Yes | n/a
V12 | Potentially outdated SW-Library SOUPs | Yes | n/a
V13 | Developed code not validated using Static analysis | Yes | n/a
V14 | UnSafely Compiled Application software | Yes | n/a
V15 | Lack of periodic software vulnerability management | Yes | n/a
V16 | 3rd Party Component/SOUP Risk/Dependency not validated | Yes | n/a
V17 | Legacy Systems (outdated HW/SW) being used | Yes | n/a
Access points
V21 | Unprotected network port | Yes | n/a
V22 | Unprotected external USB Port | Yes | n/a
V24 | Any unprotected hardware | Yes | n/a
V25 | Non-Tracing of unused external Interfaces (ports etc., ) | Yes | n/a
V26 | Exposure of sensitive product information | Yes | n/a
Application Data Management
V31 | Unencrypted ePHI at rest | Yes | n/a
V32 | Unencrypted ePHI in transit (as output, input) | Yes | n/a
V33 | Insecure Storage of Sensitive Information by application | Yes | n/a
V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | Yes | n/a
V35 | Exposure of Private Information (Health/Personal data) | Yes | n/a
V36 | Malicious activities not being recorded in administrator/user logs | Yes | n/a
V37 | Unauthorized Audit Log Manipulation, Log Injection-Tampering-Forging | Yes | n/a
V38 | Information Exposure through Log Files | Yes | n/a
V39 | Lack of input sanitization in knee intra-op application | Yes | n/a
V40 | Application not having periodic data backup & restore plan | Yes | n/a
InSecure Configurations of Resources
V41 | Lack of validation for the editable configuration controls | Yes | n/a
V42 | Software Download from malicious source | Yes | n/a
V43 | Lack of control over the Automated Software Update | Yes | n/a
V44 | External Control of System or Configuration Setting | Yes | n/a
EOL OS vulnerabilities
V51 | Unsupported OS from vendor | Yes | n/a
V52 | Unavailability of support from 3rd party tools for outdated OS | Yes | n/a
V53 | OS Security tools from vendor not getting updated | Yes | n/a
V54 | Absence of periodic backup plan | Yes | n/a
V55 | Corrupted OS and absence of restore points | Yes | n/a
V56 | Remote exploitation of the system through network | Yes | n/a
V57 | Any malicious service running in the system without being identified | Yes | n/a
V58 | Unhardened OS and network interfaces | Yes | n/a
Linux OS
V61 | Absence of periodic backup plan | No | n/a
V62 | Corrupted OS and absence of restore points | No | n/a
V63 | Remote exploitation of the system from intranet | No | n/a
V64 | Any malicious service running in the system without being identified | No | n/a
V65 | Unhardened OS and network interfaces | No | n/a
V66 | Securtiy packages not properly configured during OS build | No | n/a
V67 | Critical patches not identified and integrated within timeline | No | n/a
Cloud Backend
V71 | Unsecurely transmitted data to cloud | No | Currently not in the scope of Intra-Op application
V72 | Unsecurely encrypting the database | No | Currently not in the scope of Intra-Op application
V73 | Single Factor Authentication to cloud | No | Currently not in the scope of Intra-Op application
Nav3, 3i Cart & APM
V81 | Lack of Mandatory password policy for User(s) | Yes | n/a
V82 | Unrestricted FileSystem Access & Modification | Yes | n/a
V83 | Uniform privilege for all Installables | Yes | n/a
V84 | No Manual AV Scanning/Signature Update Option provided | Yes | n/a
V85 | Unauthorized Modification to Registry & System Settings | Yes | n/a
V86 | Open access for Application data (resources, logs, case data etc., ) | Yes | n/a
V87 | Daemon access by unauthorized resources (ex., RDP, etc., ) | Yes | n/a
V88 | No Notification mentioning outdated AV Definitions | Yes | n/a
V89 | No "on access" Scanning triggered for removable devices | Yes | n/a
V90 | Lack of authentication for externally connected devices (camera, etc.,) | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware (CAPEC-185) | Threat source delivers malware by providing removable media prepared with malware. e.g. Removable media is left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System (or) Knee Intra-Op application. Malware exploits a known vulnerability in order to gain access to restricted data or to manipulate data. Undirected attack on computer systems. | TSA-3 - Skript Kiddies, TSA-1 Individual | Yes | n/a
T02 | Deliver directed malware (CAPEC-185) | Threat source delivers malicious software on a removable media which was designed to exploit a known vulnerability of the Navigation System (or) Knee Intra-Op application. Directed attack on the Navigation System using knowledge of the Navigation System and Knee Intra-Op application. | TSA-2 Organization | Yes | n/a
T03 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies and TSA-1 Individual | Yes | n/a
T04 | Intercept network communication | Adversary manipulates the network environment to intercept network communication to gain ePHI, passwords etc. or to manipulate DICOM data (or) TDM image etc. (man-in-the middle attack) | TSA-2 Organization | No | For DICOM Image (or) TDM data input data is never transferred over network.
T05 | Conduct scavenging of ePHI at rest | Adversary gains unauthorized access to the system(Cart) and steals ePHI information | TSA-1 Individual | Yes | n/a
T06 | Theft of system or hard drives | A thief steals the navigation system and/or the hard drives as loot, which can lead to unavailability of the navigation system and/or leak of Protected Health Information (PHI). | TSA-1 Individual | Yes | n/a
T07 | Data theft via physical media | The malicious or inadvertent copy of ePHI on portable media (USB Sticks, DVD drives) from the navigation system. | TSA-2 Organization | Yes | n/a
T08 | Network-based denial of service (DoS) attack | A threat actor seeks to make a machine or network resource unavailable to its intended users by overloading the network resource , the network (intranet) or target services with malicious data traffic. | TSA-3 - Skript Kiddies TSN-4 Malfunctioning Software / OS TSN-3 Malfunctioning device | Yes | n/a
T09 | Exploit physical access of authorized staff to gain access to organizational facilities. | Adversary follows (“tailgates”) authorized individuals into secure/controlled locations with the goal of gaining access to facilities, circumventing physical security checks. | TSA-1 Individual | No | HDO responsible for physical security controls
T10 | Mis-configuration by user | A user unknowingly or maliciously mis-configures the system so that it malfunctions or no longer functions at all. | TSN-1 Accidental | Yes | n/a
T11 | Remote exploitation | A threat actor uses the network connection of the device(Cart) to target a vulnerability in the operating system and/or software to gain access to the navigation system. This is likely inadvertent, targetting general vulnerability in popular operating systems. After successful exploit, continues as malware threat. | TSA-1 Individual | Yes | n/a
T12 | Improper disposal of hard disks | A user disposes the entire system or the hard disks (for instance after a service issue or at the end of the service life of the product). The hard disks still contain patient data. | TSN-1 Accidental | Yes | n/a
T13 | Man-in-the middle attack / intercept Navigation communication | Adversary manipulates the localization communication such as infrared - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T14 | Conduct scavenging of IP at rest | Adversary gains unauthorized access to the system and steals IP off the system | TSA-1 Individual | No | HDO responsible for physical security controls
T15 | Frequency Jamming or Destruction of Wireless Communication | Adversary manipulates the wireless network environment in order to render the service unavailable. | TSA-3 - Skript Kiddies | No | Ortho DICOM Image (or) TDM data in application will not be impacted due to unavailability of wireless network
T16 | Physical Manipulation of Hardware | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. | TSA-2 Organization or TSA-1 Individual | Yes | n/a
T17 | Gather information using open source discovery of organizational information | Adversary mines publiclly accessible information to gather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. | TSA-3 - Skript Kiddies | Yes | n/a
T18 | Craft counterfeit certificates. | Adversary counterfeits or compromises a certificate authority, so that malware or connections will appear legitimate. | TSA-1 Individual | No | Device (Cart) is a pass through (or) receiver of data in transit from end point devices. The certificate validation is the responsibility of the end points.
T19 | Install persistent and targeted sniffers on organizational information systems and networks. | Adversary places sniffers within internal organizational information systems or networks software designed to (over a continuous period of time) collect (sniff) network traffic. | TSA-2 Organization | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T20 | Insert malicious scanning devices (e.g., wireless sniffers) inside facilities. | Adversary uses postal service or other commercial delivery services to deliver to organizational mailrooms a device that is able to scan wireless communications accessible from within the mailrooms and then wirelessly transmit information back to adversary. | TSA-1 Individual | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T21 | Power Failure at primary facility | Power Failure at primary facility makes facility including the navigation sytem inoperable | TSN-7 Power Failure | Yes | n/a
T22 | Theft of SSH key for root access of Navigation Systems | Adverary uses SSH root key to get unauthorized access to the full system. | TSA-1 Individual | No | SSH root key is stored safely at Stryker facility and never delivered to the HDO / customer.
T23 | Manipulation of tool firmware / memory | Adversary manipulates the tool related to calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | No | For DICOM images/TDM data, no tool is used for calibration.
T24 | Manipulation of navigation camera firmware / memory | Adversary manipulates the navgation camera related calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T25 | Manipulation of EM box firmware / memory | Adversary manipulates the EM box firmware - thereby influencing the navigation accuracy | TSA-1 Individual | No | EM is not used by DICOM image/TDM data handling applications
T26 | Unexpected/Damaged input to CT Knee Intra-op | Input from the Knee planning got modified in transit and unexpected input received by Intra-Op | TSA-1 Individual | Yes | n/a
T27 | Malware affected input device | An unscanned input device can inject a malware into the intra-op device | TSA-1 Individual | Yes | n/a
T28 | Critical data transfer involving the removable device | Write protect should be enable to avoid the transfer of critical data | TSA-1 Individual | Yes | n/a
T29 | Lack of encryption for input data to Intra-op | If input data is not encrypted then data can be easily accessible by anyone | TSA-1 Individual | Yes | n/a
T30 | Lack of authentication for the connected removable device | All the connected devices should not be allowed to provide input. Proper authentication should be carried out | TSA-1 Individual | Yes | n/a
T31 | Absence of data validation | Any input data provided to the Intra-op should be properly validated | TSA-1 Individual | Yes | n/a
T32 | Unsecured API interface | All API request need to have a secure header implementation | TSA-1 Individual | Yes | n/a
T33 | Broken authentication (user name, password…) | BRUTE force/Dictonary attacks are possible | TSA-1 Individual | Yes | n/a
T34 | BRUTE force/Dictonary attacks | BRUTE force/Dictonary attacks are possible if system not locked after certain retries | TSA-1 Individual | Yes | n/a
T35 | Unauthorized modification of application resources | Application resources such as binaries, data, etc.. getting modified | TSA-3 - Skript Kiddies, TSA-1 Individual | Yes | n/a
T36 | Unpatched/Unmanaged SOUP's | Outdated/Unpatched softwares can be easily exploited | TSA-1 Individual | Yes | n/a
T37 | Exploitation of unsecured network elements | Unused ports/other network interfaces has to be identified | TSA-1 Individual | Yes | n/a
T38 | Outdated cryptographic components & techniques | Keys, certificates, hashing techniques, etc... should be up-to-date with the latest standards | TSA-1 Individual | Yes | n/a
T39 | Unauthorized access/modification of secure data | Audit (log) data is considered as a sensitive data. It should be secured with proper access control | TSA-1 Individual | Yes | n/a
T40 | Risk of using windows 8.1 | Windows 8.1 & prior versions are considered as an EOL OS. Vendor no more provide support for these OS's | TSA-1 Individual | Yes | n/a
T41 | Risk of using Linux | Linux is an open-source OS. Some vendors has linux flavors which still needs to be considered for risk analysis | TSA-1 Individual | No | n/a
T42 | Application failed to restore data after any abnormal activity | Application should adapt the data backup policy. Data has to be encrypted and stored in a secure location at periodic timeline intervals | TSA-1 Individual | Yes | n/a
T43 | Authentication bypass | Application should have a proper authentication mechanism | TSA-1 Individual | Yes | n/a
T44 | Complete file system access | File system access control is necessary to prevent exploits that can result in an attacker assuming full control of a system. | TSA-1 Individual | Yes | n/a
T45 | All kinds of applications/scripts running | All kinds of applications/scripts should not run in the system. Running of malicious applications/scripts should be removed/uninstalled | TSA-1 Individual | Yes | n/a
T46 | Unidentified virus existence | Unidentified virus can make the system more vulnerable and easier to attack. | TSA-1 Individual | Yes | n/a
T47 | All time system settings modification (during running & bootup) | Modification of registry and system settings (during running & bootup) | TSA-1 Individual | Yes | n/a
T48 | Deletion of critical/sensitive data | Loss of critical/sensitive data (such as resources, logs, case data, etc.,) | TSA-1 Individual | Yes | n/a
T49 | Illegal OS services access | Several resources can be accessed & altered | TSA-1 Individual | Yes | n/a
T50 | Injection of latest malware/virus | If system if affected with latest malware/virus, AV should identify and should intimate the user. | TSA-1 Individual | Yes | n/a
T51 | Deliver direct/undirect malware | Threat source delivers malicious software on a removable media which was designed to exploit a known vulnerability of the Navigation System (or) Knee Intra-Op application. Directed attack on the Navigation System using knowledge of the Navigation System and Knee Intra-Op application. Threat source delivers malware by providing removable media prepared with malware. e.g. Removable media is left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System (or) Knee Intra-Op application. Malware exploits a known vulnerability in order to gain access to restricted data or to manipulate data. Undirected attack on computer systems. | TSA-2 Organization TSA-1 Individual TSA-3 - Skript Kiddies | Yes | n/a
T52 | Unvalidated external connected devices | An unvalidated externally connected device can effect the system. | TSA-1 Individual | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V05 | No proper validation for the removable devices | A01 | Nav3,3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). |  | Low | High | None | Local | Low | Low | Required | Unchanged | 1.3382 | 0.6568 | 4.2167 | 5.6 | Low | 0.2 | 4.5 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
2 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Developed code not validated using Static analysis | A01 | Nav3,3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) due to discovered defect or bug in Stryker software. |  | Low | Low | High | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | Very Low | 0.04 | 4.8 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T02 | Deliver directed malware (CAPEC-185) | V03 | Untrained/Malicious User | A01 | Nav3,3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. |  | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
4 | T02 | Deliver directed malware (CAPEC-185) | v02 | Ineffective management of admin credentials | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs malware in order to gain admin credentials |  | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities |  | Low | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.3916 | 2.5141 | 4.8 | Moderate | 0.5 | 3.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T05 | Conduct scavenging of ePHI at rest | V01 | Ineffective management of user credentials | A05 | Patient health information at rest (PHI) | ePHI stored on Navigation System is exposed to third parties. |  | Low | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Low | 0.2 | 2.9 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T05 | Conduct scavenging of ePHI at rest | V04 | Unauthorized modification to system bios | A05 | Patient health information at rest (PHI) | The threat actor actively steals data off the navigation unit by booting the system with an manipulated / self owned operating system to access the harddrive data and bypass the user role system provided by the original Stryker operating system. |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T06 | Theft of system or hard drives | V24 | Any unprotected hardware | A01 | Nav3,3i cart/ System running with windows 8.1 | When parts of the Navigation system are stolen the navigation system might be rendered unavailable for use. |  | Low | None | Low | Physical | low | low | none | unchanged | 0.66712 | 0.3916 | 2.5141 | 3.2 | Low | 0.2 | 2.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T07 | Data theft via physical media | V22 | Unprotected external USB Port | A05 | Patient health information at rest (PHI) | The threat actor actively steals data off the navigation unit by inserting physical media such as USB Sticks or External Drives. |  | High | None | None | Physical | low | low | none | unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T08 | Network-based denial of service (DoS) attack | V21 | Unprotected network port | A01 | Nav3,3i cart/ System running with windows 8.1 | The navigation system is attacked using a DoS attack. In the worst case the system is unavailable to the user. |  | None | None | Low | Network | Low | High | None | Unchanged | 1.2347 | 0.22 | 1.4124 | 2.7 | Moderate | 0.5 | 2.1 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T10 | Mis-configuration by user | V03 | Untrained/Malicious User | A02 | Admin Password / Credentials / System Configuration / Certificates | If a user misconfigures the system, this might lead to unavailability of the navigation system or make the system less performant for the user. |  | None | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.3916 | 2.5141 | 3.9 | Low | 0.2 | 2.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T11 | Remote exploitation | V21 | Unprotected network port | A09 | Navigation Accuracy | A threat actor uses the hospital network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system in order to manipulate navigation accuracy. |  | None | Low | Low | Network | High | Low | None | Changed | 1.7769 | 0.3916 | 2.7268 | 4.9 | Low | 0.2 | 3.4 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T11 | Remote exploitation | V21 | Unprotected network port | A01 | Nav3,3i cart/ System running with windows 8.1 | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. In the worst case the system is unavailable to the user. |  | None | Low | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
14 | T11 | Remote exploitation | V11 | Unpatched OS | A02 | Admin Password / Credentials / System Configuration / Certificates | If a remote exploit (e.g. "worm") is able to circumvent the operating system defenses to gain local access, then several other attack scenarios are enabled. In the worst case the system is unavailable to the user. |  | Low | Low | Low | Network | Low | High | None | Unchanged | 1.2347 | 0.52545 | 3.3734 | 4.7 | Low | 0.2 | 3.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T12 | Improper disposal of hard disks | V24 | Any unprotected hardware | A05 | Patient health information at rest (PHI) | ePHI might be exposed to third parties if the storage medium of the navigation unit is improperly disposed of. |  | High | None | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T13 | Man-in-the middle attack / intercept Navigation communication | V24 | Any unprotected hardware | A09 | Navigation Accuracy | Man-in-middle attack on localization communication such as infrared - Adversary manipulates navigation accuracy |  | None | Low | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | Very Low | 0.04 | 2.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T16 | Physical Manipulation of Hardware | V24 | Any unprotected hardware | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. Once defenses are circumvented in order to gain local access, several other Assets such as user credentials or PHI may be threatened. |  | Low | Low | none | Physical | Low | High | none | Unchanged | 0.29052 | 0.3916 | 2.5141 | 2.9 | Low | 0.2 | 2.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T16 | Physical Manipulation of Hardware | V24 | Any unprotected hardware | A09 | Navigation Accuracy | Adversary with direct mal-intent gains local access to the navigation flex cart platform and manipulates hardware such as tools and navigation camera. |  | Low | Low | none | Physical | Low | High | None | Unchanged | 0.29052 | 0.3916 | 2.5141 | 2.9 | Low | 0.2 | 2.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T17 | Gather information using open source discovery of organizational information | V26 | Exposure of sensitive product information | A01 | Nav3,3i cart/ System running with windows 8.1 | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. That information is used for an attack via network. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Low | 0.2 | 5.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T21 | Power Failure at primary facility | V24 | Any unprotected hardware | A01 | Nav3,3i cart/ System running with windows 8.1 | Adversary attacks power supply of the hospital or OR member cuts of power supply of navigation system by acccident. In both cases navigation system goes off to missing power. |  | None | None | High | Local | Low | None | None | Unchanged | 2.5151 | 0.56 | 3.5952 | 6.2 | Very Low | 0.04 | 3.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
21 | T24 | Manipulation of navigation camera firmware / memory | V24 | Any unprotected hardware | A09 | Navigation Accuracy | The threat actor manipulates the memory / firmware of navigation camera with the goal to bring the calibration off / manipulate the firmware to reduce navigation accuracy. |  | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | Low | 0.2 | 3.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
22 | T26 | Unexpected/Damaged input to CT Knee Intra-op | V39 | Lack of input sanitization in knee intra-op application | A14 | Patient Health Information in transit (PHI) | Deviation from the desired input data affects functionality and also creates major risk to patient health during surgery |  | None | High | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
23 | T27 | Malware affected input device | V22 | Unprotected external USB Port | A13 | Input device to Intra-OP | Device should be sanitized before connecting to the cart/system. Malware in the input device can cause severe threats |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
24 | T28 | Critical data transfer involving the removable device | V22 | Unprotected external USB Port | A14 | Patient Health Information in transit (PHI) | The removable devices shoud be restricted based on their purpose i.e, either as an input/output device. The input device should not take any data from the system intentionally/unintentionally |  | Low | Low | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.3916 | 2.5141 | 3.2 | Low | 0.2 | 2.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
25 | T28 | Critical data transfer involving the removable device | V36 | Malicious activities not being recorded in administrator/user logs | A14 | Patient Health Information in transit (PHI) | The removable devices shoud be restricted based on their purpose i.e, either as an input/output device. The input device should not take any data from the system intentionally/unintentionally |  | High | Low | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Very Low | 0.04 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
26 | T29 | Lack of encryption for input data to Intra-op | V32 | Unencrypted ePHI in transit (as output, input) | A14 | Patient Health Information in transit (PHI) | Data should not get modified in transit between two applications. The output from one application should maintain integrity in transition before reaching the other application |  | High | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.8064 | 5.1771 | 7.1 | Low | 0.2 | 5.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
27 | T29 | Lack of encryption for input data to Intra-op | V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | A14 | Patient Health Information in transit (PHI) | Data should not get modified in transit between two applications. The output from one application should maintain integrity in transition before reaching the other application |  | High | Low | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Low | 0.2 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
28 | T30 | Lack of authentication for the connected removable device | V05 | No proper validation for the removable devices | A13 | Input device to Intra-OP | Only hospital authenticated devices should be allowed to connect with the Intra-op application. If authentication is not enabled, any removable device can be connected to the system and trust can't be established |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
29 | T31 | Absence of data validation | V39 | Lack of input sanitization in knee intra-op application | A14 | Patient Health Information in transit (PHI) | Deviation from the desired input data affects functionality and also creates major risk to patient health during surgery |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Low | 0.2 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
30 | T33 | Broken authentication (user name, password…) | V06 | Lack of multi-factor authentication | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. Broken authentication threat can happen with single factor authentication |  | None | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.3916 | 2.5141 | 4.4 | Very Low | 0.04 | 2.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
31 | T34 | BRUTE force/Dictonary attacks | V06 | Lack of multi-factor authentication | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. BRUTE force attacks are possible with single factor authentication |  | None | Low | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
32 | T34 | BRUTE force/Dictonary attacks | V07 | Application allowing unlimited login attempts | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. BRUTE force attacks are possible with single factor authentication |  | None | Low | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
33 | T34 | BRUTE force/Dictonary attacks | V36 | Malicious activities not being recorded in administrator/user logs | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. BRUTE force attacks are possible with single factor authentication |  | None | Low | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
34 | T35 | Unauthorized modification of application resources | V36 | Malicious activities not being recorded in administrator/user logs | A11 | THOR Knee Intra-op Application | Every session has to get timed-out after certain period. If this is not enabled, then a continuous session exists without reauthentication and whole objective of authentication is bypassed |  | None | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.6568 | 4.2167 | 5.3 | Very Low | 0.04 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
35 | T36 | Unpatched/Unmanaged SOUP's | V11 | Unpatched OS | A01 | Nav3,3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
36 | T36 | Unpatched/Unmanaged SOUP's | V11 | Unpatched OS | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
37 | T36 | Unpatched/Unmanaged SOUP's | V12 | Potentially outdated SW-Library SOUPs | A01 | Nav3,3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
38 | T36 | Unpatched/Unmanaged SOUP's | V12 | Potentially outdated SW-Library SOUPs | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
39 | T36 | Unpatched/Unmanaged SOUP's | V16 | 3rd Party Component/SOUP Risk/Dependency not validated | A01 | Nav3,3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
40 | T36 | Unpatched/Unmanaged SOUP's | V16 | 3rd Party Component/SOUP Risk/Dependency not validated | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
41 | T36 | Unpatched/Unmanaged SOUP's | V17 | Legacy Systems (outdated HW/SW) being used | A01 | Nav3,3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
42 | T36 | Unpatched/Unmanaged SOUP's | V17 | Legacy Systems (outdated HW/SW) being used | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
43 | T37 | Exploitation of unsecured network elements | V21 | Unprotected network port | A07 | Computer/OS network identification | Adversary uses different kinds of tools to exploit the unused/unprotected network resouces should be scanned and protection should be enabled to safeguard from attacks |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
44 | T37 | Exploitation of unsecured network elements | V22 | Unprotected external USB Port | A07 | Computer/OS network identification | Adversary uses different kinds of tools to exploit the unused/unprotected network resouces should be scanned and protection should be enabled to safeguard from attacks |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
45 | T37 | Exploitation of unsecured network elements | V25 | Non-Tracing of unused external Interfaces (ports etc., ) | A07 | Computer/OS network identification | Adversary uses different kinds of tools to exploit the unused/unprotected network resouces should be scanned and protection should be enabled to safeguard from attacks |  | None | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.3916 | 2.5141 | 3.2 | Very Low | 0.04 | 2.6 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
46 | T38 | Outdated cryptographic components & techniques | V31 | Unencrypted ePHI at rest | A05 | Patient health information at rest (PHI) | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
47 | T38 | Outdated cryptographic components & techniques | V31 | Unencrypted ePHI at rest | A11 | THOR Knee Intra-op Application | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
48 | T38 | Outdated cryptographic components & techniques | V33 | Insecure Storage of Sensitive Information by application | A05 | Patient health information at rest (PHI) | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
49 | T38 | Outdated cryptographic components & techniques | V33 | Insecure Storage of Sensitive Information by application | A11 | THOR Knee Intra-op Application | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
50 | T38 | Outdated cryptographic components & techniques | V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | A05 | Patient health information at rest (PHI) | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
51 | T38 | Outdated cryptographic components & techniques | V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | A11 | THOR Knee Intra-op Application | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
52 | T39 | Unauthorized access/modification of secure data | V37 | Unauthorized Audit Log Manipulation, Log Injection-Tampering-Forging | A11 | THOR Knee Intra-op Application | Audit log data should be maintained to identify security breaches. With no proper authentication to the access of audit logs the content can be modified/deleted. |  | Low | High | Low | Local | Low | High | None | Unchanged | 0.79893 | 0.7323 | 4.7014 | 5.6 | Low | 0.2 | 4.9 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
53 | T40 | Risk of using windows 8.1 | V51 | Unsupported OS from vendor | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Monitor the vulnerability database to capture the high severity issues. 2. Check for any patches available & their compatibility to the OS. 3. If no patches available, inhouse patch development needs to be considered. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
54 | T40 | Risk of using windows 8.1 | V52 | Unavailability of support from 3rd party tools for outdated OS | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Audit & prepare a list of 3rd party tools. 2. Remove the unused 3rd party tools. 3. Prepare an alternate plan (same tool/new tool) 4. Based on the plan prepare the approach (by contacting the vendor). | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
55 | T40 | Risk of using windows 8.1 | V53 | OS Security tools from vendor not getting updated | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Audit & prepare a list of microsoft tools. 2. Remove the unused microsoft tools. 3. Prepare an alternate plan (same tool/new tool) 4. Based on the plan prepare the approach (by contacting the vendor). | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
56 | T40 | Risk of using windows 8.1 | V54 | Absence of periodic backup plan | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Customised & hardened os needs to be taken a backup. 2. Incase of a corrupted os, the above mentioned os can be reinstalled. 3. No need of periodic backup. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
57 | T40 | Risk of using windows 8.1 | V55 | Corrupted OS and absence of restore points | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Customised & hardened os needs to be taken a backup. 2. Incase of a corrupted os, the above mentioned os can be reinstalled. 3. No need of periodic backup. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
58 | T40 | Risk of using windows 8.1 | V56 | Remote exploitation of the system through network | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Network | Low | Low | None | Unchanged | 2.8353 | 0.6568 | 4.2167 | 7.1 | Moderate | 0.5 | 5.7 | MEDIUM | 1. Uninstall/remove all the networking software (ex: drivers). 2. Harden/disable the network interfaces of the cart. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
59 | T40 | Risk of using windows 8.1 | V57 | Any malicious service running in the system without being identified | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Periodically monitor the audit logs to identify any abnormal behaviour in the system. 2. Based on the analysis, observation needs to be reported. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
60 | T40 | Risk of using windows 8.1 | V58 | Unhardened OS and network interfaces | A01 | Nav3,3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Steps involved in hardening the os needs to be documented. 2. Uninstall/remove all the networking software (ex: drivers). 3. Harden/disable the network interfaces of the cart. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
61 | T42 | Application failed to restore data after any abnormal activity | V40 | Application not having periodic data backup & restore plan | A11 | THOR Knee Intra-op Application | Due to any abnormal activity application lost the patient/case data then without backup plan implemented can't resotre to the previous state before the abnormal activity |  | Low | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | High | 0.8 | 6.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
62 | T43 | Authentication bypass | V81 | Lack of Mandatory password policy for User(s) | A01 | Nav3,3i cart/ System running with windows 8.1 | It helps the attacker to attack/penetrate the system easily. |  | High | High | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.91482 | 5.8731 | 7.8 | High | 0.8 | 7.4 | HIGH | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
63 | T44 | Complete file system access | V82 | Unrestricted FileSystem Access & Modification | A01 | Nav3,3i cart/ System running with windows 8.1 | Adversary can gain access & alter the system files and able to modify the registry and system settings. |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
64 | T45 | All kinds of applications/scripts running | V83 | Uniform privilege for all Installables | A01 | Nav3,3i cart/ System running with windows 8.1 | Installation of malicious applications can make the system vulnerable and easier to attack. |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
65 | T46 | Unidentified virus existence | V84 | No Manual AV Scanning/Signature Update Option provided | A01 | Nav3,3i cart/ System running with windows 8.1 | If anti virus is outdated, no manual scanning can identify the hidden vulnerabilities in the system. Hence attacker can gain the system access easily. |  | Low | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
66 | T47 | All time system settings modification (during running & bootup) | V85 | Unauthorized Modification to Registry & System Settings | A15 | Nav3,3i device configurable information (for eg: bios, etc.,) | Adversary can gain access & alter the system files and able to modify the registry & system settings. |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
67 | T48 | Deletion of critical/sensitive data | V86 | Open access for Application data (resources, logs, case data etc., ) | A01 | Nav3,3i cart/ System running with windows 8.1 | Anyone can access/modify/delete the application data (resources, logs, case data, etc.,) |  | High | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.8064 | 5.1771 | 7.1 | Moderate | 0.5 | 6.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
68 | T49 | Illegal OS services access | V87 | Daemon access by unauthorized resources (ex., RDP, etc., ) | A01 | Nav3,3i cart/ System running with windows 8.1 | Adversary can gain access to the complete system through existing daemons. |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
69 | T50 | Injection of latest malware/virus | V88 | No Notification mentioning outdated AV Definitions | A01 | Nav3,3i cart/ System running with windows 8.1 | If anti virus is outdated, no manual scanning can identify the hidden vulnerabilities in the system. Hence attacker can gain the system access easily. |  | Low | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
70 | T51 | Deliver direct/undirect malware | V89 | No "on access" Scanning triggered for removable devices | A01 | Nav3,3i cart/ System running with windows 8.1 | If removable devices scanning is not performed, adversary can gain access to the system with direct/undirect malware. |  | Low | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
71 | T52 | Unvalidated external connected devices | V90 | Lack of authentication for externally connected devices (camera, etc.,) | A01 | Nav3,3i cart/ System running with windows 8.1 | Only authenticated devices should be allowed to connect. If authentication is not enabled, any external device can be connected to the system and trust can't be established |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Security Risk Control Measures
Security Risk Control Measures
Security Risk Control Measures
ID | Risk Control Measure | Category (optional) | Reference to more detailed specification
Note on A5: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V05 | No proper validation for the removable devices | A01 | Nav3,3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). |  | MEDIUM
2 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Developed code not validated using Static analysis | A01 | Nav3,3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) due to discovered defect or bug in Stryker software. |  | MEDIUM
3 | T02 | Deliver directed malware (CAPEC-185) | V03 | Untrained/Malicious User | A01 | Nav3,3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. |  | LOW
4 | T02 | Deliver directed malware (CAPEC-185) | v02 | Ineffective management of admin credentials | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs malware in order to gain admin credentials |  | LOW
5 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities |  | LOW
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Privileged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Document References (Optional)
Product Security Risk Table
References
ID | Title | Rev.* | Doc. No.
Auto
Auto
Note on A7: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.
Note on D10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: cross check if this is correct doc ids

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls