## Risk Table
Risk Table Document Number: | [5 empty] | Approvals | Name | Title | Signature | Date
Risk Table Revision: | [5 empty] | Risk Table Author
Risk Table Date: | [5 empty] | Risk Table Approvers (Risk Management Team)
Description/Identification of Medical Device/System:
Reference to Risk Management Plan (RMP):
Reference to DIOVV (DIOV and DIOTV): |  |  | D005020015
Reference to Intended Use (statement):
Scope of Risk Table: |  |  | <Reference Functional Analysis or similar doc.> |  |  | Procedure used: | SGTC-QPR-RSK-001
01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
Risk ID # | Functional Output | Hazard | Hazardous Situation | Harm | Person Affected | Causes | Planned/Existing Risk Control Measures | Severity Level | Occurrence Level | Risk Region | Additional Risk Control Measures | Design Input (Optional) | Severity Level | Revised Occurrence Level | Revised Risk Region | Benefit-risk analysis (High/Medium Risk) | Verification of Risk Control Measures (Implementation) | Verification of Risk Control Measures (Effectiveness) | Comments/Rationale (if necessary)
[8 empty] | (SHARM) | (OHARM) |  |  |  | (SHARM) | (OHARM)
T0-01 | Overall System Hazard Analysis | HZ10: Navigation information permanently not available | User is not able to operate on the system prior to the surgery | Patient can not be treated using navigation | Patient | Prior to surgery, the system cannot be started due to, a defect of a component (e.g. platform, camera, software, navigated instruments, tracker not working (damaged or broken) or tracker can not be used for surgery. |  | S0 | O3 | Low |  |  | S0
T0-02 | Overall System Hazard Analysis | HZ10: Navigation information permanently not available | Patient's surgery delayed prior to the surgery | Patient can not be treated using navigation | Service Technician/Sales Rep | Application can not be installed due to, 1. Installation info only available in english. 2. Installation Instruction/ Protocol (any user documentation) not available, missed in packaging due to mishandling or lost in storage/ transportation. 3. Pen drive is damaged, system is not able to read. 4. Wrong version of software is installed. |  | S0 | O2 | Low |  |  | S0
T0-03 | Overall System Hazard Analysis | HZ15: Navigation information delayed or temporarily not available | Usage of navigation system is interrupted during surgery | Slight OR time extension (≤15 minutes) or patient can not be treated using navigation | Patient | Software application crashes or stopped temporarily due to software failure (e.g. insufficient system memory, power breakdown, cybersecurity attack) Unintended touch on the input/output screens |  | S0 | O2 | Low |  |  | S0
T0-04 | Overall System Hazard Analysis | HZ10: Navigation information permanently not available | Abortion of navigated surgery | Patient can not be treated using navigation | Patient | Network connection causes software failure. Eg: memory full, screen blur, flickering, malware attack etc. |  | S0 | O1 | Low |  |  | S0
T0-05 | Overall System Hazard Analysis | HZ10: Navigation information permanently not available | Abortion of navigated surgery | Patient can not be treated using navigation | Patient | 1. One-time-use license (Single case license) expired because case took longer than license life time. 2. System crashed/faulty in during surgery and single case license can not be used. |  | S0 | O1 | Low |  |  | S0
T0-06 | Overall System Hazard Analysis | HZ30: Navigation system displays incorrect image/data | Positioning of the instruments is set to wrong location | Small prosthesis misalignment (≤5° but >3°) | Patient | Foreseeable misuse: Surgeon uses Pointer as a tool (e.g. for scratching); the tool bends and navigation is inaccurate |  | S2 | O2 | Low |  |  | S2
T0-07 | Overall System Hazard Analysis | HZ25: Parts or particals remain in the patient body | User did not notice parts falling off | Inflammation or infection | Patient | During impaction, parts of the Navigated Instruments fall off and remain in the patient (e.g. metal fragments). |  | S3 | O2 | Medium |  |  | S3
T1-01 | T1: Configure System | HZ30: Navigation system displays incorrect image/data | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) | Patient | User manipulates critical parameters in the system setting files (e.g. threshold for tool validation). |  | S3 | O2 | Medium |  |  | S3
T2-01 | T2 Edit User Preferences User Preferences | HZ15: Navigation information delayed or temporarily not available | User not able to see correct user data during navigation procedure | Slight OR time extension (≤15 minutes) | Patient | User has selected accidentally different settings in the User Preferences than the intended one. |  | S0 | O4 | Low |  |  | S0
T3-01 | T3 Enter Patient Data Patient Data | HZ30: Navigation system displays incorrect image/data | User not able to see correct user data during navigation procedure | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | The user has selected accidentally the wrong leg side in the patient data dialog. This could lead to wrong values which are not detected (e.g. varus/valgus swapping). |  | S3 | O3 | Medium |  |  | S3
T3-02 | T3 Enter Patient Data Patient Data | HZ30: Navigation system displays incorrect image/data | User not able to see correct user data during navigation procedure | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Patient data set loaded does not correspond to the treated patient, eg: data was deleted or overwritten due to cybersecurity attack Not detected by the user and surgery performed with wrong referencing data. Deleted records can not be restored caused by technical defect or poor cybersecurtiy measures. |  | S3 | O2 | Medium |  |  | S3
T3-03 | T3 Enter Patient Data Patient Data | HZ30: Navigation system displays incorrect image/data | User not able to see correct user data during navigation procedure | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Current patient has the same name as a former patient treated with the system. Patient files are swapped, manipulated or corrupted and wrong reference data are used for navigation. |  | S3 | O2 | Medium |  |  | S3
T3-04 | T3 Enter Patient Data Patient Data | HZ10: Navigation information permanently not available | User is not able to operate on the system prior to the surgery | Patient can not be treated using navigation | Patient | Patient data set entered or loaded does not correspond to the treated patient but detected by the user, -User is unaware to the feature that virtual keyboard is launched by S/W and not hardware key, thus faces issue in entering patient data -virtual keyboard freezes or updates at a slow rate, or there is no response at all and user is unable to enter patient data |  | S0 | O3 | Low |  |  | S0
T3-05 | T3 Enter Patient Data Patient Data | HZ30: Navigation system displays incorrect image/data | User is not able to do correct registration | Slightly misbalanced knee or slightly misaligned cuts | Patient | Different implant family is accidentally selected in patient data as used during surgery. |  | S2 | O3 | Low |  |  | S2
T4-01 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ50: Navigation system displays tool in wrong position | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Poor registration due to compromised pointer geometry, camera callibration, (e.g. corrupted tool data, poor pointer calibration, poor camera validation/ callibration) |  | S3 | O3 | Medium |  |  | S3
T4-02 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ50: Navigation system displays tool in wrong position | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | LED positions changed (e.g. moving due to extreme mechanical or thermical forces). |  | S3 | O2 | Medium |  |  | S3
T4-03 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ50: Navigation system displays tool in wrong position | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Camera system provides inaccurate emitter position data. Inherent safety measures implemented in the camera itself cannot detect inaccuracy. |  | S3 | O2 | Medium |  |  | S3
T4-04 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ50: Navigation system displays tool in wrong position | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | The system displays incorrect navigational information because the screen freezes or updates at a slow rate, or there is no response at all. This is not recognized by the surgeon. |  | S3 | O3 | Medium |  |  | S3
T4-05 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ50: Navigation system displays tool in wrong position | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | The system displays incorrect navigational information because one or more smart tools are located inaccurately in 3D space. |  | S3 | O3 | Medium |  |  | S3
T4-06 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ50: Navigation system displays tool in wrong position | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | The system displays incorrect navigational information because the accuracy of the position relative to the patient tracker degrades for distances bigger than 250mm (consider lever arm). |  | S2 | O3 | Low |  |  | S2
T4-07 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ10: Navigation information permanently not available | User is not able to operate on the system | Patient can not be treated using navigation | Patient | Prior to surgery, the system cannot be activated due to a defect of a system component (e.g. platform, camera, software, navigated instruments) |  | S0 | O3 | Low |  |  | S0
T4-08 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ10: Navigation information permanently not available | User is not able to operate on the system | Patient can not be treated using navigation | Patient | During the surgery, the system breaks down due to a defect of a system component (e.g. platform, camera, firmware, navigated instruments, tracker, empty battery, cybersecurity attack, computer virus and cannot be reactivated anymore) or its functionality is compromised (e.g. wrong screen colors, screen not readable). This is recognized, the system cannot be reactivated anymore and the surgery is aborted. |  | S0 | O3 | Low |  |  | S0
T4-09 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ15: Navigation information delayed or temporarily not available | User is interrupted while operating | Slight OR time extension (≤15 minutes) | Patient | Major software failure (e.g. crash, bad performance of the computer platform, Insufficient system memory, cybersecurity attacks) |  | S0 | O4 | Low |  |  | S0
T4-10 | T4 Setup System Setup System Dialog Validate Pointer Dialog | HZ15: Navigation information delayed or temporarily not available | User is interrupted while operating | Large OR time extension (>15 minutes) | Patient | During the surgery, the system breaks down due to a defect of a system component (e.g. platform, monitor, camera, software, or navigated instruments) or its functionality is compromised (e.g. wrong screen colors, screen not readable). This is recognized and the system is reactivated after more than 15 minutes. |  | S0 | O4 | Low |  |  | S0
T5-01 | T5 Register Patient (Surface based) | HZ50: Navigation system displays tool in wrong position | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Pin moves relatively to anatomy during surgery due to intervention forces on bone (sawing, drilling, etc.), no collision with instruments or persons assumed |  | S3 | O3 | Medium |  |  | S3
T5-02 | T5 Register Patient (Surface based) | HZ50: Navigation system displays tool in wrong position | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Pin collision with instrument or implant resulting in a Tracker movement. |  | S3 | O3 | Medium |  |  | S3
T5-03 | T5 Register Patient (Surface based) | HZ50: Navigation system displays tool in wrong position | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Tracker was detached from pin and reattached in a different position |  | S3 | O4 | High |  |  | S3
T5-04 | T5 Register Patient (Surface based) | HZ50: Navigation system displays tool in wrong position | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Either the 1st release of Universal Tracker (interface # HF105369) or the 1st release of Lightweight Tracker (HF193682 S01), which both does not have a guiding profile to support proper connection onto pin, is used as backup tracker. It is detached from pin and reattached in a slightly different position. |  | S3 | O2 | Medium |  |  | S3
T5-05 | T5 Register Patient (Surface based) | HZ50: Navigation system displays tool in wrong position | Anchoring system not placed correctly | Damage of critical soft tissue, blood vessels or nerves | Patient | During Pin insertion critical soft tissue is damaged (e.g. blood vessels are hit) |  | S3 | O3 | Medium |  |  | S3
T5-06 | T5 Register Patient (Surface based) | HZ50: Navigation system displays tool in wrong position | Anchoring system not placed correctly | Pin fragment remains in the patient. A medical intervention is required to remove the fragment | Patient | Pin breaks during insertion or removal and a fragment sticks in the bone. A medical intervention is required to remove the fragment Large pelvis movement, irreversible patient movement or camera movement resulting in inaccurate hip center determination. |  | S3 | O3 | Medium |  |  | S3
T5-07 | T5 Register Patient (Surface based) | HZ50: Navigation system displays tool in wrong position | User is not able to establish correct reference system | Small prosthesis misalignment (≤5° but >3°) | Patient | Movement of the femur at digitization just in one straight line instead of circles with different radii, resulting in inaccurate hip center determination. |  | S2 | O3 | Low |  |  | S2
T5-08 | T5 Register Patient (Surface based) | HZ30: Navigation system displays incorrect image/data | User is not able to do correct registration | Small prosthesis misalignment (≤5° but >3°) | Patient | Center of femoral head could not be calculated correctly because tibial tracker and femoral tracker were swapped. |  | S2 | O3 | Low |  |  | S2
T5-09 | T5 Register Patient (Surface based) | HZ15: Navigation information delayed or temporarily not available | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Landmarks digitized inaccurately (valid for all landmarks). -Insufficient coverage of surface during digitization -Lot of airpoints captured during digitization |  | S3 | O4 | High |  |  | S3
T5-10 | T5 Register Patient (Surface based) | HZ15: Navigation information delayed or temporarily not available | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Femoral center digitized inaccurately |  | S1 | O3 | Low |  |  | S1
T5-11 | T5 Register Patient (Surface based) | HZ15: Navigation information delayed or temporarily not available | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | During Femur Distal Condyle recording “air points” are digitized leading in wrong depth values. |  | S3 | O3 | Medium |  |  | S3
T5-12 | T5 Register Patient (Surface based) | HZ15: Navigation information delayed or temporarily not available | Registration is showing incorrect results | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | During Femur Posterior Condyle recording “air points” are digitized leading in wrong posterior condyles. |  | S3 | O3 | Medium |  |  | S3
T5-13 | T5 Register Patient (Surface based) | HZ15: Navigation information delayed or temporarily not available | Registration is showing incorrect results | Small prosthesis misalignment (≤5° but >3°) | Patient | Tibia center digitized inaccurately. |  | S2 | O3 | Low |  |  | S2
T5-14 | T5 Register Patient (Surface based) | HZ15: Navigation information delayed or temporarily not available | Registration is showing incorrect results | Small prosthesis misalignment (≤5° but >3°) | Patient | Tibial compartments digitized inaccurately; e.g. “air points” resulting in wrong depth. |  | S2 | O3 | Low |  |  | S2
T5-15 | T5 Register Patient (Surface based) | HZ15: Navigation information delayed or temporarily not available | Registration is showing incorrect results | Small prosthesis misalignment (≤5° but >3°) | Patient | Landmarks needed for ankle center calculation digitized inaccurately due to difficulties finding the surface at the draped ankle. |  | S2 | O3 | Low |  |  | S2
T7-01 | T7: Optional Position Implant | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Poor ROM or instable knee joint post operatively | Patient | System calculates wrong femoral implant size. |  | S3 | O3 | Medium |  |  | S3
T7-02 | T7: Optional Position Implant | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Poor ROM or instable knee joint post operatively | Patient | Soft tissue balancing is performed based on the values displayed in the implant planning dialog. If this is done without applying equal (and for the soft tissue situation of this knee adequate) force in flexion and extension following the navigation values could lead to inadequate ligament stability. |  | S3 | O3 | Medium |  |  | S3
T7-03 | T7: Optional Position Implant | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Slightly reduced ROM or slightly instable knee postoperatively | Patient | Inconsistency between implant database and actual implant |  | S2 | O3 | Low |  |  | S2
T7-04 | T7: Optional Position Implant | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Slightly reduced ROM or slightly instable knee postoperatively | Patient | Foreseeable misuse: the surgeon uses implants which are not supported by the system. |  | S2 | O3 | Low |  |  | S2
T7-05 | T7: Optional Position Implant | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Functional flexion axis (FFA) calculated inaccurate; e.g. due to ROM-recording with extreme internal- or external rotated tibia. |  | S3 | O3 | Medium |  |  | S3
T8-01 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Femoral or tibial tracker moved since landmark digitization. |  | S3 | O4 | High |  |  | S3
T8-02 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Patient Tracker moves due to vibrations during drilling and sawing. |  | S3 | O3 | Medium |  |  | S3
T8-03 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Large but reversible prosthesis misalignment (≤15° but >5°) requiring potentially a revision | Patient | Tracker was not attached correctly onto interface pin of navigated instrument. Tracker gets stuck on instrument not at correct position and this is not noticed by the surgeon |  | S3 | O2 | Medium |  |  | S3
T8-04 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Small prosthesis misalignment (≤5° but >3°) | Patient | Swapping the Resection Plane Probe with Posterior Plane Probe results in wrong resection depth. |  | S2 | O3 | Low |  |  | S2
T8-05 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Small prosthesis misalignment (≤5° but >3°) | Patient | The Stryker Plane Probe gets bent resulting in a wrong cut alignment. |  | S2 | O2 | Low |  |  | S2
T8-06 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Small prosthesis misalignment (≤5° but >3°) | Patient | Notching at anterior cortex of distal femur due to strong femur curvature in lateral view (posterior referencing technique only). |  | S2 | O3 | Low |  |  | S2
T8-07 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | Slight OR time extension (≤15 minutes) or patient can not be treated using navigation | Patient | Stryker Plane Probe does not fit in the slots of the cutting jig due to incompatibility of resection plane probes and cutting block or mechanical damages. |  | S0 | O3 | Low |  |  | S0
T8-08 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | <Gap Balancing> Misbalanced knee | Patient | Cuts, which have to be performed prior to gap monitoring, are recorded wrong. The calculation of the gap is based on the inaccurate cutting plane. |  | S3 | O3 | Medium |  |  | S3
T8-09 | T8: Resect Bones Dialogs | HZ40: The user selects the incorrect implant e.g. cup/stem | Positioning of the instruments is set to wrong location | <Gap Balancing> Misbalanced knee | Patient | Soft tissue balancing is performed based on the values displayed in gap balancing dialog without applying same adequate force in flexion and extension |  | S3 | O3 | Medium |  |  | S3
T8-10 | T8: Resect Bones Dialogs | HZ50: Navigation system displays tool in wrong position | User is not able to establish correct reference system | Small prosthesis misalignment (≤5° but >3°) | Patient | Patient not placed in the optimal working area / camera misplaced. |  | S2 | O3 | Low |  |  | S2
T8-11 | T8: Resect Bones Dialogs | HZ30: Navigation system displays incorrect image/data | Positioning of the instruments is set to wrong location | Small prosthesis misalignment (≤5° but >3°) | Patient | Resections performed by the surgeon do not match with the computed alignment of the navigated cutting jigs (e.g. sliding of the sawblade) |  | S2 | O5 | Medium |  |  | S2
T8-12 | T8: Resect Bones Dialogs | HZ30: Navigation system displays incorrect image/data | Positioning of the instruments is set to wrong location | Small prosthesis misalignment (≤5° but >3°) | Patient | The Stryker Plane Probe does not fit exactly into the cutting block slot (due to missing compatibility) resulting in a play and a small angular error |  | S2 | O3 | Low |  |  | S2
T8-13 | T8: Resect Bones Dialogs | HZ25: Parts or particals remain in the patient body | User did not notice parts falling off | Inflammation or infection | Patient | Parts fall off the tools or accessories, land in the patient’s wound and remain in the patient | [6 empty] | 0
T6-01 T9-01 T10-01 | T6/T9/T10 Analyze alignment (optional) Analyze Initial Alignment Analyze Trial Alignment Analyze Final Alignment | HZ00: No Hazard | None | No harm | N/A | None | [6 empty] | 0
T11-01 | T11: Report Create Report | HZ00: No Hazard | None | No harm | N/A | None | [6 empty] | 0

## Functional Analysis
User input
T1
Edit system setting file
Screen display
System setting file
User input
T2
Edit User preferences
Screen display
User setting file
User input
T3
Enter patient data
Screen display
Patient file
User input
T4
System setup
Screen display
Validate pointer
Localizer system
User input
T5 Registration
Register patient in general
Screen display
Register Femur
System setting file
User setting file
Localizer system
Register Tibia
Verify registration
Co-ordinate system of the Femur
Co-ordinate system of the Tibia
T6, T9, T10 Analyze alignment
Analyze trail alignment
Screen display
User input
Analyze initial alignment
Localizer system
Analyze final alignment
Kinematics overview dialog
Co-ordinate system of the Femur
Co-ordinate system of the Tibia
Record table
Analyze Varus-Valgus
Analyze curves
T7 Optional position implant
Optional - Resect proximal Tibia
User input
Optional – soft tissue balance
Screen display
Localizer system
Position implant
Co-ordinate system of the Femur
Co-ordinate system of the Tibia
Optional and separate license
Implant data base
Implant positioning plan
Update of Tibia coordinate system
Update of femur coordinate system
User input
T11
Create Report
Report
Screen display
T8 Resect bone
Resect bone overview dialog
User input
Resect femur dialog
Screen display
Localizer system
Resect tibia dialog
Co-ordinate system of the Femur
Co-ordinate system of the Tibia
Kinematic dialog
Implant data base
Tibia Implant system
Femur Implant system
User Input
Localizer system
Implant database
Screen Display
Report
T3 – Enter patient data
T4 – Setup system
T5 – Registration
T6 – Optional: Analyze initial alignment
T7 – Optional: Position implant
T8 – Optional: Resect bone
T9 – Optional: Analyze trail alignment
T10 – Optional: Analyze Final alignment
T11 – Create Report
Patient file/ Report

## Open points
1 | Planning needs to installed in Nav system??
2 | Pateint Data can be pulled directly from CT knee planning output
3 | Surgeon can modify the data?
4 | Registration (Surface based or Landmark based-Hip centre and Ankle centre)
5 | Pateint File?
6 | Message or Risk mitigation warning with Navigation system to be added to Intra-Op software

## Risk Matrix
Risk Matrix - with Planned/Existing Risk Control Measures (through DR4)
PROBABILITY of OCCURRENCE of HARM | High | O5 | 0 | 0 | 0 | 0 | 0 | 0
| Moderate | O4 | 0 | 0 | 0 | 0 | 0 | 0
| Low | O3 | 0 | 0 | 0 | 0 | 0 | 0
| Remote | O2 | 0 | 0 | 0 | 0 | 0 | 0
| Negligible | O1 | 0 | 0 | 0 | 0 | 0 | 0
|  |  | S0 | S1 | S2 | S3 | S4 | S5
|  |  | None | Limited | Temporary or Reversible | Intervention Required | Permanent Impairment | Death
SEVERITY of HARM
|  |  | No adverse health consequences | Transient, minor impairment or (minor) complaints | Temporary injury or impairment (not requiring professional medical intervention) | Necessitates professional medical or surgical intervention (to prevent permanent impairment) | Permanent impairment of body function or permanent damage to a body structure | Death
Revised Risk Matrix - with All (Planned/Existing and Additional) Risk Control Measures
PROBABILITY of OCCURRENCE of HARM | High | O5 | 0 | 0 | 0 | 0 | 0 | 0
| Moderate | O4 | 0 | 0 | 0 | 0 | 0 | 0
| Low | O3 | 0 | 0 | 0 | 0 | 0 | 0
| Remote | O2 | 0 | 0 | 0 | 0 | 0 | 0
| Negligible | O1 | 0 | 0 | 0 | 0 | 0 | 0
|  |  | S0 | S1 | S2 | S3 | S4 | S5
|  |  | None | Limited | Temporary or Reversible | Intervention Required | Permanent Impairment | Death
SEVERITY of HARM
|  |  | No adverse health consequences | Transient, minor impairment or (minor) complaints | Temporary injury or impairment (not requiring professional medical intervention) | Necessitates professional medical or surgical intervention (to prevent permanent impairment) | Permanent impairment of body function or permanent damage to a body structure | Death
Risk Matrix
PROBABILITY of OCCURRENCE of HARM | High | O5 | Low | Medium | Medium | High | High | High
| Moderate | O4 | Low | Low | Medium | High | High | High
| Low | O3 | Low | Low | Low | Medium | High | High
| Remote | O2 | Low | Low | Low | Medium | Medium | High
| Negligible | O1 | Low | Low | Low | Low | Medium | Medium
|  |  | S0 | S1 | S2 | S3 | S4 | S5
|  |  | None | Limited | Temporary or Reversible | Intervention Required | Permanent Impairment | Death
SEVERITY of HARM
|  |  | No adverse health consequences | Transient, minor impairment or (minor) complaints | Temporary injury or impairment (not requiring professional medical intervention) | Necessitates professional medical or surgical intervention (to prevent permanent impairment) | Permanent impairment of body function or permanent damage to a body structure | Death

## Reference (hidden)
NOTE: DO NOT EDIT THIS DATA AS IT IS USED TO GENERATE RISK TABLE FEATURES!
| Severity |  | Occurrence
1 |  | 1 |  |  | Blank | S0 | S1 | S2 | S3 | S4 | S5
2 | S5 | 2 | O5 | O5 |  | Low | Medium | Medium | High | High | High
3 | S4 | 3 | O4 | O4 |  | Low | Low | Medium | High | High | High
4 | S3 | 4 | O3 | O3 |  | Low | Low | Low | Medium | High | High
5 | S2 | 5 | O2 | O2 |  | Low | Low | Low | Medium | Medium | High
6 | S1 | 6 | O1 | O1 |  | Low | Low | Low | Low | Medium | Medium
7 | S0 |  |  | Blank