## Header (Optional)
Product Security Risk Table
Ortho Guidance Precision Knee 6.0 Software
Document Control
| Author: | Siva Kumar
| Document No.: | D0000055902
Revision No.:
| Project: | CT Knee Intra-Op | DC-0000002518
| Project Lead: | Akhil Gupta
Change Records
| Rev. No. | Comment | Date | Author
| AA | Initial revision | See onePLM
Note on B19: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## Document References (Optional)
Product Security Risk Table
References
ID | Title | Rev.* | Doc. No.
SRS | SRS - OrthoMap® Precision Knee 6.0 Software - (DC-0000002518 Precision Knee 6.0) [SRS] | Auto | D0000039056
RARC | Risk Assessment - Ortho Guidance Precision Knee 6.0 Software - (DC-0000002518 Precision Knee 6.0) [RARC] | Auto | D0000039062
Note on A7: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Stryker Flex Cart Navigation System
Scope: |  | OrthoMap Precision Knee
Date: |  | <yyyyy-mm-dd>
Conducted by: |  | <Author Name / Function / Organization> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description
A00 | Physical Asset | New cart running with Linux OS | Contains software module which controls the overall functionality of Knee System i.e. intra-operative planning and guidance.
A01 | Physical Asset | Nav3i cart/ System running with windows 8.1 | Contains software module which controls the overall functionality of Knee System i.e. intra-operative planning and guidance.
A02 | Information Asset | Admin Password / Credentials / System Configuration / Certificates | Admin Passwords/Credentials allow modifying config files, accessing ePHI, software binaries, computer configuration files, log files, and computer resources as well as installing applications, security certificates, OS updates.
A03 | Information Asset | User Password / Credentials / System Configuration | User Passwords/Credentials provide access to ePHI, Software Binaries, read access to config files.
A05 | Information asset | Patient health information at rest | Datastored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A07 | Information asset | Computer/OS network identification | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A08 | Information Asset | Intellectual Property | Algorithms and other core-elements of interest to competitors.
A09 | Physical Asset | Navigation Accuracy | Navigation Accuracy of FP8000 Camera, critical components for essential performance
A10 | Information Asset | SSH credentials for system admin access | Only using an 2048bit SSH key allows full access to all system functions. That SSH key is not provided to the customer and can be used by Stryker internals only.
A11 | Information asset | THOR Knee Intra-op Application | Contains Knee intra-op software which controls the overall features for segmentation, landmark planning, implant size calculation etc.
A12 | Information asset | Patient Personal data | Contains patient personal details such as patient identity, DOB, Age/Gender, image data
A13 | Physical Asset | Input to Intra-Op | Any kind of removable device which is connected from knee intra-op to knee planning application
A14 | Physical Asset | Removable Devices | Removable devices provide input to knee intra-op
A15 | Physical Asset | System with hw resources/configurable bios | Knee Intra-Op running Cart/System having HW resources and Configurable Bios
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user credentials | Yes | n/a
V02 | Ineffective management of admin credentials | Yes | n/a
V03 | Untrained/Malicious User | Yes | n/a
V04 | Unauthorized modification to system bios | Yes | n/a
V05 | No proper validation for the removable devices | Yes | n/a
V06 | Lack of multi-factor authentication | Yes | n/a
V07 | Application allowing unlimited login attempts | Yes | n/a
V08 | No session time out implemented for the application | Yes | n/a
SBOM/Software Management
V11 | Unpatched OS | Yes | n/a
V12 | Potentially outdated SW-Library SOUPs | Yes | n/a
V13 | Developed code not validated using static/Dynamic Tools | Yes | n/a
V14 | UnSafely Compiled Application software | Yes | n/a
V15 | Lack of periodic software vulnerability management | Yes | n/a
V16 | 3rd Party Component/SOUP Risk/Dependency not validated | Yes | n/a
V17 | Legacy Systems (outdated HW/SW) being used | Yes | n/a
Access points
V21 | Unprotected network port | Yes | n/a
V22 | Unprotected external USB Port | Yes | n/a
V24 | Any unprotected hardware | Yes | n/a
V25 | Non-Tracing of unused external Interfaces (ports etc., ) | Yes | n/a
Data
V31 | Unencrypted ePHI at rest | Yes | n/a
V32 | Unencrypted ePHI in flight (as output, input) | Yes | n/a
V33 | Insecure Storage of Sensitive Information by application | Yes | n/a
V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | Yes | n/a
V35 | Exposure of Private Information (Health/Personal data) | Yes | n/a
V36 | Malicious activities not being recorded in administrator/user logs | Yes | n/a
V37 | Unauthorized Audit Log Manipulation, Log Injection-Tampering-Forging | Yes | n/a
V38 | Information Exposure through Log Files | Yes | n/a
V39 | Lack of input sanitization in knee intra-op application | Yes | n/a
InSecure Configurations of Resources
V40 | Lack of validation for the changed configuration controls | Yes | n/a
V41 | Software Download from malicious source | Yes | n/a
V42 | Lack of control over the Automated Software Update | Yes | n/a
V43 | External Control of System or Configuration Setting | Yes | n/a
OS vulnerabilities
V51 | Unsupported OS from vendor | Yes | n/a
V52 | Unavailability of support from 3rd party tools for outdated OS | Yes | n/a
V53 | OS tools from vendor not getting updated | Yes | n/a
V54 | Absence of periodic backup plan | Yes | n/a
V55 | Corrupted OS and absence of restore points | Yes | n/a
V56 | Remote exploitation of the system from intranet | Yes | n/a
V57 | Any malicious service running in the system without being identified | Yes | n/a
V58 | Unhardened OS and network interfaces | Yes | n/a
Linux OS
V61 | Absence of periodic backup plan | Yes | n/a
V62 | Corrupted OS and absence of restore points | Yes | n/a
V63 | Remote exploitation of the system from intranet | Yes | n/a
V64 | Any malicious service running in the system without being identified | Yes | n/a
V65 | Unhardened OS and network interfaces | Yes | n/a
V66 | Securtiy packages not properly configured during OS build | Yes | n/a
V67 | Critical patches not identified and integrated within timeline | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware (CAPEC-185) | Threat source delivers malware by providing removable media prepared with malware. e.g. Removable media is left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System. Malware exploits a known vulnerability in order to gain access to restricted data or to manipulate data. Undirected attack on computer systems. | TSA-3 - Skript Kiddies, TSA-1 Individual | Yes | n/a
T02 | Deliver directed malware (CAPEC-185) | Threat source delivers malicious software on a removable media which was designed to exploit a known vulnerability of the Navigation System. Directed attack on the Navigation System using knowledge of the Navigation System. | TSA-2 Organization | Yes | n/a
T03 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies and TSA-1 Individual | Yes | n/a
T04 | Intercept network communication | Adversary manipulates the network environment to intercept network communication to gain ePHI, passwords, credit card information etc. or to manipulate DICOM data etc. (man-in-the middle attack) | TSA-2 Organization | No | For Imageless applications, data is never transferred over network.
T05 | Conduct scavenging of ePHI at rest | Adversary gains unauthorized access to the system and steals ePHI information | TSA-1 - Individual | Yes | n/a
T06 | Theft of system or hard drives | A thief steals the navigation system and/or the hard drives as loot, which can lead to unavailability of the navigation system and/or leak of Protected Health Information (PHI). | TSA-1 - Individual | Yes | n/a
T07 | Data theft via physical media | The malicious or inadvertent copy of ePHI on portable media (USB Sticks, DVD drives) from the navigation system. | TSA-2 Organization | Yes | n/a
T08 | Network-based denial of service (DoS) attack | A threat actor seeks to make a machine or network resource unavailable to its intended users by overloading the network resource, the network or target services with malicious data traffic. | TSA-3 - Skript Kiddies TSN-4 Malfunctioning Software / OS TSN-3 Malfunctioning device | Yes | n/a
T09 | Exploit physical access of authorized staff to gain access to organizational facilities. | Adversary follows (“tailgates”) authorized individuals into secure/controlled locations with the goal of gaining access to facilities, circumventing physical security checks. | TSA-1 Individual | No | HDO responsible for physical security controls
T10 | Mis-configuration by user | A user unknowingly or maliciously mis-configures the system so that it malfunctions or no longer functions at all. | TSN-1 Accidental | Yes | n/a
T11 | Remote exploitation | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. This is likely inadvertent, targetting general vulnerability in popular operating systems. After successful exploit, continues as malware threat. | TSA-1 Individual | Yes | n/a
T12 | Improper disposal of hard disks | A user disposes the entire system or the hard disks (for instance after a service issue or at the end of the service life of the product). The hard disks still contain patient data. | TSN-1 Accidental | Yes | n/a
T13 | Man-in-the middle attack / intercept Navigation communication | Adversary manipulates the localization communication such as infrared - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T14 | Conduct scavenging of IP at rest | Adversary gains unauthorized access to the system and steals IP off the system | TSA-1 Individual | No | HDO responsible for physical security controls
T15 | Frequency Jamming or Destruction of Wireless Communication | Adversary manipulates the wireless network environment in order to render the service unavailable. | TSA-3 - Skript Kiddies | No | Ortho Imageless application will not be impacted due to unavailability of wireless network
T16 | Physical Manipulation of Hardware | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. | TSA-2 Organization or TSA-1 Individual | Yes | n/a
T17 | Gather information using open source discovery of organizational information | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. | TSA-3 - Skript Kiddies | Yes | n/a
T18 | Craft counterfeit certificates. | Adversary counterfeits or compromises a certificate authority, so that malware or connections will appear legitimate. | TSA-1 Individual | No | Device is a pass through of data in transit from end point devices. The certificate validation is the responsibility of the end points.
T19 | Install persistent and targeted sniffers on organizational information systems and networks. | Adversary places within internal organizational information systems or networks software designed to (over a continuous period of time) collect (sniff) network traffic. | TSA-2 Organization | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T20 | Insert malicious scanning devices (e.g., wireless sniffers) inside facilities. | Adversary uses postal service or other commercial delivery services to deliver to organizational mailrooms a device that is able to scan wireless communications accessible from within the mailrooms and then wirelessly transmit information back to adversary. | TSA-1 Individual | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T21 | Power Failure at primary facility | Power Failure at primary facility makes facility including the navigation sytem inoperable | TSN-7 Power Failure | Yes | n/a
T22 | Theft of SSH key for root access of Navigation Systems | Adverary uses SSH root key to get unauthorized access to the full system. | TSA-1 Individual | No | SSH root key is stored safely at Stryker facility and never delivered to the HDO / customer.
T23 | Manipulation of tool firmware / memory | Adversary manipulates the tool related calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | No | For Imageless applications, no tool is calibrated.
T24 | Manipulation of navigation camera firmware / memory | Adversary manipulates the navgation camera related calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T25 | Manipulation of EM box firmware / memory | Adversary manipulates the EM box firmware - thereby influencing the navigation accuracy | TSA-1 Individual | No | EM is not used by Ortho Imageless applications
T26 | Unexpected/Damaged input to CT Knee Intra-op | Input from the Knee planning got modified in transit and unexpected input received by Intra-Op | TSA-1 Individual | Yes | n/a
T27 | Malware affected input device | An unscanned input device can inject a malware into the intra-op device | TSA-1 Individual | Yes | n/a
T28 | Critical data transfer to the removable device | Write protect should be enable to avoid the transfer of critical data | TSA-1 Individual | Yes | n/a
T29 | Lack of encryption for input data to Intra-op | If input data is not encrypted then data can be easily accessible by anyone | TSA-1 Individual | Yes | n/a
T30 | Lack of authentication for the connected removable device | All the connected devices should not be allowed to provide input. Proper authentication should be carried out | TSA-1 Individual | Yes | n/a
T31 | Absence of data validation | Any input data provided to the Intra-op should be properly validated | TSA-1 Individual | Yes | n/a
T32 | Unsecured API interface | All API request need to have a secure header implementation | TSA-1 Individual | Yes | n/a
T33 | Broken authentication (user name, password…) | BRUTE force/Dictonary attacks are possible | TSA-1 Individual | Yes | n/a
T34 | BRUTE force/Dictonary attacks | BRUTE force/Dictonary attacks are possible if system not locked after certain retries | TSA-1 Individual | Yes | n/a
T35 | No session expiry | Session should get terminated after specified time-out | TSA-1 Individual | Yes | n/a
T36 | Unpatched OS/SOUP's | Outdated/Unpatched softwares can be easily exploited | TSA-1 Individual | Yes | n/a
T37 | Exploitation of unsecured network elements | Unused ports/other network interfaces has to be identified | TSA-1 Individual | Yes | n/a
T38 | Outdated cryptographic components & techniques | Keys, certificates, hashing techniques, etc... should be up-to-date with the latest standards | TSA-1 Individual | Yes | n/a
T39 | Unauthorized access/modification of secure data | Audit (log) data is considered as a sensitive data. It should be secured with proper access control | TSA-1 Individual | Yes | n/a
T40 | Risk of using windows 8.1 | Windows 8.1 & prior versions are considered as an EOL OS. Vendor no more provide support for these OS's | TSA-1 Individual | Yes | n/a
T41 | Risk of using Linux | Linux is an open-source OS. Some vendors has linux flavors which still needs to be considered for risk analysis | TSA-1 Individual | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V11 | Unpatched OS | A01 | Nav3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). | [RARC.RI040] | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Moderate | 0.5 | 4.1 | MEDIUM | Application Whitelisting Firewall Virus Scan User authentication | SRS.SR01 SRS.SR02 SRS.SR03 SRS.SR07 | TestReport #123 | Low | None | None | Network | Low | None | None | Changed | 3.887 | 0.22 | 1.4363 | 5.8 | 3.7 | LOW
2 | T02 | Deliver directed malware (CAPEC-185) | V24 | Any unprotected hardware | A01 | Nav3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. | [RARC.RI040] | Low | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Moderate | 0.5 | 5.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T10 | Mis-configuration by user | V03 | Untrained/Malicious User | A02 | Admin Password / Credentials / System Configuration / Certificates | If a user misconfigures the system, this might lead to unavailability of the navigation system or make the system less performant for the user. | [RARC.RI110] | None | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.3916 | 2.5141 | 3.9 | Low | 0.2 | 2.8 | LOW | Application Whitelisting Firewall Virus Scan User authentication | SRS.SR01 SRS.SR02 SRS.SR03 SRS.SR07 | TestReport #123 | High | High | Low | Physical | High | High | Required | Unchanged | 0.12109 | 0.84899 | 5.4505 | 5.6 | 5.5 | MEDIUM | Justification
4 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities | [RARC.RI110] | Low | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.3916 | 2.5141 | 4.8 | Moderate | 0.5 | 3.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T05 | Conduct scavenging of ePHI at rest | V01 | Ineffective management of user credentials | A05 | Patient health information at rest | ePHI stored on Navigation System is exposed to third parties. | [RARC.RI180] | High | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Low | 0.2 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T02 | Deliver directed malware (CAPEC-185) | v02 | Ineffective management of admin credentials | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs malware in order to gain admin credentials | [RARC.RI040] | Low | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Moderate | 0.5 | 5.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T13 | Man-in-the middle attack / intercept Navigation communication | V24 | Any unprotected hardware | A09 | Navigation Accuracy | Man-in-middle attack on localization communication such as infrared - Adversary manipulates navigation accuracy | [RARC.RI180] | Low | High | High | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Very Low | 0.04 | 5.5 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T08 | Network-based denial of service (DoS) attack | V21 | Unprotected network port | A01 | Nav3i cart/ System running with windows 8.1 | The navigation system is attacked using a DoS attack. In the worst case the system is unavailable to the user. | [RARC.RI110] | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Moderate | 0.5 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T11 | Remote exploitation | V21 | Unprotected network port | A09 | Navigation Accuracy | A threat actor uses the hospital network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system in order to manipulate navigation accuracy. | [RARC.RI180] | None | High | Low | Network | High | Low | None | Changed | 1.7769 | 0.6568 | 4.7173 | 7.1 | Low | 0.2 | 5.5 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T11 | Remote exploitation | V21 | Unprotected network port | A01 | Nav3i cart/ System running with windows 8.1 | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. In the worst case the system is unavailable to the user. | [RARC.RI110] | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Moderate | 0.5 | 5.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T11 | Remote exploitation | V11 | Unpatched OS | A02 | Admin Password / Credentials / System Configuration / Certificates | If a remote exploit (e.g. "worm") is able to circumvent the operating system defenses to gain local access, then several other attack scenarios are enabled. In the worst case the system is unavailable to the user. | [RARC.RI110] | High | High | Low | Network | Low | High | None | Unchanged | 1.2347 | 0.84899 | 5.4505 | 6.7 | Moderate | 0.5 | 6.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T12 | Improper disposal of hard disks | V24 | Any unprotected hardware | A05 | Patient health information at rest | ePHI might be exposed to third parties if the storage medium of the navigation unit is improperly disposed of. | N/A | High | None | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T06 | Theft of system or hard drives | V24 | Any unprotected hardware | A01 | Nav3i cart/ System running with windows 8.1 | When parts of the Navigation system are stolen the navigation system might be rendered unavailable for use. | N/A | High | none | High | Physical | low | low | None | unchanged | 0.66712 | 0.8064 | 5.1771 | 5.9 | Low | 0.2 | 5.4 | MEDIUM | Encrpyted ePHI in flight | SRS.SR04 | TestReport #789 | Low | None | None | Network | High | None | None | Unchanged | 2.2212 | 0.22 | 1.4124 | 3.7 | 1.9 | LOW
14 | T07 | Data theft via physical media | V22 | Unprotected external USB Port | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by inserting physical media such as USB Sticks or External Drives. | N/A | High | none | none | Physical | low | low | None | unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T16 | Physical Manipulation of Hardware | V24 | Any unprotected hardware | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. Once defenses are circumvented in order to gain local access, several other Assets such as user credentials or PHI may be threatened. | [RARC.RI180] | Low | High | none | Physical | Low | High | None | Unchanged | 0.29052 | 0.6568 | 4.2167 | 4.6 | Low | 0.2 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T21 | Power Failure at primary facility | V24 | Any unprotected hardware | A01 | Nav3i cart/ System running with windows 8.1 | Adversary attacks power supply of the hospital or OR member cuts of power supply of navigation system by acccident. In both cases navigation system goes off to missing power. | [RARC.RI110] | None | None | High | Local | Low | None | None | Unchanged | 2.5151 | 0.56 | 3.5952 | 6.2 | Very Low | 0.04 | 3.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T17 | Gather information using open source discovery of organizational information | V21 | Unprotected network port | A01 | Nav3i cart/ System running with windows 8.1 | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. That information is used for an attack via network. | [RARC.RI040] | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Low | 0.2 | 5.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Developed code not validated using static/Dynamic Tools | A01 | Nav3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) due to discovered defect or bug in Stryker software. | [RARC.RI040] | Low | Low | High | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | Very Low | 0.04 | 4.8 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T05 | Conduct scavenging of ePHI at rest | V04 | Unauthorized modification to system bios | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by booting the system with an manipulated / self owned operating system to access the harddrive data and bypass the user role system provided by the original Stryker operating system. | [RARC.RI180] | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Low | 0.2 | 5.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T24 | Manipulation of navigation camera firmware / memory | V24 | Any unprotected hardware | A09 | Navigation Accuracy | The threat actor manipulates the memory / firmware of navigation camera with the goal to bring the calibration off / manipulate the firmware to reduce navigation accuracy. | [RARC.RI180] | High | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Low | 0.2 | 5 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
21 | T16 | Physical Manipulation of Hardware | V24 | Any unprotected hardware | A09 | Navigation Accuracy | Adversary with direct mal-intent gains local access to the navigation flex cart platform and manipulates hardware such as tools and navigation camera. | [RARC.RI180] | Low | High | none | Physical | Low | High | None | Unchanged | 0.29052 | 0.6568 | 4.2167 | 4.6 | Low | 0.2 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
22 | T26 | Unexpected/Damaged input to CT Knee Intra-op | V39 | Lack of input sanitization in knee intra-op application | A13 | Input to Intra-Op | Deviation from the desired input data affects functionality and also creates major risk to patient health during surgery |  | None | None | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | 1. Application should validate the input data 2. Data format and expected input should be properly documented | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
23 | T27 | Malware affected input device | V41 | Software Download from malicious source | A13 | Input to Intra-Op | Device should be sanitized before connecting to the cart/system. Malware in the input device can cause severe threats |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | All removable devices connecting to the system should be scanned | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
24 | T27 | Malware affected input device | V41 | Software Download from malicious source | A14 | Removable Devices | Device should be sanitized before connecting to the cart/system. Malware in the input device can cause severe threats |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | All removable devices connecting to the system should be scanned | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
25 | T28 | Critical data transfer to the removable device | V22 | Unprotected external USB Port | A13 | Input to Intra-Op | The removable devices shoud be restricted based on their purpose i.e, either as an input/output device. The input device should not take any data from the system intentionally/unintentionally |  | High | Low | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | 1. Critical/Sensitive data should be protected. 2. Data transfer to the removable devices should happen with proper authentication. 3. System security elements such as firewall, antivirus etc.. should be properly identified | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
26 | T28 | Critical data transfer to the removable device | V22 | Unprotected external USB Port | A14 | Removable Devices | The removable devices shoud be restricted based on their purpose i.e, either as an input/output device. The input device should not take any data from the system intentionally/unintentionally |  | High | Low | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | 1. Critical/Sensitive data should be protected. 2. Data transfer to the removable devices should happen with proper authentication. 3. System security elements such as firewall, antivirus etc.. should be properly identified | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
27 | T28 | Critical data transfer to the removable device | V36 | Malicious activities not being recorded in administrator/user logs | A13 | Input to Intra-Op | The removable devices shoud be restricted based on their purpose i.e, either as an input/output device. The input device should not take any data from the system intentionally/unintentionally |  | None | None | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0 | 0 | 0 | Very Low | 0.04 | 0 | NONE | 1. Critical/Sensitive data should be protected. 2. Data transfer to the removable devices should happen with proper authentication. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
28 | T28 | Critical data transfer to the removable device | V36 | Malicious activities not being recorded in administrator/user logs | A14 | Removable Devices | The removable devices shoud be restricted based on their purpose i.e, either as an input/output device. The input device should not take any data from the system intentionally/unintentionally |  | None | None | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0 | 0 | 0 | Very Low | 0.04 | 0 | NONE | 1. Critical/Sensitive data should be protected. 2. Data transfer to the removable devices should happen with proper authentication. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
29 | T29 | Lack of encryption for input data to Intra-op | V32 | Unencrypted ePHI in flight (as output, input) | A13 | Input to Intra-Op | Data should not get modified in transit between two applications. The output from one application should maintain integrity in transition before reaching the other application |  | High | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.8064 | 5.1771 | 7.1 | Low | 0.2 | 5.6 | MEDIUM | 1. Critical/Sensitive data should be protected. 2. Data transfer from one device to another device should be encrypted. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
30 | T29 | Lack of encryption for input data to Intra-op | V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | A13 | Input to Intra-Op | Data should not get modified in transit between two applications. The output from one application should maintain integrity in transition before reaching the other application |  | High | Low | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Low | 0.2 | 4.6 | MEDIUM | 1. Critical/Sensitive data should be protected. 2. Data transfer from one device to another device should be encrypted. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
31 | T30 | Lack of authentication for the connected removable device | V05 | No proper validation for the removable devices | A13 | Input to Intra-Op | Only hospital authenticated devices should be allowed to connect with the Intra-op application. If authentication is not enabled, any removable device can be connected to the system and trust can't be established |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | 1. All the removable devices should not be allowed to connect with the systsem 2. Authentication/Validation of input device should be mandatory | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
32 | T30 | Lack of authentication for the connected removable device | V05 | No proper validation for the removable devices | A14 | Removable Devices | Only hospital authenticated devices should be allowed to connect with the Intra-op application. If authentication is not enabled, any removable device can be connected to the system and trust can't be established |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | 1. All the removable devices should not be allowed to connect with the systsem 2. Authentication/Validation of input device should be mandatory | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
33 | T31 | Absence of data validation | V39 | Lack of input sanitization in knee intra-op application | A13 | Input to Intra-Op | Deviation from the desired input data affects functionality and also creates major risk to patient health during surgery |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Low | 0.2 | 4.6 | MEDIUM | 1. Application should validate the input data 2. Data format and expected input should be properly documented | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
34 | T33 | Broken authentication (user name, password…) | V06 | Lack of multi-factor authentication | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. Broken authentication threat can happen with single factor authentication |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Very Low | 0.04 | 4.3 | MEDIUM | 1. Require multi-factor authentication 2. Limit authentication attempts (rate Limiting) 3. Maintain Access Logs 4. Maintain Server Security Logs 5. Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
35 | T34 | BRUTE force/Dictonary attacks | V06 | Lack of multi-factor authentication | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. BRUTE force attacks are possible with single factor authentication |  | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Moderate | 0.5 | 5.1 | MEDIUM | 1. Require multi-factor authentication 2. Limit authentication attempts (rate Limiting) 3. Maintain Access Logs 4. Maintain Server Security Logs" | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
36 | T34 | BRUTE force/Dictonary attacks | V07 | Application allowing unlimited login attempts | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. BRUTE force attacks are possible with single factor authentication |  | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Moderate | 0.5 | 5.1 | MEDIUM | 1. Require multi-factor authentication 2. Limit authentication attempts (rate Limiting) 3. Maintain Access Logs 4. Maintain Server Security Logs" | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
37 | T34 | BRUTE force/Dictonary attacks | V36 | Malicious activities not being recorded in administrator/user logs | A11 | THOR Knee Intra-op Application | Automated/robotic attacks are possible if only single factor (user id, password) or with any other login. BRUTE force attacks are possible with single factor authentication |  | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Moderate | 0.5 | 5.1 | MEDIUM | 1. Require multi-factor authentication 2. Limit authentication attempts (rate Limiting) 3. Maintain Access Logs 4. Maintain Server Security Logs" | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
38 | T35 | No session expiry | V08 | No session time out implemented for the application | A11 | THOR Knee Intra-op Application | Every session has to get timed-out after certain period. If this is not enabled, then a continuous session exists without reauthentication and whole objective of authentication is bypassed |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Low | 0.2 | 4.6 | MEDIUM | 1. User configuration settings should be properly validated from security perspective 2. There should be limit/control over the user settings. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
39 | T36 | Unpatched OS/SOUP's | V11 | Unpatched OS | A01 | Nav3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
40 | T36 | Unpatched OS/SOUP's | V11 | Unpatched OS | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
41 | T36 | Unpatched OS/SOUP's | V12 | Potentially outdated SW-Library SOUPs | A01 | Nav3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
42 | T36 | Unpatched OS/SOUP's | V12 | Potentially outdated SW-Library SOUPs | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
43 | T36 | Unpatched OS/SOUP's | V16 | 3rd Party Component/SOUP Risk/Dependency not validated | A01 | Nav3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
44 | T36 | Unpatched OS/SOUP's | V16 | 3rd Party Component/SOUP Risk/Dependency not validated | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
45 | T36 | Unpatched OS/SOUP's | V17 | Legacy Systems (outdated HW/SW) being used | A01 | Nav3i cart/ System running with windows 8.1 | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
46 | T36 | Unpatched OS/SOUP's | V17 | Legacy Systems (outdated HW/SW) being used | A11 | THOR Knee Intra-op Application | Every OS has to get latest security patches based on the vulnerabilities logged into the CVE database & CVSS score. If this is not provided then extreme care has to be taken in protecting the resource (system) |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Moderate | 0.5 | 4.6 | MEDIUM | 1. OS updates should be properly planned for running the Intra-op application 2. Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
47 | T37 | Exploitation of unsecured network elements | V21 | Unprotected network port | A07 | Computer/OS network identification | Adversary uses different kinds of tools to exploit the unused/unprotected network resouces should be scanned and protection should be enabled to safeguard from attacks |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | 1. All the provided physical network components should be identified and unused should be identified 2. Unused componens should be properly disabled/limited 3. System security elements such as firewall, antivirus etc.. should be properly identified | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
48 | T37 | Exploitation of unsecured network elements | V22 | Unprotected external USB Port | A07 | Computer/OS network identification | Adversary uses different kinds of tools to exploit the unused/unprotected network resouces should be scanned and protection should be enabled to safeguard from attacks |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Low | 0.2 | 4.4 | MEDIUM | 1. All the provided physical network components should be identified and unused should be identified 2. Unused componens should be properly disabled/limited 3. System security elements such as firewall, antivirus etc.. should be properly identified | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
49 | T37 | Exploitation of unsecured network elements | V25 | Non-Tracing of unused external Interfaces (ports etc., ) | A07 | Computer/OS network identification | Adversary uses different kinds of tools to exploit the unused/unprotected network resouces should be scanned and protection should be enabled to safeguard from attacks |  | None | Low | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Very Low | 0.04 | 4.3 | MEDIUM | 1. All the provided physical network components should be identified and unused should be identified 2. Unused componens should be properly disabled/limited 3. System security elements such as firewall, antivirus etc.. should be properly identified | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
50 | T38 | Outdated cryptographic components & techniques | V31 | Unencrypted ePHI at rest | A05 | Patient health information at rest | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | 1. Cryptography should be properly maintained for sensitive data 2. Latest crypto algorithms should be selected for ensuring proper security | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
51 | T38 | Outdated cryptographic components & techniques | V31 | Unencrypted ePHI at rest | A11 | THOR Knee Intra-op Application | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | 1. Cryptography should be properly maintained for sensitive data 2. Latest crypto algorithms should be selected for ensuring proper security | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
52 | T38 | Outdated cryptographic components & techniques | V33 | Insecure Storage of Sensitive Information by application | A05 | Patient health information at rest | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | 1. Cryptography should be properly maintained for sensitive data 2. Latest crypto algorithms should be selected for ensuring proper security | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
53 | T38 | Outdated cryptographic components & techniques | V33 | Insecure Storage of Sensitive Information by application | A11 | THOR Knee Intra-op Application | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | 1. Cryptography should be properly maintained for sensitive data 2. Latest crypto algorithms should be selected for ensuring proper security | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
54 | T38 | Outdated cryptographic components & techniques | V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | A05 | Patient health information at rest | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Cryptography should be properly maintained for sensitive data 2. Latest crypto algorithms should be selected for ensuring proper security | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
55 | T38 | Outdated cryptographic components & techniques | V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength in Intra-op application | A11 | THOR Knee Intra-op Application | Encryption/cryptographic techniques are important in securing the sensitive data. Outdated cryptographic algorithms should be avoided to improve the attack complexity, as this can be easily cracked by the adversary |  | High | Low | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Cryptography should be properly maintained for sensitive data 2. Latest crypto algorithms should be selected for ensuring proper security | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
56 | T39 | Unauthorized access/modification of secure data | V37 | Unauthorized Audit Log Manipulation, Log Injection-Tampering-Forging | A11 | THOR Knee Intra-op Application | Audit log data should be maintained to identify security breaches. With no proper authentication to the access of audit logs the content can be modified/deleted. |  | Low | High | Low | Local | Low | High | None | Unchanged | 0.79893 | 0.7323 | 4.7014 | 5.6 | Low | 0.2 | 4.9 | MEDIUM | 1. Audit log data should be maintained for security 2. Access with proper authentication is needed 3. Unauthorized modification/copying should not be allowed | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
57 | T40 | Risk of using windows 8.1 | V51 | Unsupported OS from vendor | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. MDM should provide the details of updating the virus definitions, firewall rules etc… for the cart management. 2. Hospital IT staff should manage the same | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
58 | T40 | Risk of using windows 8.1 | V52 | Unavailability of support from 3rd party tools for outdated OS | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. List down all the 3rd party components and their EOL support 2.Assess the risk & find the replacement for outdated/unsupported 3rd parties based on their dependency | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
59 | T40 | Risk of using windows 8.1 | V53 | OS tools from vendor not getting updated | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. List down all the OS tools/components and their EOL support 2.Assess the risk & find the alternatives for outdated/unsupported vendor tools based on their dependency | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
60 | T40 | Risk of using windows 8.1 | V54 | Absence of periodic backup plan | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Configure the backup plan based on timeline from vendor if available 2. Design & enable the plan for backup if vendor not supporting 3. Ensure a safe and secure backup design done if vendor not supporting | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
61 | T40 | Risk of using windows 8.1 | V55 | Corrupted OS and absence of restore points | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Configure the restore points based on timeline from vendor if available 2. Design & enable the plan for multiple restore points if vendor not supporting 3. Planning for the restore points should be suggested from MDM for the hospital 4. Ensure a safe and secure restore points design if vendor not supporting | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
62 | T40 | Risk of using windows 8.1 | V56 | Remote exploitation of the system from intranet | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Network | Low | Low | None | Unchanged | 2.8353 | 0.6568 | 4.2167 | 7.1 | Moderate | 0.5 | 5.7 | MEDIUM | 1. Firewall rules/other resources should be properly maintained 2. Incoming network traffic should be monitored with the help of tools | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
63 | T40 | Risk of using windows 8.1 | V57 | Any malicious service running in the system without being identified | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Audit logs should be enabled and capture all the user, network etc... activities 2. Periodic monitoring should not be ignored 3. Periodic plan for monitoring should be designed | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
64 | T40 | Risk of using windows 8.1 | V58 | Unhardened OS and network interfaces | A01 | Nav3i cart/ System running with windows 8.1 | EOL OS, outdated tools/components, no patches/updates from vendor, automatic updation of virus definitions, firewall rules not available |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. The vendor provided OS (8.1) should not be used as it is provided in the market 2. Customization of OS is required 3. OS services, frameworks & network interfaces which are not being used can be avoided from running 4. Unnecessary OS services, frameworks & network interfaces should be hardened or has to limited in functionality | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
65 | T41 | Risk of using Linux | V61 | Absence of periodic backup plan | A00 | New cart running with Linux OS | All security packages to be included during OS build. In absence of inbuilt security tools, security breaches can occur. Patch management should be palnned in reference to vulnerable database. If not planned then vulnerabilities can enable the attacker to compromise the system. Hardening checklist to be designed & ensure it has been enforced. If not provided then exploitation of vulnerabilities can make the system unstable |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Configure the backup plan based on timeline from vendor if available 2. Design & enable the plan for backup if vendor not supporting 3. Ensure a safe and secure backup design done if vendor not supporting | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
66 | T41 | Risk of using Linux | V62 | Corrupted OS and absence of restore points | A00 | New cart running with Linux OS | All security packages to be included during OS build. In absence of inbuilt security tools, security breaches can occur. Patch management should be palnned in reference to vulnerable database. If not planned then vulnerabilities can enable the attacker to compromise the system. Hardening checklist to be designed & ensure it has been enforced. If not provided then exploitation of vulnerabilities can make the system unstable |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Configure the restore points based on timeline from vendor if available 2. Design & enable the plan for multiple restore points if vendor not supporting 3. Planning for the restore points should be suggested from MDM for the hospital 4. Ensure a safe and secure restore points design if vendor not supporting | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
67 | T41 | Risk of using Linux | V63 | Remote exploitation of the system from intranet | A00 | New cart running with Linux OS | All security packages to be included during OS build. In absence of inbuilt security tools, security breaches can occur. Patch management should be palnned in reference to vulnerable database. If not planned then vulnerabilities can enable the attacker to compromise the system. Hardening checklist to be designed & ensure it has been enforced. If not provided then exploitation of vulnerabilities can make the system unstable |  | None | Low | High | Network | Low | Low | None | Unchanged | 2.8353 | 0.6568 | 4.2167 | 7.1 | Moderate | 0.5 | 5.7 | MEDIUM | 1. Firewall rules/other resources should be properly maintained 2. Incoming network traffic should be monitored with the help of tools | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
68 | T41 | Risk of using Linux | V64 | Any malicious service running in the system without being identified | A00 | New cart running with Linux OS | All security packages to be included during OS build. In absence of inbuilt security tools, security breaches can occur. Patch management should be palnned in reference to vulnerable database. If not planned then vulnerabilities can enable the attacker to compromise the system. Hardening checklist to be designed & ensure it has been enforced. If not provided then exploitation of vulnerabilities can make the system unstable |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Audit logs should be enabled and capture all the user, network etc... activities 2. Periodic monitoring should not be ignored 3. Periodic plan for monitoring should be designed | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
69 | T41 | Risk of using Linux | V65 | Unhardened OS and network interfaces | A00 | New cart running with Linux OS | All security packages to be included during OS build. In absence of inbuilt security tools, security breaches can occur. Patch management should be palnned in reference to vulnerable database. If not planned then vulnerabilities can enable the attacker to compromise the system. Hardening checklist to be designed & ensure it has been enforced. If not provided then exploitation of vulnerabilities can make the system unstable |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. The vendor provided OS (8.1) should not be used as it is provided in the market 2. Customization of OS is required 3. OS services, frameworks & network interfaces which are not being used can be avoided from running 4. Unnecessary OS services, frameworks & network interfaces should be hardened or has to limited in functionality | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
70 | T41 | Risk of using Linux | V66 | Securtiy packages not properly configured during OS build | A00 | New cart running with Linux OS | All security packages to be included during OS build. In absence of inbuilt security tools, security breaches can occur. Patch management should be palnned in reference to vulnerable database. If not planned then vulnerabilities can enable the attacker to compromise the system. Hardening checklist to be designed & ensure it has been enforced. If not provided then exploitation of vulnerabilities can make the system unstable |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | Packages like ip tables, ep tables, firewall-D & other services that can provide a stronger security environment should be identified/enabled for linux OS | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
71 | T41 | Risk of using Linux | V67 | Critical patches not identified and integrated within timeline | A00 | New cart running with Linux OS | Patch management should be palnned in reference to vulnerable database. If not planned then vulnerabilities can enable the attacker to compromise the system. Hardening checklist to be designed & ensure it has been enforced. If not provided then exploitation of vulnerabilities can make the system unstable |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | Automatic/Manual pulling of the patches for severe vulnerabilities should be planned. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Security Risk Control Measures
Security Risk Control Measures
Security Risk Control Measures
ID | Risk Control Measure | Category (optional) | Reference to more detailed specification
Note on A5: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V11 | Unpatched OS | A01 | Nav3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). | [RARC.RI040] | MEDIUM | Application Whitelisting Firewall Virus Scan User authentication | LOW
2 | T02 | Deliver directed malware (CAPEC-185) | V24 | Any unprotected hardware | A01 | Nav3i cart/ System running with windows 8.1 | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. | [RARC.RI040] | MEDIUM
3 | T10 | Mis-configuration by user | V03 | Untrained/Malicious User | A02 | Admin Password / Credentials / System Configuration / Certificates | If a user misconfigures the system, this might lead to unavailability of the navigation system or make the system less performant for the user. | [RARC.RI110] | LOW | Application Whitelisting Firewall Virus Scan User authentication | MEDIUM | Justification
4 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities | [RARC.RI110] | LOW
5 | T05 | Conduct scavenging of ePHI at rest | V01 | Ineffective management of user credentials | A05 | Patient health information at rest | ePHI stored on Navigation System is exposed to third parties. | [RARC.RI180] | MEDIUM
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Privileged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls