## Header (Optional)
Product Security Risk Table
Ortho Guidance Precision Knee 6.0 Software
Document Control
| Author: | Siva Kumar
| Document No.: | D0000055902
Revision No.:
| Project: | CT Knee Intra-Op | DC-0000002518
| Project Lead: | Akhil Gupta
Change Records
| Rev. No. | Comment | Date | Author
| AA | Initial revision | See onePLM
Note on B19: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## Document References (Optional)
Product Security Risk Table
References
ID | Title | Rev.* | Doc. No.
DI | DI - Flex - Platform (0000013017 - Flex Platform) [DI] | Auto | D0000063236NAV
SRS | SRS - OrthoMap® Precision Knee 6.0 Software - (DC-0000002518 Precision Knee 6.0) [SRS] | Auto | D0000039056
RARC | Risk Assessment - Ortho Guidance Precision Knee 6.0 Software - (DC-0000002518 Precision Knee 6.0) [RARC] | Auto | D0000039062
SRT | SRT - Flex Cart Navigation System - Flex Platform (0000013017 - Flex Platform) [SRT] | Auto | D0000043989
Note on A7: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Stryker Flex Cart Navigation System
Scope: |  | OrthoMap Precision Knee
Date: |  | <yyyyy-mm-dd>
Conducted by: |  | <Author Name / Function / Organization> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description
A01 | Physical Asset | System resources | Contains software module which controls the overall functionality of Knee System i.e. intra-operative planning and guidance.
A02 | Information Asset | Admin Password / Credentials / System Configuration / Certificates | Admin Passwords/Credentials allow modifying config files, accessing ePHI, software binaries, computer configuration files, log files, and computer resources as well as installing applications, security certificates, OS updates.
A03 | Information Asset | User Password / Credentials / System Configuration | User Passwords/Credentials provide access to ePHI, Software Binaries, read access to config files.
A05 | Information asset | Patient health information at rest | Datastored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A07 | Information asset | Computer/OS network identification | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A08 | Information Asset | Intellectual Property | Algorithms and other core-elements of interest to competitors.
A09 | Physical Asset | Navigation Accuracy | Navigation Accuracy of FP8000 Camera, critical components for essential performance
A10 | Information Asset | SSH credentials for system admin access | Only using an 2048bit SSH key allows full access to all system functions. That SSH key is not provided to the customer and can be used by Stryker internals only.
A11 | Information asset | THOR Knee Planning Application | Contains Knee Planning software which controls the overall features for segmentation, landmark planning, implant size calculation etc.
A12 | Information asset | Patient Personal and Health data | DICOM data stored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A13 | Physical Asset | Network Location | Network location where the CT Dicom data is stored
A14 | Physical Asset | input to Intra-Op
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user credentials | Yes | n/a
V02 | Ineffective management of admin credentials | Yes | n/a
V03 | Uneducated/ Malicious User | Yes | n/a
V04 | Insecure Boot / Boot for external media | Yes | n/a
SBOM
V11 | Unpatched OS , Application SW | Yes | n/a
V12 | Potentially outdated SW-Library SOUPs | Yes | n/a
V13 | Developed code not validated using static/Dynamic Tools | Yes | n/a
V14 | UnSafely Compiled Application software | Yes | n/a
V15 | Lack of software vulnerability management | Yes | n/a
V16 | 3rd Party Component/SOUP Dependency not validated | Yes | n/a
V17 | Legacy Systems being used | Yes | n/a
Access points
V21 | Unprotected network port | Yes | n/a
V22 | Unprotected external USB Port | Yes | n/a
V24 | Unprotected hardware | Yes | n/a
V25 | Unused/Unprotected external Interfaces (ports etc., ) | Yes | n/a
Data
V31 | Unencrypted ePHI at rest | Yes | n/a
V32 | Unencrypted ePHI in flight
V33 | Insecure Storage of Sensitive Information
V34 | Use of a Broken or Risky Cryptographic Algorithm or Inadequate encryption strength
V35 | Exposure of Private Information
V36 | Sensitive Data Storage in Improperly Locked Memory
V37 | Lack of fault reports recorded in administrator and operator logs
V38 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging
V39 | Information Exposure Through Log Files
InSecure Configurations of Resources
V40 | Lack of efficient configuration change control | Yes
V41 | Malicious Software Download | Yes
V42 | Malicious Automated Software Update | Yes
V43 | External Control of System or Configuration Setting | Yes
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware | Threat source delivers malware by providing removable media prepared with malware. e.g. Removable media is left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System. Malware exploits a known vulnerability in order to gain access to restricted data or to manipulate data. Undirected attack on computer systems. | TSA-3 - Skript Kiddies, TSA-1 Individual | Yes | n/a
T02 | Deliver directed malware | Threat source delivers malicious software on a removable media which was designed to exploit a known vulnerability of the Navigation System. Directed attack on the Navigation System using knowledge of the Navigation System. | TSA-2 Organization | Yes | n/a
T03 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies and TSA-1 Individual | Yes | n/a
T04 | Intercept network communication | Adversary manipulates the network environment to intercept network communication to gain ePHI, passwords, credit card information etc. or to manipulate DICOM data etc. (man-in-the middle attack) | TSA-2 Organization | No | For Imageless applications, data is never transferred over network.
T05 | Conduct scavenging of ePHI at rest | Adversary gains unauthorized access to the system and steals ePHI information | TSA-1 - Individual | Yes | n/a
T06 | Theft of system or hard drives | A thief steals the navigation system and/or the hard drives as loot, which can lead to unavailability of the navigation system and/or leak of Protected Health Information (PHI). | TSA-1 - Individual | Yes | n/a
T07 | Data theft via physical media | The malicious or inadvertent copy of ePHI on portable media (USB Sticks, DVD drives) from the navigation system. | TSA-2 Organization | Yes | n/a
T08 | Network-based denial of service (DoS) attack | A threat actor seeks to make a machine or network resource unavailable to its intended users by overloading the network resource, the network or target services with malicious data traffic. | TSA-3 - Skript Kiddies TSN-4 Malfunctioning Software / OS TSN-3 Malfunctioning device | Yes | n/a
T09 | Exploit physical access of authorized staff to gain access to organizational facilities. | Adversary follows (“tailgates”) authorized individuals into secure/controlled locations with the goal of gaining access to facilities, circumventing physical security checks. | TSA-1 Individual | No | HDO responsible for physical security controls
T10 | Mis-configuration by user | A user unknowingly or maliciously mis-configures the system so that it malfunctions or no longer functions at all. | TSN-1 Accidental | Yes | n/a
T11 | Remote exploit | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. This is likely inadvertent, targetting general vulnerability in popular operating systems. After successful exploit, continues as malware threat. | TSA-1 Individual | Yes | n/a
T12 | Improper disposal of hard disks | A user disposes the entire system or the hard disks (for instance after a service issue or at the end of the service life of the product). The hard disks still contain patient data. | TSN-1 Accidental | Yes | n/a
T13 | Man-in-the middle attack / intercept Navigation communication | Adversary manipulates the localization communication such as infrared - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T14 | Conduct scavenging of IP at rest | Adversary gains unauthorized access to the system and steals IP off the system | TSA-1 Individual | No | HDO responsible for physical security controls
T15 | Frequency Jamming or Destruction of Wireless Communication | Adversary manipulates the wireless network environment in order to render the service unavailable. | TSA-3 - Skript Kiddies | No | Ortho Imageless application will not be impacted due to unavailability of wireless network
T16 | Physical Manipulation of Hardware | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. | TSA-2 Organization or TSA-1 Individual | Yes | n/a
T17 | Gather information using open source discovery of organizational information | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. | TSA-3 - Skript Kiddies | Yes | n/a
T18 | Craft counterfeit certificates. | Adversary counterfeits or compromises a certificate authority, so that malware or connections will appear legitimate. | TSA-1 Individual | No | Device is a pass through of data in transit from end point devices. The certificate validation is the responsibility of the end points.
T19 | Install persistent and targeted sniffers on organizational information systems and networks. | Adversary places within internal organizational information systems or networks software designed to (over a continuous period of time) collect (sniff) network traffic. | TSA-2 Organization | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T20 | Insert malicious scanning devices (e.g., wireless sniffers) inside facilities. | Adversary uses postal service or other commercial delivery services to deliver to organizational mailrooms a device that is able to scan wireless communications accessible from within the mailrooms and then wirelessly transmit information back to adversary. | TSA-1 Individual | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T21 | Power Failure at primary facility | Power Failure at primary facility makes facility including the navigation sytem inoperable | TSN-7 Power Failure | Yes | n/a
T22 | Theft of SSH key for root access of Navigation Systems | Adverary uses SSH root key to get unauthorized access to the full system. | TSA-1 Individual | No | SSH root key is stored safely at Stryker facility and never delivered to the HDO / customer.
T23 | Manipulation of tool firmware / memory | Adversary manipulates the tool related calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | No | For Imageless applications, no tool is calibrated.
T24 | Manipulation of navigation camera firmware / memory | Adversary manipulates the navgation camera related calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T25 | Manipulation of EM box firmware / memory | Adversary manipulates the EM box firmware - thereby influencing the navigation accuracy | TSA-1 Individual | No | EM is not used by Ortho Imageless applications
T26 | unexpected/Damaged input to CT Knee Intra-op | Input from the Knee planning got modified in transit and unexpected input received by Intra-Op | TSA-1 Individual | Yes
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V11 | Unpatched OS , Application SW | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). | [RARC.RI040] | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Moderate | 0.5 | 4.1 | MEDIUM | Application Whitelisting Firewall Virus Scan User authentication | SRS.SR01 SRS.SR02 SRS.SR03 SRS.SR07 | TestReport #123 | Low | None | None | Network | Low | None | None | Changed | 3.887 | 0.22 | 1.4363 | 5.8 | 3.7 | LOW
2 | T02 | Deliver directed malware | V24 | Unprotected hardware | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. | [RARC.RI040] | Low | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Moderate | 0.5 | 5.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T10 | Mis-configuration by user | V03 | Uneducated/ Malicious User | A02 | Admin Password / Credentials / System Configuration / Certificates | If a user misconfigures the system, this might lead to unavailability of the navigation system or make the system less performant for the user. | [RARC.RI110] | None | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.3916 | 2.5141 | 3.9 | Low | 0.2 | 2.8 | LOW | Application Whitelisting Firewall Virus Scan User authentication | SRS.SR01 SRS.SR02 SRS.SR03 SRS.SR07 | TestReport #123 | High | High | Low | Physical | High | High | Required | Unchanged | 0.12109 | 0.84899 | 5.4505 | 5.6 | 5.5 | MEDIUM | Justification
4 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities | [RARC.RI110] | Low | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.3916 | 2.5141 | 4.8 | Moderate | 0.5 | 3.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T05 | Conduct scavenging of ePHI at rest | V01 | Ineffective management of user credentials | A05 | Patient health information at rest | ePHI stored on Navigation System is exposed to third parties. | [RARC.RI180] | High | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Low | 0.2 | 4.6 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T02 | Deliver directed malware | v02 | Ineffective management of admin credentials | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs malware in order to gain admin credentials | [RARC.RI040] | Low | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Moderate | 0.5 | 5.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T13 | Man-in-the middle attack / intercept Navigation communication | V24 | Unprotected hardware | A09 | Navigation Accuracy | Man-in-middle attack on localization communication such as infrared - Adversary manipulates navigation accuracy | [RARC.RI180] | Low | High | High | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Very Low | 0.04 | 5.5 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T08 | Network-based denial of service (DoS) attack | V21 | Unprotected network port | A01 | System resources | The navigation system is attacked using a DoS attack. In the worst case the system is unavailable to the user. | [RARC.RI110] | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Moderate | 0.5 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T11 | Remote exploit | V21 | Unprotected network port | A09 | Navigation Accuracy | A threat actor uses the hospital network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system in order to manipulate navigation accuracy. | [RARC.RI180] | None | High | Low | Network | High | Low | None | Changed | 1.7769 | 0.6568 | 4.7173 | 7.1 | Low | 0.2 | 5.5 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T11 | Remote exploit | V21 | Unprotected network port | A01 | System resources | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. In the worst case the system is unavailable to the user. | [RARC.RI110] | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Moderate | 0.5 | 5.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T11 | Remote exploit | V11 | Unpatched OS , Application SW | A02 | Admin Password / Credentials / System Configuration / Certificates | If a remote exploit (e.g. "worm") is able to circumvent the operating system defenses to gain local access, then several other attack scenarios are enabled. In the worst case the system is unavailable to the user. | [RARC.RI110] | High | High | Low | Network | Low | High | None | Unchanged | 1.2347 | 0.84899 | 5.4505 | 6.7 | Moderate | 0.5 | 6.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T12 | Improper disposal of hard disks | V24 | Unprotected hardware | A05 | Patient health information at rest | ePHI might be exposed to third parties if the storage medium of the navigation unit is improperly disposed of. | N/A | High | None | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T06 | Theft of system or hard drives | V24 | Unprotected hardware | A01 | System resources | When parts of the Navigation system are stolen the navigation system might be rendered unavailable for use. | N/A | High | none | High | Physical | low | low | none | unchanged | 0.66712 | 0.8064 | 5.1771 | 5.9 | Low | 0.2 | 5.4 | MEDIUM | Encrpyted ePHI in flight | SRS.SR04 | TestReport #789 | Low | None | None | Network | High | None | None | Unchanged | 2.2212 | 0.22 | 1.4124 | 3.7 | 1.9 | LOW
14 | T07 | Data theft via physical media | V22 | Unprotected external USB Port | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by inserting physical media such as USB Sticks or External Drives. | N/A | High | none | none | Physical | low | low | none | unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T16 | Physical Manipulation of Hardware | V24 | Unprotected hardware | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. Once defenses are circumvented in order to gain local access, several other Assets such as user credentials or PHI may be threatened. | [RARC.RI180] | Low | High | none | Physical | Low | High | none | Unchanged | 0.29052 | 0.6568 | 4.2167 | 4.6 | Low | 0.2 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T21 | Power Failure at primary facility | V24 | Unprotected hardware | A01 | System resources | Adversary attacks power supply of the hospital or OR member cuts of power supply of navigation system by acccident. In both cases navigation system goes off to missing power. | [RARC.RI110] | None | None | High | Local | Low | None | None | Unchanged | 2.5151 | 0.56 | 3.5952 | 6.2 | Very Low | 0.04 | 3.7 | LOW | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T17 | Gather information using open source discovery of organizational information | V21 | Unprotected network port | A01 | System resources | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. That information is used for an attack via network. | [RARC.RI040] | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Low | 0.2 | 5.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T01 | Deliver undirected malware | V13 | Developed code not validated using static/Dynamic Tools | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) due to discovered defect or bug in Stryker software. | [RARC.RI040] | Low | Low | High | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | Very Low | 0.04 | 4.8 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T05 | Conduct scavenging of ePHI at rest | V04 | Insecure Boot / Boot for external media | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by booting the system with an manipulated / self owned operating system to access the harddrive data and bypass the user role system provided by the original Stryker operating system. | [RARC.RI180] | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Low | 0.2 | 5.1 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T24 | Manipulation of navigation camera firmware / memory | V24 | Unprotected hardware | A09 | Navigation Accuracy | The threat actor manipulates the memory / firmware of navigation camera with the goal to bring the calibration off / manipulate the firmware to reduce navigation accuracy. | [RARC.RI180] | High | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Low | 0.2 | 5 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
21 | T16 | Physical Manipulation of Hardware | V24 | Unprotected hardware | A09 | Navigation Accuracy | Adversary with direct mal-intent gains local access to the navigation flex cart platform and manipulates hardware such as tools and navigation camera. | [RARC.RI180] | Low | High | none | Physical | Low | High | none | Unchanged | 0.29052 | 0.6568 | 4.2167 | 4.6 | Low | 0.2 | 4.3 | MEDIUM | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Security Risk Control Measures
Security Risk Control Measures
Security Risk Control Measures
ID | Risk Control Measure | Category (optional) | Reference to more detailed specification
Note on A5: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V11 | Unpatched OS , Application SW | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). | [RARC.RI040] | MEDIUM | Application Whitelisting Firewall Virus Scan User authentication | LOW
2 | T02 | Deliver directed malware | V24 | Unprotected hardware | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. | [RARC.RI040] | MEDIUM
3 | T10 | Mis-configuration by user | V03 | Uneducated/ Malicious User | A02 | Admin Password / Credentials / System Configuration / Certificates | If a user misconfigures the system, this might lead to unavailability of the navigation system or make the system less performant for the user. | [RARC.RI110] | LOW | Application Whitelisting Firewall Virus Scan User authentication | MEDIUM | Justification
4 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities | [RARC.RI110] | LOW
5 | T05 | Conduct scavenging of ePHI at rest | V01 | Ineffective management of user credentials | A05 | Patient health information at rest | ePHI stored on Navigation System is exposed to third parties. | [RARC.RI180] | MEDIUM
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Privileged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls