## System & Asset Identification
System & Asset Identification
Medical Device / System:
Scope:
Date: |  | <2021-07-12>
Conducted by: |  | <Author Name / Function / Organization> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description
A01 | Physical Asset | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | Utilizing computer resources and computing power by adversary, allows various general purpose attacks, such as incl. Ransomware deployment, Bitcoin Mining, abuse of peripheral devices such as WebCam, Microphones, etc., .
A02 | Information asset | Tablet OS/network details & Tablet Application | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A03 | Physical Assets | Smart medic (Stryker device) System Component | Monitors local bed status information, alerting caregivers visually, audibly or remotely if preset parameters are compromised.
A04 | Information asset | Authenication/Authorisation data | Information related to authenication/authorisation data (password/pins/MFA/Biometrics)
A05 | Physical Assets | Device Maintainence tool (Hardware/Software) | Device Maintainence tool (Hardware/Software) that patchs and updates Smart Medic Device and Application related to Security
A06 | Information asset | Electronic Health Records (EHR)/ Device Component status | Smart device components health status information
A07 | Information asset | Interface/API Communication | Communication middleware enables communication and data management for distributed applications.
A08 | Physical Assets | Wireless Network device | Devices that are used for communication among the Smart Medic project component.
A09 | Information asset | Data at Rest | Use strong encryption algorthim to store data on cloud platform (Smartmedic Device)/tablet
A10 | Information asset | Data in Motion | Use strong encryption algorthim to data moving on tablet to cloud platform(Smartmedic Device)/tablet
A11 | Information asset | Smart medic app (Stryker Azure Cloud Web Application) | Smart medic application for nurse/health worker (Stryker Azure Cloud Web Application)
A12 | Information asset | Smart medic app (Azure Portal Administrator) | Azure Portal Administrator for Smart medic app
A13 | Information asset | Azure Cloud DataBase | Azure Cloud DataBase related to Smart Medic app
A14 | Information asset | Health vital data | Health vital data Body temperature. Pulse rate. Respiration rate,weight data, position data, etc.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
D05788-1, Ver 1
Note on D19: P Muni Sai Sivakumar: moving from tablet

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
AuthN,AuthZ management
V01 | Devices with default passwords needs to be checked for bruteforce attacks | Yes | n/a
V02 | Lack of Strong Guidelines Password Policy
V03 | Error Info containing sensitive data for Failed Authentication attempts
V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | Yes | n/a
V05 | Checking authentication modes for possible hacks and bypasses | Yes | n/a
V06 | Absence of Multiple layer Authentication (MFA)
V07 | Lack of Authorization policies & Management
SOM
V07 | Lack of Asset location digaram in security operations manual | Yes | n/a
V08 | Lack of configuration controls for IT assets in the informaion system plan | Yes | n/a
V09 | Ineffective patch management of firware, OS and applications thoughout the information system plan | Yes | n/a
V10 | Lack of plan for periodic Software Vulnerability Management | Yes | n/a
V11 | The static connection digaram between devices and applications with provision for periodic updation as per changes | Yes | n/a
V12 | Assest counting system for all instances of product implementation | Yes | n/a
Access points
V12 | Unprotected network port(s) on network devices and connection points | Yes | n/a
V13 | Unprotected external USB Port on the tablet/devices. | Yes | n/a
V14 | Unencrypted Network segment throught the information flow | Yes | n/a
V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | Yes | n/a
V16 | Controlled Use of Administrative Privileges over the network | Yes | n/a
Data
V17 | Unencrypted data at rest in all possible locations | Yes | n/a
V18 | Unencrypted data in flight in all flowchannels | Yes | n/a
V19 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | Yes | n/a
V20 | Non-Standard/outdated algorithms with insufficient crypto elements
V21 | Weak Algorthim implementation with respect cipher key size | Yes | n/a
InSecure Configurations of Resources
V21 | InSecure/not recommended Configuration for Mobile Devices, Laptops, Workstations, and Servers | Yes | n/a
V22 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | Yes | n/a
V23 | Legacy system identification if any | Yes | n/a
V24 | Outdated - Software/Hardware | Yes | n/a
V25 | Insecure communications in networks (hospital) | Yes | n/a
SBOM/software
V27 | Presence of Re-used/Outdated components in developed software | Yes
V28 | UnProtected/Unsafely Compiled Application Binaries | Yes | n/a
V29 | Irregular patch updates | Yes
Logging/Monitoring
V30 | Insufficient Logging information
V31 | Insufficient Access permissions for accessing Log files
Keys & Certificates
V32 | Generation of keys without using a FIPS certified /Standard RNG
V33 | Generation and Signing of certificates without standards
V34 | Improper security (for ex.,Storage & Access) for Keys and Certificates
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
Note on B28: P Muni Sai Sivakumar: Data in Transit

## Threat Assessment
Threat Assessment
# | Threat Category | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | CAPEC-185 | Deliver undirected malware | Thread source delivers malware by providing removable media prepared with malware. Removable media is e.g. left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System. Malware eploits known a known vulnerability and e.g. gains admin priviledges. Undirected attack on computer systems. | TSA-3 - Skript Kiddies | Yes | n/a
T02 | CAPEC-185 | Deliver directed malware | Thread source delivers malware on a removable media which was designed to exploit a known vulnerability of the Navigation System. Directed attack on the Navigation System using knowledge about the Navigation System. | TSA-2 Organization | Yes | n/a
T03 | [S]TRID[E] | Gaining Access | This phase is where an attacker breaks into the system/network using various tools or methods. After entering into a system, he has to increase his privilege to administrator level so he can install an application he needs or modify data or hide data | TSA-2 Organization | Yes | n/a
T04 | TTP | Maintaining Access | The aim is to maintain the access to the target until he finishes the tasks he planned to accomplish in that target. | TSA-2 Organization | Yes | n/a
T05 | TTP | Clearing Track | This involves modifying/corrupting/deleting the values of Logs, modifying registry values and uninstalling all applications he used and deleting all folders he created | TSA-2 Organization | Yes | n/a
T06 | STRID[E] | Elevation of privilege | Identify weaknesses of segregation in terms of administrative and user-level privileges | TSA-2 Organization | Yes | n/a
T07 | STRI(D)E | Denial of service | Find ways to exhaust or drown out legitimate requests | TSA-3 - Skript Kiddies | Yes | n/a
T08 | STR(I)DE | Information disclosure | Fuzz application parameters or arguments to impact application error disclosures. Identify open ports with their respective services. Incite confidentiality and integrity in the browser interface. Identify clear text communications. Review usage of HTTP headers and user-agent profile. Pinpoint usages of API endpoints and application backend technologies. | TSA-2 Organization | Yes | n/a
T09 | STR[I]DE | Data Access | Access user and application data e.g. by a malicious application or script | TSA-3 - Skript Kiddies | Yes | n/a
T10 | TTP | Open network port exploit | Penetrate Open and Unsecured Ports | TSA-3 - Skript Kiddies | Yes | n/a
T11 | CAPEC-112 | Brute-force Attack | The brute-force attack contained a dictionary of well-known directories and authentication paradigms present in common webservers. | TSA-2 Organization | Yes | n/a
T12 | TTP | Social Engineering | create custom phishing scams, phone-based attacks and even physical intrusion to test your organization’s level of security. | TSA-3 - Skript Kiddies | Yes | n/a
T13 | ST[R]IDE | Data logs with insufficient access permissions | Data logs are generally exploited through untrusted and unknown access permissions. | TSA-2 Organization | Yes | n/a
T14 | TTP | Perform Network scanning | Perform Perimeter network reconnaisance / Network Sniffing | TSA-3 - Skript Kiddies | Yes | n/a
T15 | S[T]RIDE | Unauthorized Alterations | This involves modifying registry values, deleting/encrypting Confidential info and uninstalling Any secure applications and renaming/deleting all files/folders | TSA-2 Organization | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
Session invalidation after certain time-interval
Note on D13: P Muni Sai Sivakumar: Un-Used Ports

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  |  | Low | None | None | Network | Low | None | None | Changed | 3.887 | 0.22 | 1.4363 | 5.8 | 2.4 | LOW
2 | T01 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T01 | CAPEC-185 | V02 | Lack of Strong Guidelines Password Policy | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
4 | T01 | CAPEC-185 | V02 | Lack of Strong Guidelines Password Policy | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T01 | CAPEC-185 | V22 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T01 | CAPEC-185 | V22 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T01 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T01 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T01 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T01 | CAPEC-185 | V09 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T01 | CAPEC-185 | V09 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T01 | CAPEC-185 | V09 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T01 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
14 | T01 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | High | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T01 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T01 | CAPEC-185 | V17 | Unencrypted data at rest in all possible locations | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T01 | CAPEC-185 | V17 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T01 | CAPEC-185 | V23 | Legacy system identification if any | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T01 | CAPEC-185 | V23 | Legacy system identification if any | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T01 | CAPEC-185 | V23 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
21 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | High | 0.8 | 4.5 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  |  | High | High | Low | Physical | High | High | Required | Unchanged | 0.12109 | 0.84899 | 5.4505 | 5.6 | 5.6 | MEDIUM | Justification
22 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.84899 | 5.4505 | 6.8 | High | 0.8 | 6.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
23 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
24 | T02 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Moderate | 0.5 | 3.7 | LOW | • Stateful Firewall • Maintain access control list | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
25 | T02 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | High | 0.8 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
26 | T02 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | • Use Secure tunnel Communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
27 | T02 | CAPEC-185 | V02 | Lack of Strong Guidelines Password Policy | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | High | 0.8 | 5.1 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
28 | T02 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
29 | T02 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
30 | T02 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
31 | T02 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.84899 | 5.4505 | 8.3 | Moderate | 0.5 | 6.9 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
32 | T02 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.84899 | 5.4505 | 8.3 | Low | 0.2 | 6.1 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
33 | T02 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.84899 | 5.4505 | 8.3 | Low | 0.2 | 6.1 | MEDIUM | • Disable the device network discoverable • Maintain Access Control List. • Block all unrequired ports • Stateful firewall | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
34 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | Low | Low | Local | High | Low | Required | Unchanged | 0.76466 | 0.7323 | 4.7014 | 5.5 | High | 0.8 | 5.4 | MEDIUM | • Implement automated configuration monitoring systems • Use Secure tunnel Communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
35 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
36 | T02 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  |  | Low | None | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.3916 | 2.5141 | 3.2 | 2.7 | LOW
37 | T02 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A02 | Tablet OS/network details & Tablet Application | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
38 | T02 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | • Encrypting the storage subsystem • Encryption methods such as HTTPS, SSL, and TLS are often used to protect data in motion. • Stateful firewall • Maintain access control list. • Use strong encrption algorithm | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
39 | T03 | [S]TRID[E] | V12 | Assest counting system for all instances of product implementation | A02 | Tablet OS/network details & Tablet Application | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
40 | T03 | [S]TRID[E] | V12 | Assest counting system for all instances of product implementation | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | High | 0.8 | 7 | HIGH | • Configuring account lockout policies cannot be exploited to lock out well known service accounts. • Ensuring application is capable of handling required volumes of traffic and that thresholds are in place to handle optimal loads. • Review your application's failover functionality. • Maintain Access Control List. • Block all unrequired ports • Stateful firewall | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
41 | T03 | [S]TRID[E] | V12 | Assest counting system for all instances of product implementation | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  |  | Low | None | None | Network | High | None | None | Unchanged | 2.2212 | 0.22 | 1.4124 | 3.7 | 1.9 | LOW
42 | T03 | [S]TRID[E] | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
43 | T03 | [S]TRID[E] | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.7323 | 4.7014 | 5.4 | High | 0.8 | 5.3 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
44 | T03 | [S]TRID[E] | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A07 | Interface/API Communication | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
45 | T03 | [S]TRID[E] | V03 | Error Info containing sensitive data for Failed Authentication attempts | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
46 | T03 | [S]TRID[E] | V03 | Error Info containing sensitive data for Failed Authentication attempts | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | High | 0.8 | 5.8 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
47 | T03 | [S]TRID[E] | V03 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | High | 0.8 | 5.8 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
48 | T03 | [S]TRID[E] | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
49 | T03 | [S]TRID[E] | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
50 | T03 | [S]TRID[E] | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
51 | T03 | [S]TRID[E] | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
52 | T03 | [S]TRID[E] | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi-factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
53 | T03 | [S]TRID[E] | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
54 | T04 | TTP | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.7323 | 4.7014 | 5.4 | High | 0.8 | 5.3 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
55 | T04 | TTP | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.7323 | 4.7014 | 5.4 | High | 0.8 | 5.3 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
56 | T04 | TTP | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
57 | T04 | TTP | V03 | Error Info containing sensitive data for Failed Authentication attempts | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
58 | T04 | TTP | V03 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
59 | T04 | TTP | V03 | Error Info containing sensitive data for Failed Authentication attempts | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
60 | T04 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
61 | T04 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
62 | T04 | TTP | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | High | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.84899 | 5.4505 | 7.6 | Moderate | 0.5 | 6.5 | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
63 | T04 | TTP | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
64 | T05 | TTP | V21 | Weak Algorthim implementation with respect cipher key size | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
65 | T05 | TTP | V21 | Weak Algorthim implementation with respect cipher key size | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | High | Low | Required | Unchanged | 0.76466 | 0.52545 | 3.3734 | 4.2 | Low | 0.2 | 3.6 | LOW | • Establish secure configuration • Deploy system configuration management tool • Implement automated configuration monitoring systems • Establish internal and external information sources for threat intelligence and vulnerability data, monitoring them regularly and taking appropriate action for high-priority items • Upgrades the software, firmware • Never use credentials such as date of birth, spouse, or child’s or pet’s name • Stateful Firewall | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
66 | T05 | TTP | V23 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
67 | T05 | TTP | V07 | Lack of Authorization policies & Management | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
68 | T05 | TTP | V07 | Lack of Authorization policies & Management | A05 | Device Maintainence tool (Hardware/Software) |  |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
69 | T05 | TTP | V08 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
70 | T05 | TTP | V08 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
71 | T05 | TTP | V08 | Lack of configuration controls for IT assets in the informaion system plan | A02 | Tablet OS/network details & Tablet Application | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
72 | T05 | TTP | V10 | Lack of plan for periodic Software Vulnerability Management | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
73 | T05 | TTP | V10 | Lack of plan for periodic Software Vulnerability Management | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
74 | T06 | STRID[E] | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A04 | Authenication/Authorisation data | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
75 | T06 | STRID[E] | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A12 | Smart medic app (Azure Portal Administrator) | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | High | High | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.84899 | 5.4505 | 7.6 | High | 0.8 | 7.2 | HIGH | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
76 | T07 | STRI(D)E | V12 | Assest counting system for all instances of product implementation | A08 | Wireless Network device | 1) Bring down the service availability 2) Blocking the end user usage |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | • Eliminating rogue access points • Properly configuring all authorized access points • Firewall • Limit Network Access | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
77 | T07 | STRI(D)E | V12 | Assest counting system for all instances of product implementation | A02 | Tablet OS/network details & Tablet Application | 1) Bring down the service availability 2) Blocking the end user usage |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
78 | T08 | STR(I)DE | V16 | Controlled Use of Administrative Privileges over the network | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.84899 | 5.4505 | 6.5 | Moderate | 0.5 | 6 | MEDIUM | • Encrypting the storage subsystem • Encryption methods such as HTTPS, SSL, and TLS are often used to protect data in motion. • Stateful firewall • Host Hardening • Maintain access control list. • Use strong encrption algorithm | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
79 | T08 | STR(I)DE | V17 | Unencrypted data at rest in all possible locations | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Use secure tunnel communication channel • Maintain access control list. • Use strong encrption algorithm | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
80 | T08 | STR(I)DE | V18 | Unencrypted data in flight in all flowchannels | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.84899 | 5.4505 | 6.5 | Moderate | 0.5 | 6 | MEDIUM | • Server-side encryption using Service-Managed keys • Network access control • Encryption for sensitive data in transit, such as when files are attached to an email message or moved to cloud storage, removable drives • Transfer over encrypted tunnel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
81 | T08 | STR(I)DE | V18 | Unencrypted data in flight in all flowchannels | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Server-side encryption using Service-Managed keys • Use secure tunnel communication channel • Network access control • Encryption for sensitive data in transit, such as when files are attached to an email message or moved to cloud storage, removable drives • Transfer over encrypted tunnel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
82 | T08 | STR(I)DE | V19 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.84899 | 5.4505 | 6.5 | Moderate | 0.5 | 6 | MEDIUM | • Implement strong encryption descrption method at data at rest • Use secure tunnel communication channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
83 | T08 | STR(I)DE | V19 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Implement strong sysmetric or unsysmetric algorhythn • Implementing VPN and stateful Firewall | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
84 | T08 | STR(I)DE | V21 | Weak Algorthim implementation with respect cipher key size | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
85 | T08 | STR(I)DE | V14 | Unencrypted Network segment throught the information flow | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Anonymization/Pseudomyzation • Encyrption • Mainitaing Access Logs • Maintain Server Security Logs | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
86 | T08 | STR(I)DE | V05 | Checking authentication modes for possible hacks and bypasses | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Use strong encryption. • Secure communication links with protocols that provide message confidentiality. • Secure sensitive data in the channel flow • Protect authentication cookies with Secure Sockets Layer (SSL) or TLS. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
87 | T09 | STR[I]DE | V12 | Assest counting system for all instances of product implementation | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
88 | T09 | STR[I]DE | V12 | Assest counting system for all instances of product implementation | A08 | Wireless Network device | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Firewall • Secure communication links with protocols that provide message confidentiality. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
89 | T09 | STR[I]DE | V12 | Assest counting system for all instances of product implementation | A02 | Tablet OS/network details & Tablet Application | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
90 | T09 | STR[I]DE | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A09 | Data at Rest | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Firewall • Do not store sensitive data in plaintext. • Use strong encrption algorithm. • Apply salting over sensitive data. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
91 | T09 | STR[I]DE | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | High | 0.8 | 6.1 | MEDIUM | • Firewall • Do not store sensitive data in plaintext. • Use strong encrption algorithm. • Apply salting over sensitive data. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
92 | T09 | STR[I]DE | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A10 | Data in Motion | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Firewall • Do not store sensitive data in plaintext. • Use strong encrption algorithm. • Apply salting over sensitive data. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
93 | T09 | STR[I]DE | V03 | Error Info containing sensitive data for Failed Authentication attempts | A09 | Data at Rest | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Stronger authentication methods • Use strong encrption algorithm. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
94 | T09 | STR[I]DE | V03 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Stronger authentication methods • Use strong encrption algorithm. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
95 | T09 | STR[I]DE | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A12 | Smart medic app (Azure Portal Administrator) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
96 | T09 | STR[I]DE | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Physical | High | High | None | Unchanged | 0.16601 | 0.84899 | 5.4505 | 5.7 | High | 0.8 | 5.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
97 | T10 | TTP | V12 | Assest counting system for all instances of product implementation | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
98 | T10 | TTP | V12 | Assest counting system for all instances of product implementation | A08 | Wireless Network device | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Using port security on Firewall • Application level Firewall • Check ing internet provider’s or router manufacturer’s wireless security options. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
99 | T10 | TTP | V14 | Unencrypted Network segment throught the information flow | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
100 | T10 | TTP | V14 | Unencrypted Network segment throught the information flow | A08 | Wireless Network device | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • IP Whitelisting • Use strong encryption algorithyms | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
101 | T10 | TTP | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A12 | Smart medic app (Azure Portal Administrator) | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
102 | T10 | TTP | V17 | Unencrypted data at rest in all possible locations | A10 | Data in Motion | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Use secure tunnel communication channel • Configure and upgrade routers for the n/w security • Configure firewalls to reject any packets with spoofed addresses. • Maintain access control list. • Use strong encrption algorithm | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
103 | T10 | TTP | V05 | Checking authentication modes for possible hacks and bypasses | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
104 | T11 | CAPEC-112 | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
105 | T11 | CAPEC-112 | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
106 | T11 | CAPEC-112 | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
107 | T11 | CAPEC-112 | V03 | Error Info containing sensitive data for Failed Authentication attempts | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
108 | T11 | CAPEC-112 | V03 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
109 | T11 | CAPEC-112 | V03 | Error Info containing sensitive data for Failed Authentication attempts | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
110 | T11 | CAPEC-112 | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Set your spam filters setting options to high • Secure your computing devices | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
111 | T11 | CAPEC-112 | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Set your spam filters setting options to high • Secure your computing devices | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
112 | T11 | CAPEC-112 | V18 | Unencrypted data in flight in all flowchannels | A09 | Data at Rest | 1) An attacker may attempt to discover a weak encryption by systematically trying every possible combination of decryption key. |  | High | High | Low | Local | High | High | None | Unchanged | 0.45653 | 0.84899 | 5.4505 | 6 | Moderate | 0.5 | 5.7 | MEDIUM | • Server-side encryption using Service-Managed keys • Network access control • Encryption for sensitive data in transit, such as when files are attached to an email message or moved to cloud storage, removable drives • Transfer over encrypted tunnel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
113 | T11 | CAPEC-112 | V18 | Unencrypted data in flight in all flowchannels | A10 | Data in Motion | 1) An attacker may attempt to discover a weak encryption by systematically trying every possible combination of decryption key. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Firewall • Encrypting data at rest • Protect your data with multi factor authentication • Use secure tunnel communication channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
114 | T12 | TTP | V21 | Weak Algorthim implementation with respect cipher key size | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Implement automated configuration monitoring systems • Stateful firewall | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
115 | T12 | TTP | V22 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Delete any request for personal information • Firewall • Disable device network discoverable • Maintain access control list | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
116 | T12 | TTP | V03 | Error Info containing sensitive data for Failed Authentication attempts | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Stronger authentication methods | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
117 | T12 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A07 | Interface/API Communication | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Set your spam filters setting options to high • Firewalls, • Use secure tunnel communication channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
118 | T12 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Set your spam filters setting options to high • Firewalls, • Use Multi Factor Authentication • Use secure tunnel communication channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
119 | T12 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A13 | Azure Cloud DataBase | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Set your spam filters setting options to high • Firewalls, • Use Multi Factor Authentication • Use secure tunnel communication channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
120 | T12 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Firewall • Use Multi Factor Authentication • Use secure tunnel communication channel | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | LOW
2 | T01 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
3 | T01 | CAPEC-185 | V02 | Lack of Strong Guidelines Password Policy | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
4 | T01 | CAPEC-185 | V02 | Lack of Strong Guidelines Password Policy | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
5 | T01 | CAPEC-185 | V22 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
6 | T01 | CAPEC-185 | V22 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
7 | T01 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
8 | T01 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
9 | T01 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
10 | T01 | CAPEC-185 | V09 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
11 | T01 | CAPEC-185 | V09 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
12 | T01 | CAPEC-185 | V09 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
13 | T01 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
14 | T01 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
15 | T01 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
16 | T01 | CAPEC-185 | V17 | Unencrypted data at rest in all possible locations | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
17 | T01 | CAPEC-185 | V17 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
18 | T01 | CAPEC-185 | V23 | Legacy system identification if any | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
19 | T01 | CAPEC-185 | V23 | Legacy system identification if any | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
20 | T01 | CAPEC-185 | V23 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
21 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | MEDIUM | Justification
22 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
23 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
24 | T02 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | • Stateful Firewall • Maintain access control list
25 | T02 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
26 | T02 | CAPEC-185 | V13 | Unprotected external USB Port on the tablet/devices. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | • Use Secure tunnel Communications channel
27 | T02 | CAPEC-185 | V02 | Lack of Strong Guidelines Password Policy | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
28 | T02 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
29 | T02 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
30 | T02 | CAPEC-185 | V08 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
31 | T02 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
32 | T02 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
33 | T02 | CAPEC-185 | V12 | Assest counting system for all instances of product implementation | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | • Disable the device network discoverable • Maintain Access Control List. • Block all unrequired ports • Stateful firewall
34 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | • Implement automated configuration monitoring systems • Use Secure tunnel Communications channel
35 | T02 | CAPEC-185 | V21 | Weak Algorthim implementation with respect cipher key size | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
36 | T02 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | LOW
37 | T02 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A02 | Tablet OS/network details & Tablet Application | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
38 | T02 | CAPEC-185 | V16 | Controlled Use of Administrative Privileges over the network | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | • Encrypting the storage subsystem • Encryption methods such as HTTPS, SSL, and TLS are often used to protect data in motion. • Stateful firewall • Maintain access control list. • Use strong encrption algorithm
39 | T03 | [S]TRID[E] | V12 | Assest counting system for all instances of product implementation | A02 | Tablet OS/network details & Tablet Application | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
40 | T03 | [S]TRID[E] | V12 | Assest counting system for all instances of product implementation | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | HIGH | • Configuring account lockout policies cannot be exploited to lock out well known service accounts. • Ensuring application is capable of handling required volumes of traffic and that thresholds are in place to handle optimal loads. • Review your application's failover functionality. • Maintain Access Control List. • Block all unrequired ports • Stateful firewall
41 | T03 | [S]TRID[E] | V12 | Assest counting system for all instances of product implementation | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | LOW
42 | T03 | [S]TRID[E] | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs
43 | T03 | [S]TRID[E] | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
44 | T03 | [S]TRID[E] | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A07 | Interface/API Communication | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
45 | T03 | [S]TRID[E] | V03 | Error Info containing sensitive data for Failed Authentication attempts | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
46 | T03 | [S]TRID[E] | V03 | Error Info containing sensitive data for Failed Authentication attempts | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
47 | T03 | [S]TRID[E] | V03 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
48 | T03 | [S]TRID[E] | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
49 | T03 | [S]TRID[E] | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
50 | T03 | [S]TRID[E] | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
51 | T03 | [S]TRID[E] | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
52 | T03 | [S]TRID[E] | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require that administrators establish multi-factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
53 | T03 | [S]TRID[E] | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
54 | T04 | TTP | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
55 | T04 | TTP | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
56 | T04 | TTP | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
57 | T04 | TTP | V03 | Error Info containing sensitive data for Failed Authentication attempts | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
58 | T04 | TTP | V03 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
59 | T04 | TTP | V03 | Error Info containing sensitive data for Failed Authentication attempts | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
60 | T04 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses.
61 | T04 | TTP | V04 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses.
62 | T04 | TTP | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
63 | T04 | TTP | V15 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
64 | T05 | TTP | V21 | Weak Algorthim implementation with respect cipher key size | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
65 | T05 | TTP | V21 | Weak Algorthim implementation with respect cipher key size | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | • Establish secure configuration • Deploy system configuration management tool • Implement automated configuration monitoring systems • Establish internal and external information sources for threat intelligence and vulnerability data, monitoring them regularly and taking appropriate action for high-priority items • Upgrades the software, firmware • Never use credentials such as date of birth, spouse, or child’s or pet’s name • Stateful Firewall
66 | T05 | TTP | V23 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
67 | T05 | TTP | V07 | Lack of Authorization policies & Management | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
68 | T05 | TTP | V07 | Lack of Authorization policies & Management | A05 | Device Maintainence tool (Hardware/Software) |  |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
69 | T05 | TTP | V08 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
70 | T05 | TTP | V08 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
71 | T05 | TTP | V08 | Lack of configuration controls for IT assets in the informaion system plan | A02 | Tablet OS/network details & Tablet Application | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
72 | T05 | TTP | V10 | Lack of plan for periodic Software Vulnerability Management | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls