## System & Asset Identification
System & Asset Identification
Medical Device / System:
Scope:
Date: |  | <2021-07-12>
Conducted by: |  | <Author Name / Function / Organization> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description
A01 | Physical Asset | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | Utilizing computer resources and computing power by adversary, allows various general purpose attacks, such as incl. Ransomware deployment, Bitcoin Mining, abuse of peripheral devices such as WebCam, Microphones, etc., .
A02 | Information asset | Tablet OS/network details & Tablet Application | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A03 | Physical Assets | Smart medic (Stryker device) System Component | Monitors local bed status information, alerting caregivers visually, audibly or remotely if preset parameters are compromised.
A04 | Information asset | Authenication/Authorisation data | Information related to authenication/authorisation data (password/pins/MFA/Biometrics)
A05 | Physical Assets | Device Maintainence tool (Hardware/Software) | Device Maintainence tool (Hardware/Software) that patchs and updates Smart Medic Device and Application related to Security
A06 | Information asset | Electronic Health Records (EHR)/ Device Component status | Smart device components health status information
A07 | Information asset | Interface/API Communication | Communication middleware enables communication and data management for distributed applications.
A08 | Physical Assets | Wireless Network device | Devices that are used for communication among the Smart Medic project component.
A09 | Information asset | Data at Rest | Use strong encryption algorthim to store data on cloud platform (Smartmedic Device)/tablet
A10 | Information asset | Data in Motion | Use strong encryption algorthim to data moving on tablet to cloud platform(Smartmedic Device)/tablet
A11 | Information asset | Smart medic app (Stryker Azure Cloud Web Application) | Smart medic application for nurse/health worker (Stryker Azure Cloud Web Application)
A12 | Information asset | Smart medic app (Azure Portal Administrator) | Azure Portal Administrator for Smart medic app
A13 | Information asset | Azure Cloud DataBase | Azure Cloud DataBase related to Smart Medic app
A14 | Information asset | Health vital data | Health vital data Body temperature. Pulse rate. Respiration rate,weight data, position data, etc.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
D05788-1, Ver 1

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Devices with default passwords needs to be checked for bruteforce attacks | Yes | n/a
V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | Yes | n/a
V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | Yes | n/a
V04 | Checking authentication modes for possible hacks and bypasses | Yes | n/a
V05 | Insecure communications in networks (hospital) | Yes | n/a
SBOM
V06 | Lack of Asset location digaram in security operations manual | Yes | n/a
V07 | Lack of configuration controls for IT assets in the informaion system plan | Yes | n/a
V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | Yes | n/a
V09 | Lack of plan for periodic Software Vulnerability Management | Yes | n/a
V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | Yes | n/a
V11 | Assest counting system for all instances of product implementation | Yes | n/a
Access points
V12 | Unprotected network port(s) on network devices and connection points | Yes | n/a
V13 | Unprotected external USB Port on the tablet/devices. | Yes | n/a
V14 | Unencrypted Network segment throught the information flow | Yes | n/a
V15 | Controlled Use of Administrative Privileges over the network | Yes | n/a
Data
V16 | Unencrypted data at rest in all possible locations | Yes | n/a
V17 | Unencrypted data in flight in all flowchannels | Yes | n/a
V18 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | Yes | n/a
V19 | Weak Algorthim implementation with respect cipher key size | Yes | n/a
InSecure Configurations of Resources
V20 | InSecure/not recommended Configuration for Mobile Devices, Laptops, Workstations, and Servers | Yes | n/a
V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | Yes | n/a
V22 | Legacy system identification if any | Yes | n/a
V23 | Outdated - Software/Hardware | Yes | n/a
AuthN management
V24 | Error Info containing sensitive data for Failed Authentication attempts | Yes | n/a
V25 | Absence of additional security factor along with user identification | Yes | n/a
V26 | Having no limit on the login attempts | Yes | n/a
V27 | No session expiry after certain time interval | Yes | n/a
Logging/Monitoring
V28 | Insufficient Logging information | Yes | n/a
V29 | Insufficient Access permissions for accessing and modifying Log files | Yes | n/a
Keys & Certificates
V30 | Improper security (for ex.,Storage & Access) for Key tokens and Certificates | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware (CAPEC-185) | Thread source delivers malware by providing removable media prepared with malware. Removable media is e.g. left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System. Malware eploits known a known vulnerability and e.g. gains admin priviledges. Undirected attack on computer systems. | TSA-3 - Skript Kiddies | Yes | n/a
T02 | Deliver directed malware (CAPEC-185) | Thread source delivers malware on a removable media which was designed to exploit a known vulnerability of the Navigation System. Directed attack on the Navigation System using knowledge about the Navigation System. | TSA-2 Organization | Yes | n/a
T03 | Gaining Access ([S]TRID[E]) | This phase is where an attacker breaks into the system/network using various tools or methods. After entering into a system, he has to increase his privilege to administrator level so he can install an application he needs or modify data or hide data | TSA-2 Organization | Yes | n/a
T04 | Maintaining Access (TTP) | The aim is to maintain the access to the target until he finishes the tasks he planned to accomplish in that target. | TSA-2 Organization | Yes | n/a
T05 | Clearing Track (TTP) | This involves modifying/corrupting/deleting the values of Logs, modifying registry values and uninstalling all applications he used and deleting all folders he created | TSA-2 Organization | Yes | n/a
T06 | Elevation of privilege (STRID[E]) | Identify weaknesses of segregation in terms of administrative and user-level privileges | TSA-2 Organization | Yes | n/a
T07 | Denial of service (STRI(D)E) | Find ways to exhaust or drown out legitimate requests | TSA-3 - Skript Kiddies | Yes | n/a
T08 | Information disclosure (STR(I)DE) | Fuzz application parameters or arguments to impact application error disclosures. Identify open ports with their respective services. Incite confidentiality and integrity in the browser interface. Identify clear text communications. Review usage of HTTP headers and user-agent profile. Pinpoint usages of API endpoints and application backend technologies. | TSA-2 Organization | Yes | n/a
T09 | Data Access (STR[I]DE) | Access user and application data e.g. by a malicious application or script | TSA-3 - Skript Kiddies | Yes | n/a
T10 | Open network port exploit (TTP) | Penetrate Open and Unsecured Ports | TSA-3 - Skript Kiddies | Yes | n/a
T11 | Brute-force Attack (CAPEC-112) | The brute-force attack contained a dictionary of well-known directories and authentication paradigms present in common webservers. | TSA-2 Organization | Yes | n/a
T12 | Social Engineering (TTP) | create custom phishing scams, phone-based attacks and even physical intrusion to test your organization’s level of security. | TSA-3 - Skript Kiddies | Yes | n/a
T13 | Lack of evidence to conclude any malicious attempt/attack (ST[R]IDE) | All the actions/events should be properly logged and the content needs to be protected by proper access rights. | TSA-2 Organization | Yes | n/a
T14 | Unauthorized Alterations (S[T]RIDE) | This involves modifying registry values, deleting/encrypting Confidential info and uninstalling Any secure applications and renaming/deleting all files/folders | TSA-2 Organization | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | Low | None | None | Network | Low | None | None | Changed | 3.887 | 0.22 | 1.4363 | 5.8 | 2.4 | LOW
2 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
4 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
14 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | High | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T01 | Deliver undirected malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. 9) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in flight in all flowchannels | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in flight in all flowchannels | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
21 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | High | 0.8 | 4.5 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | High | High | Low | Physical | High | High | Required | Unchanged | 0.12109 | 0.84899 | 5.4505 | 5.6 | 5.6 | MEDIUM | Justification
22 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.84899 | 5.4505 | 6.8 | High | 0.8 | 6.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
23 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
24 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Moderate | 0.5 | 3.7 | LOW | • Stateful Firewall • Maintain access control list |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Use sniffing tool to sniff the data at motion and MITM 7) Exploit the open ports found while assessment and information gathering. 8) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 9) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
25 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | High | 0.8 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
26 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | • Use Secure tunnel Communications channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
27 | T02 | Deliver directed malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | High | 0.8 | 5.1 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
28 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
29 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
30 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
31 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.84899 | 5.4505 | 8.3 | Moderate | 0.5 | 6.9 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
32 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.84899 | 5.4505 | 8.3 | Low | 0.2 | 6.1 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
33 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | High | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.84899 | 5.4505 | 8.3 | Low | 0.2 | 6.1 | MEDIUM | • Disable the device network discoverable • Maintain Access Control List. • Block all unrequired ports • Stateful firewall |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Use sniffing tool to sniff the data at motion and MITM 7) Exploit the open ports found while assessment and information gathering. 8) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 9) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
34 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | High | Low | Low | Local | High | Low | Required | Unchanged | 0.76466 | 0.7323 | 4.7014 | 5.5 | High | 0.8 | 5.4 | MEDIUM | • Implement automated configuration monitoring systems • Use Secure tunnel Communications channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
35 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
36 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. 9) Use sniffing tool to sniff the data at motion and MITM | Low | None | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.3916 | 2.5141 | 3.2 | 2.7 | LOW
37 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A02 | Tablet OS/network details & Tablet Application | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. 9) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
38 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | • Encrypting the storage subsystem • Encryption methods such as HTTPS, SSL, and TLS are often used to protect data in motion. • Stateful firewall • Maintain access control list. • Use strong encrption algorithm |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
39 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
40 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | High | 0.8 | 7 | HIGH | • Configuring account lockout policies cannot be exploited to lock out well known service accounts. • Ensuring application is capable of handling required volumes of traffic and that thresholds are in place to handle optimal loads. • Review your application's failover functionality. • Maintain Access Control List. • Block all unrequired ports • Stateful firewall |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
41 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | Low | None | None | Network | High | None | None | Unchanged | 2.2212 | 0.22 | 1.4124 | 3.7 | 1.9 | LOW
42 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
43 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.7323 | 4.7014 | 5.4 | High | 0.8 | 5.3 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
44 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A07 | Interface/API Communication | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
45 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
46 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | High | 0.8 | 5.8 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
47 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | High | 0.8 | 5.8 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
48 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
49 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
50 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
51 | T03 | Gaining Access ([S]TRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
52 | T03 | Gaining Access ([S]TRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi-factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
53 | T03 | Gaining Access ([S]TRID[E]) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
54 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.7323 | 4.7014 | 5.4 | High | 0.8 | 5.3 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
55 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.7323 | 4.7014 | 5.4 | High | 0.8 | 5.3 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
56 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
57 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
58 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
59 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
60 | T04 | Maintaining Access (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
61 | T04 | Maintaining Access (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
62 | T04 | Maintaining Access (TTP) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | High | High | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.84899 | 5.4505 | 7.6 | Moderate | 0.5 | 6.5 | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
63 | T04 | Maintaining Access (TTP) | V15 | Controlled Use of Administrative Privileges over the network | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
64 | T05 | Clearing Track (TTP) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
65 | T05 | Clearing Track (TTP) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | High | Low | Required | Unchanged | 0.76466 | 0.52545 | 3.3734 | 4.2 | Low | 0.2 | 3.6 | LOW | • Establish secure configuration • Deploy system configuration management tool • Implement automated configuration monitoring systems • Establish internal and external information sources for threat intelligence and vulnerability data, monitoring them regularly and taking appropriate action for high-priority items • Upgrades the software, firmware • Never use credentials such as date of birth, spouse, or child’s or pet’s name • Stateful Firewall |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
66 | T05 | Clearing Track (TTP) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
67 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
68 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) |  |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
69 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
70 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
71 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A02 | Tablet OS/network details & Tablet Application | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
72 | T05 | Clearing Track (TTP) | V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | NA | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
73 | T05 | Clearing Track (TTP) | V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
74 | T06 | Elevation of privilege (STRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A04 | Authenication/Authorisation data | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
75 | T06 | Elevation of privilege (STRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | High | High | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.84899 | 5.4505 | 7.6 | High | 0.8 | 7.2 | HIGH | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
76 | T07 | Denial of service (STRI(D)E) | V12 | Unprotected network port(s) on network devices and connection points | A08 | Wireless Network device | 1) Bring down the service availability 2) Blocking the end user usage |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | • Eliminating rogue access points • Properly configuring all authorized access points • Firewall • Limit Network Access |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Use sniffing tool to sniff the data at motion and MITM 7) Exploit the open ports found while assessment and information gathering. 8) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 9) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
77 | T07 | Denial of service (STRI(D)E) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Bring down the service availability 2) Blocking the end user usage |  | High | Low | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.7323 | 4.7014 | 7.6 | Low | 0.2 | 5.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
78 | T08 | Information disclosure (STR(I)DE) | V16 | Unencrypted data at rest in all possible locations | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.84899 | 5.4505 | 6.5 | Moderate | 0.5 | 6 | MEDIUM | • Encrypting the storage subsystem • Encryption methods such as HTTPS, SSL, and TLS are often used to protect data in motion. • Stateful firewall • Host Hardening • Maintain access control list. • Use strong encrption algorithm |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM 8) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
79 | T08 | Information disclosure (STR(I)DE) | V17 | Unencrypted data in flight in all flowchannels | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Use secure tunnel communication channel • Maintain access control list. • Use strong encrption algorithm |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
80 | T08 | Information disclosure (STR(I)DE) | V18 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.84899 | 5.4505 | 6.5 | Moderate | 0.5 | 6 | MEDIUM | • Server-side encryption using Service-Managed keys • Network access control • Encryption for sensitive data in transit, such as when files are attached to an email message or moved to cloud storage, removable drives • Transfer over encrypted tunnel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
81 | T08 | Information disclosure (STR(I)DE) | V18 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Server-side encryption using Service-Managed keys • Use secure tunnel communication channel • Network access control • Encryption for sensitive data in transit, such as when files are attached to an email message or moved to cloud storage, removable drives • Transfer over encrypted tunnel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
82 | T08 | Information disclosure (STR(I)DE) | V19 | Weak Algorthim implementation with respect cipher key size | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.84899 | 5.4505 | 6.5 | Moderate | 0.5 | 6 | MEDIUM | • Implement strong encryption descrption method at data at rest • Use secure tunnel communication channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
83 | T08 | Information disclosure (STR(I)DE) | V19 | Weak Algorthim implementation with respect cipher key size | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Implement strong sysmetric or unsysmetric algorhythn • Implementing VPN and stateful Firewall |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
84 | T08 | Information disclosure (STR(I)DE) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
85 | T08 | Information disclosure (STR(I)DE) | V14 | Unencrypted Network segment throught the information flow | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Anonymization/Pseudomyzation • Encyrption • Mainitaing Access Logs • Maintain Server Security Logs |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
86 | T08 | Information disclosure (STR(I)DE) | V05 | Insecure communications in networks (hospital) | A10 | Data in Motion | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Use strong encryption. • Secure communication links with protocols that provide message confidentiality. • Secure sensitive data in the channel flow • Protect authentication cookies with Secure Sockets Layer (SSL) or TLS. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
87 | T09 | Data Access (STR[I]DE) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
88 | T09 | Data Access (STR[I]DE) | V12 | Unprotected network port(s) on network devices and connection points | A08 | Wireless Network device | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Firewall • Secure communication links with protocols that provide message confidentiality. |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Use sniffing tool to sniff the data at motion and MITM 7) Exploit the open ports found while assessment and information gathering. 8) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 9) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
89 | T09 | Data Access (STR[I]DE) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
90 | T09 | Data Access (STR[I]DE) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A09 | Data at Rest | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Firewall • Do not store sensitive data in plaintext. • Use strong encrption algorithm. • Apply salting over sensitive data. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
91 | T09 | Data Access (STR[I]DE) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | High | 0.8 | 6.1 | MEDIUM | • Firewall • Do not store sensitive data in plaintext. • Use strong encrption algorithm. • Apply salting over sensitive data. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
92 | T09 | Data Access (STR[I]DE) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A10 | Data in Motion | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Firewall • Do not store sensitive data in plaintext. • Use strong encrption algorithm. • Apply salting over sensitive data. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
93 | T09 | Data Access (STR[I]DE) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A09 | Data at Rest | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Stronger authentication methods • Use strong encrption algorithm. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
94 | T09 | Data Access (STR[I]DE) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Stronger authentication methods • Use strong encrption algorithm. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
95 | T09 | Data Access (STR[I]DE) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Low | 0.2 | 5.6 | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
96 | T09 | Data Access (STR[I]DE) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | High | Low | Physical | High | High | None | Unchanged | 0.16601 | 0.84899 | 5.4505 | 5.7 | High | 0.8 | 5.6 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
97 | T10 | Open network port exploit (TTP) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
98 | T10 | Open network port exploit (TTP) | V12 | Unprotected network port(s) on network devices and connection points | A08 | Wireless Network device | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Using port security on Firewall • Application level Firewall • Check ing internet provider’s or router manufacturer’s wireless security options. |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Use sniffing tool to sniff the data at motion and MITM 7) Exploit the open ports found while assessment and information gathering. 8) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 9) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
99 | T10 | Open network port exploit (TTP) | V14 | Unencrypted Network segment throught the information flow | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
100 | T10 | Open network port exploit (TTP) | V14 | Unencrypted Network segment throught the information flow | A08 | Wireless Network device | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • IP Whitelisting • Use strong encryption algorithyms |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Use sniffing tool to sniff the data at motion and MITM 7) Exploit the open ports found while assessment and information gathering. 8) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 9) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
101 | T10 | Open network port exploit (TTP) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points, |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
102 | T10 | Open network port exploit (TTP) | V17 | Unencrypted data in flight in all flowchannels | A10 | Data in Motion | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | • Use secure tunnel communication channel • Configure and upgrade routers for the n/w security • Configure firewalls to reject any packets with spoofed addresses. • Maintain access control list. • Use strong encrption algorithm |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
103 | T10 | Open network port exploit (TTP) | V05 | Insecure communications in networks (hospital) | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | High | High | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.84899 | 5.4505 | 7.1 | Moderate | 0.5 | 6.3 | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel |  | 1) Create Android malware 2) Transfer the malware to tablet/Smart Medic Device 3) Malware execution on the device 4) Exploit the devices with respect to vulnerability 5) Check for the open ports 6) Exploit the open ports found while assessment and information gathering. 7) Vulnerability Assessment scanning for the identifying unknown vulnerabilities 8) Exploit the found loopholes while VA scanning using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
104 | T11 | Brute-force Attack (CAPEC-112) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
105 | T11 | Brute-force Attack (CAPEC-112) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
106 | T11 | Brute-force Attack (CAPEC-112) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
107 | T11 | Brute-force Attack (CAPEC-112) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
108 | T11 | Brute-force Attack (CAPEC-112) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
109 | T11 | Brute-force Attack (CAPEC-112) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
110 | T11 | Brute-force Attack (CAPEC-112) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Set your spam filters setting options to high • Secure your computing devices |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
111 | T11 | Brute-force Attack (CAPEC-112) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Set your spam filters setting options to high • Secure your computing devices |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
112 | T11 | Brute-force Attack (CAPEC-112) | V18 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | A09 | Data at Rest | 1) An attacker may attempt to discover a weak encryption by systematically trying every possible combination of decryption key. |  | High | High | Low | Local | High | High | None | Unchanged | 0.45653 | 0.84899 | 5.4505 | 6 | Moderate | 0.5 | 5.7 | MEDIUM | • Server-side encryption using Service-Managed keys • Network access control • Encryption for sensitive data in transit, such as when files are attached to an email message or moved to cloud storage, removable drives • Transfer over encrypted tunnel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
113 | T11 | Brute-force Attack (CAPEC-112) | V18 | Weak Encryption Implementaion in data at rest and in motion tactical and design wise | A10 | Data in Motion | 1) An attacker may attempt to discover a weak encryption by systematically trying every possible combination of decryption key. |  | High | High | Low | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Moderate | 0.5 | 5.9 | MEDIUM | • Firewall • Encrypting data at rest • Protect your data with multi factor authentication • Use secure tunnel communication channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port 7) Use sniffing tool to sniff the data at motion and MITM | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
114 | T12 | Social Engineering (TTP) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Implement automated configuration monitoring systems • Stateful firewall |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
115 | T12 | Social Engineering (TTP) | V22 | Legacy system identification if any | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Delete any request for personal information • Firewall • Disable device network discoverable • Maintain access control list |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
116 | T12 | Social Engineering (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Stronger authentication methods |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
117 | T12 | Social Engineering (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A07 | Interface/API Communication | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Set your spam filters setting options to high • Firewalls, • Use secure tunnel communication channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
118 | T12 | Social Engineering (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Set your spam filters setting options to high • Firewalls, • Use Multi Factor Authentication • Use secure tunnel communication channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
119 | T12 | Social Engineering (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A13 | Azure Cloud DataBase | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Set your spam filters setting options to high • Firewalls, • Use Multi Factor Authentication • Use secure tunnel communication channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
120 | T12 | Social Engineering (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | High | High | Low | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.84899 | 5.4505 | 5.9 | Moderate | 0.5 | 5.7 | MEDIUM | • Firewall • Use Multi Factor Authentication • Use secure tunnel communication channel |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
121 | T13 | Lack of evidence to conclude any malicious attempt/attack (ST[R]IDE) | V28 | Insufficient Logging information | A12 | Smart medic app (Azure Portal Administrator) | 1) Adversary tried to obtain knowledge about system internals 2) Attempt to find attack vectors 3) Executed malicious activities 4) Complete details related to the attacker/malicious activities not recorded |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | All the information needed for identifying the threat (malicious) activity and adversary information needed to be logged for determining the attack vector and attach surface. This helps to make the system less vulnerable |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
122 | T13 | Lack of evidence to conclude any malicious attempt/attack (ST[R]IDE) | V29 | Insufficient Access permissions for accessing and modifying Log files | A12 | Smart medic app (Azure Portal Administrator) | 1) Adversary tried to obtain knowledge about system internals 2) Attempt to find attack vectors 3) Executed malicious activities 4) Complete details related to the attacker/malicious activities not recorded |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | All the information needed for identifying the threat (malicious) activity and adversary information needed to be logged for determining the attack vector and attach surface. This helps to make the system less vulnerable |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
123 | T14 | Unauthorized Alterations (S[T]RIDE) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A13 | Azure Cloud DataBase | 1) Adversary tried to obtain knowledge about system internals 2) Identify unsecured critical data/information (eg: patient info, keys, certificates...) 3) Modify the identified content |  | Low | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | System configurations and oudated/undated software should be properly monitored for avoiding the malicious activities and tampering of data |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
124 | T14 | Unauthorized Alterations (S[T]RIDE) | V29 | Insufficient Access permissions for accessing and modifying Log files | A13 | Azure Cloud DataBase | 1) Adversary tried to obtain knowledge about system internals 2) Identify unsecured critical data/information (eg: patient info, keys, certificates...) 3) Modify the identified content |  | Low | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Moderate | 0.5 | 5.7 | MEDIUM | Strict authorization policies has to be designed based on either asset/role based and has to be enforcedfor proper protection of the data and alterations |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
125 | T03 | Gaining Access ([S]TRID[E]) | V24 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
126 | T03 | Gaining Access ([S]TRID[E]) | V24 | Error Info containing sensitive data for Failed Authentication attempts | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
127 | T03 | Gaining Access ([S]TRID[E]) | V25 | Absence of additional security factor along with user identification | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
128 | T03 | Gaining Access ([S]TRID[E]) | V25 | Absence of additional security factor along with user identification | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
129 | T03 | Gaining Access ([S]TRID[E]) | V26 | Having no limit on the login attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
130 | T03 | Gaining Access ([S]TRID[E]) | V26 | Having no limit on the login attempts | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
131 | T11 | Brute-force Attack (CAPEC-112) | V24 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
132 | T11 | Brute-force Attack (CAPEC-112) | V24 | Error Info containing sensitive data for Failed Authentication attempts | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
133 | T11 | Brute-force Attack (CAPEC-112) | V25 | Absence of additional security factor along with user identification | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
134 | T11 | Brute-force Attack (CAPEC-112) | V25 | Absence of additional security factor along with user identification | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
135 | T11 | Brute-force Attack (CAPEC-112) | V26 | Having no limit on the login attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
136 | T11 | Brute-force Attack (CAPEC-112) | V26 | Having no limit on the login attempts | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | Standard error info messages should be used. Information aiding the attacks should be avoided. Apart from user id there should be additional security factor for verification. Limit on the login attempts is mandatory. |  | 1) Vulnerability Assessment scanning for the identifying unknown vulnerabilities of web application. 2) Check for the open ports using nmap, other kali tools 3) Exploit the found loopholes while VA Scanning using the Burpsuite, kali tools. 4) Exlpoit the open ports using kali tools. 5) Brute Force attempt for the authentication/authorization on the open port 6) Exploit related to Brute Force attempt for the authentication/authorization on the open port | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | LOW
2 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
3 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
4 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
5 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
6 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
7 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
8 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
9 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
10 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
11 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
12 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
13 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
14 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
15 | T01 | Deliver undirected malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
16 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in flight in all flowchannels | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
17 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in flight in all flowchannels | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
18 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
19 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
20 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
21 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | MEDIUM | Justification
22 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
23 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
24 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | • Stateful Firewall • Maintain access control list
25 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
26 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | • Use Secure tunnel Communications channel
27 | T02 | Deliver directed malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
28 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
29 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
30 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
31 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
32 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
33 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A08 | Wireless Network device | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | • Disable the device network discoverable • Maintain Access Control List. • Block all unrequired ports • Stateful firewall
34 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | • Implement automated configuration monitoring systems • Use Secure tunnel Communications channel
35 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
36 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | LOW
37 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A02 | Tablet OS/network details & Tablet Application | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
38 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | • Encrypting the storage subsystem • Encryption methods such as HTTPS, SSL, and TLS are often used to protect data in motion. • Stateful firewall • Maintain access control list. • Use strong encrption algorithm
39 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
40 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | HIGH | • Configuring account lockout policies cannot be exploited to lock out well known service accounts. • Ensuring application is capable of handling required volumes of traffic and that thresholds are in place to handle optimal loads. • Review your application's failover functionality. • Maintain Access Control List. • Block all unrequired ports • Stateful firewall
41 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel | LOW
42 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs
43 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
44 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A07 | Interface/API Communication | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
45 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
46 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
47 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
48 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
49 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
50 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Encrypt authentication data in storage and transit to mitigate risk of information disclosure and authentication protocol attacks • Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks • Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks •Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
51 | T03 | Gaining Access ([S]TRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
52 | T03 | Gaining Access ([S]TRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require that administrators establish multi-factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
53 | T03 | Gaining Access ([S]TRID[E]) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
54 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
55 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
56 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
57 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
58 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
59 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require multi-factor authentication • limit authentication attempts (rate Limiting) • Maintain Access Logs • Maintain Server Security Logs • Stronger authentication methods
60 | T04 | Maintaining Access (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses.
61 | T04 | Maintaining Access (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | •Never use credentials such as date of birth, spouse, or child’s or pet’s name •Lockout an account subjected to too many incorrect credential guesses.
62 | T04 | Maintaining Access (TTP) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
63 | T04 | Maintaining Access (TTP) | V15 | Controlled Use of Administrative Privileges over the network | A04 | Authenication/Authorisation data | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | • Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. • Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. • Change all default credentials for applications, operating systems, routers, firewalls, wireless access points,
64 | T05 | Clearing Track (TTP) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
65 | T05 | Clearing Track (TTP) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Azure Cloud Web Application) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | • Establish secure configuration • Deploy system configuration management tool • Implement automated configuration monitoring systems • Establish internal and external information sources for threat intelligence and vulnerability data, monitoring them regularly and taking appropriate action for high-priority items • Upgrades the software, firmware • Never use credentials such as date of birth, spouse, or child’s or pet’s name • Stateful Firewall
66 | T05 | Clearing Track (TTP) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
67 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
68 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) |  |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
69 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | MEDIUM | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
70 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
71 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A02 | Tablet OS/network details & Tablet Application | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
72 | T05 | Clearing Track (TTP) | V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | ▪Asset should be behind stateful firewall • Use secure tunnel communications channel
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls