## System & Asset Identification
System & Asset Identification
Medical Device / System:
Scope:
Date: |  | <2021-07-12>
Conducted by: |  | <Author Name / Function / Organization> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description
A01 | Physical Asset | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | Utilizing computer resources and computing power by adversary, allows various general purpose attacks, such as incl. Ransomware deployment, Bitcoin Mining, abuse of peripheral devices such as WebCam, Microphones, etc., .
A02 | Information asset | Tablet OS/network details & Tablet Application | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A03 | Physical Assets | Smart medic (Stryker device) System Component | Monitors local bed status information, alerting caregivers visually, audibly or remotely if preset parameters are compromised.
A04 | Information asset | Authentication/Authorisation method of all device(s)/app | Information related to authenication/authorisation data (password/pins/MFA/Biometrics)
A05 | Physical Assets | Device Maintainence tool (Hardware/Software) | Device Maintainence tool (Hardware/Software) that patchs and updates Smart Medic Device and Application related to Security
A06 | Information asset | Electronic Health Records (EHR)/ Device Component status | Smart device components health status information
A07 | Information asset | Interface/API Communication | Communication middleware enables communication and data management for distributed applications.
A08 | Physical Assets | Wireless Network device (Scope of HDO) | Devices that are used for communication among the Smart Medic project component.
A09 | Information asset | Data at Rest | Use strong encryption algorthim to store data on cloud platform (Smartmedic Device)/tablet
A10 | Information asset | Data in Transit | Use strong encryption algorthim to data moving on tablet to cloud platform(Smartmedic Device)/tablet
A11 | Information asset | Smart medic app (Stryker Admin Web Application) | Smart medic application for nurse/health worker (Stryker Admin Web Application)
A12 | Information asset | Smart medic app (Azure Portal Administrator) | Azure Portal Administrator for Smart medic app
A13 | Information asset | Azure Cloud DataBase | Azure Cloud DataBase related to Smart Medic app
A14 | Information asset | Health vital data | Health vital data Body temperature. Pulse rate. Respiration rate,weight data, position data, etc.
A15 | Information asset | Nurse Station Application | Smart medic web application for nurse/health worker running on the Nurse Station
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
D05788-1, Ver 1

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Devices with default passwords needs to be checked for bruteforce attacks | Yes | n/a
V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | Yes | n/a
V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | Yes | n/a
V04 | Checking authentication modes for possible hacks and bypasses | Yes | n/a
V05 | Insecure communications in networks (hospital) | Yes | n/a
SBOM
V06 | Lack of Asset location digaram in security operations manual | Yes | n/a
V07 | Lack of configuration controls for IT assets in the informaion system plan | Yes | n/a
V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | Yes | n/a
V09 | Lack of plan for periodic Software Vulnerability Management | Yes | n/a
V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | Yes | n/a
V11 | Assest counting system for all instances of product implementation | Yes | n/a
Access points
V12 | Unprotected network port(s) on network devices and connection points | Yes | n/a
V13 | Unprotected external USB Port on the tablet/devices. | Yes | n/a
V14 | Unencrypted Network segment through out the information flow | Yes | n/a
V15 | Controlled Use of Administrative Privileges over the network | Yes | n/a
Data
V16 | Unencrypted data at rest in all possible locations | Yes | n/a
V17 | Unencrypted data in transit in all flowchannels | Yes | n/a
V18 | Weak Encryption Implementaion in data at rest and in transit tactical and design wise | Yes | n/a
V19 | Weak Algorthim implementation with respect cipher key size | Yes | n/a
InSecure Configurations of Resources
V20 | InSecure/not recommended Configuration for Mobile Devices, Laptops, Workstations, and Servers | Yes | n/a
V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | Yes | n/a
V22 | Legacy system identification if any | Yes | n/a
V23 | Outdated - Software/Hardware | Yes | n/a
V31 | Improper/insufficient provisioning of IOT hub | Yes | n/a
V32 | Unsecured communication with unauthenticated 3rd party devices | Yes | n/a
AuthN management
V24 | Error Info containing sensitive data for Failed Authentication attempts | Yes | n/a
V25 | Absence of additional security factor along with user identification | Yes | n/a
V26 | Having no limit on the login attempts | Yes | n/a
V27 | No session expiry after certain time interval | Yes | n/a
Logging/Monitoring
V28 | Insufficient Logging information | Yes | n/a
V29 | Insufficient Access permissions for accessing and modifying Log files | Yes | n/a
Keys & Certificates
V30 | Improper security (for ex.,Storage & Access) for Key tokens and Certificates | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware (CAPEC-185) | Thread source delivers malware by providing removable media prepared with malware. Removable media is e.g. left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System. Malware eploits known a known vulnerability and e.g. gains admin priviledges. Undirected attack on computer systems. | TSA-3 - Skript Kiddies | Yes | n/a
T02 | Deliver directed malware (CAPEC-185) | Thread source delivers malware on a removable media which was designed to exploit a known vulnerability of the Navigation System. Directed attack on the Navigation System using knowledge about the Navigation System. | TSA-2 Organization | Yes | n/a
T03 | Gaining Access ([S]TRID[E]) | This phase is where an attacker breaks into the system/network using various tools or methods. After entering into a system, he has to increase his privilege to administrator level so he can install an application he needs or modify data or hide data | TSA-2 Organization | Yes | n/a
T04 | Maintaining Access (TTP) | The aim is to maintain the access to the target until he finishes the tasks he planned to accomplish in that target. | TSA-2 Organization | Yes | n/a
T05 | Clearing Track (TTP) | This involves modifying/corrupting/deleting the values of Logs, modifying registry values and uninstalling all applications he used and deleting all folders he created | TSA-2 Organization | Yes | n/a
T06 | Elevation of privilege (STRID[E]) | Identify weaknesses of segregation in terms of administrative and user-level privileges | TSA-2 Organization | Yes | n/a
T07 | Denial of service (STRI(D)E) | Find ways to exhaust or drown out legitimate requests | TSA-3 - Skript Kiddies | Yes | n/a
T08 | Information disclosure (STR(I)DE) | Fuzz application parameters or arguments to impact application error disclosures. Identify open ports with their respective services. Incite confidentiality and integrity in the browser interface. Identify clear text communications. Review usage of HTTP headers and user-agent profile. Pinpoint usages of API endpoints and application backend technologies. | TSA-2 Organization | Yes | n/a
T09 | Data Access (STR[I]DE) | Access user and application data e.g. by a malicious application or script | TSA-3 - Skript Kiddies | Yes | n/a
T10 | Open network port exploit (TTP) | Penetrate Open and Unsecured Ports | TSA-3 - Skript Kiddies | Yes | n/a
T11 | Brute-force Attack (CAPEC-112) | The brute-force attack contained a dictionary of well-known directories and authentication paradigms present in common webservers. | TSA-2 Organization | Yes | n/a
T12 | Social Engineering (TTP) | create custom phishing scams, phone-based attacks and even physical intrusion to test your organization’s level of security. | TSA-3 - Skript Kiddies | Yes | n/a
T13 | Lack of evidence to conclude any malicious attempt/attack (ST[R]IDE) | All the actions/events should be properly logged and the content needs to be protected by proper access rights. | TSA-2 Organization | Yes | n/a
T14 | Unauthorized Alterations (S[T]RIDE) | This involves modifying registry values, deleting/encrypting Confidential info and uninstalling Any secure applications and renaming/deleting all files/folders | TSA-2 Organization | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel |  | Low | None | None | Network | Low | None | None | Changed | 3.887 | 0.22 | 1.4363 | 5.8 | 2.4 | LOW
2 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
4 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Low | 0.2 | 3.9 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
14 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Low | 0.2 | 3.9 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T01 | Deliver undirected malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in transit in all flowchannels | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Low | 0.2 | 3.9 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in transit in all flowchannels | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Low | 0.2 | 3.9 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
21 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Local | Low | High | Required | Unchanged | 0.58275 | 0.56 | 3.5952 | 4.2 | Moderate | 0.5 | 3.9 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel |  | High | High | Low | Physical | High | High | Required | Unchanged | 0.12109 | 0.84899 | 5.4505 | 5.6 | 5.6 | MEDIUM | Justification
22 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Local | Low | High | Required | Unchanged | 0.58275 | 0.56 | 3.5952 | 4.2 | Moderate | 0.5 | 3.9 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
23 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
24 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A08 | Wireless Network device (Scope of HDO) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Moderate | 0.5 | 3.7 | LOW | SOM responsibility 1. Statefull Firewall 2. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SOM D001020115 - 05. Access control policy and management | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
25 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Moderate | 0.5 | 3.7 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
26 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A11 | Smart medic app (Stryker Admin Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
27 | T02 | Deliver directed malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Network | Low | High | Required | Unchanged | 0.90061 | 0.56 | 3.5952 | 4.5 | Moderate | 0.5 | 4.1 | MEDIUM | 1. Only stryker made/authenticated devices should communicate with smart medic device & tablet 2. Asset should be behind stateful firewall 3. Use secure tunnel communications channel | 1. NA 2. SOM D001020115 - 23. Malware Detection/Protection 3. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
28 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
29 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
30 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
31 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Moderate | 0.5 | 4.3 | MEDIUM | 1. Only Stryker/HDO authenticated devices should communicate with smart medic device & tablet 2. Asset should be behind stateful firewall 3. Use secure tunnel communications channel | 1.NA- SOM - (Firewall and HDO ) 2.SOM D001020115 - 23. Malware Detection/Protection 3. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
32 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Low | 0.2 | 3.9 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
33 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Admin Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | None | None | High | Local | High | High | Required | Unchanged | 0.333 | 0.56 | 3.5952 | 4 | Moderate | 0.5 | 3.8 | LOW | 1. Deployed (V&V) secure system configuration model needs to be mentioned in the installation manual. 2. Deploy system configuration management tool 3. Implement automated configuration monitoring systems 4. Establish internal and external information sources for threat intelligence and vulnerability data, monitoring them regularly and taking appropriate action for high-priority items 5. Use upgraded software, firmware 6 Never create/use credentials with personal details such as date of birth, spouse, or child’s or pet’s name 7. Stateful Firewall | 1.NA 2.NA 3.NA 4.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 5.SRS D001020097 –2.1.7.2 The application shall allow to upgrade the tablet application. 6.NA 7.SOM D001020115 - 23. Malware Detection/Protection | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
34 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
35 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel |  | Low | None | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.3916 | 2.5141 | 3.2 | 2.7 | LOW
36 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A02 | Tablet OS/network details & Tablet Application | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
37 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A11 | Smart medic app (Stryker Admin Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Identification of the sensitive data in storage and encryption of storage subsystem 2. Stateful firewall 3. Hardening of the host system containing sensitive data at rest 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. Use strong encrption algorithm | 1. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 2. SOM D001020115 - 23. Malware Detection/Protection 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SOM D001020115 - 05 Access control policy and management 5.SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
38 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | None | None | Low | Network | Low | Low | None | Unchanged | 2.8353 | 0.22 | 1.4124 | 4.3 | Low | 0.2 | 2 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
39 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A11 | Smart medic app (Stryker Admin Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | None | Low | High | Network | Low | High | None | Unchanged | 1.2347 | 0.6568 | 4.2167 | 5.5 | Low | 0.2 | 4.5 | MEDIUM | 1. Admin application can be accessed by login credentials & MFA. Hence, strong password policies & management are required 2. Data transfer between the admin application and the smart medic components needs to be encrypted & secured. 3. Any vulnerable network ports and connection points should be identified and hardened. 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. Stateful firewall | 1. SRS D001020024 -2.1.2.2 If the Admin Pin is valid and then the application shall navigate further. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha. 2.SRS D001020023-2.13.1Communication between browser application and backend shall be secured. 3. SRS D001020024 -2.17.4 The Application shall establish technical controls to mitigate the potential for compromise to the integrity and confidentiality of health data stored on the product or removable media 4.NA 5. SOM D001020115 - 23. Malware Detection/Protection | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
40 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | None | Low | None | Network | Low | Low | None | Unchanged | 2.8353 | 0.22 | 1.4124 | 4.3 | Low | 0.2 | 2 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel |  | Low | None | None | Network | High | None | None | Unchanged | 2.2212 | 0.22 | 1.4124 | 3.7 | 1.9 | LOW
41 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | None | High | Physical | Low | Low | None | Unchanged | 0.66712 | 0.6568 | 4.2167 | 4.9 | Moderate | 0.5 | 4.6 | MEDIUM | 1.During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Maintain Access Logs | 1. NA 2.SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
42 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A07 | Interface/API Communication | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | None | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.3916 | 2.5141 | 3.2 | Low | 0.2 | 2.7 | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods | 1. Installation Manual -Tablet, 2.SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5.NA 6.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
43 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | None | High | Local | Low | Low | Required | Unchanged | 1.3382 | 0.6568 | 4.2167 | 5.6 | Low | 0.2 | 4.5 | MEDIUM | 1. If devices/apps being accessed by login credentials & MFA. Then, strong password policies & management are required 2. Require multi-factor authentication 3 Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 5 Maintain Server Security Logs 6. Stronger authentication methods | 1. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 2. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 5. NA 6. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
44 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. Encrypt the authentication data in storage & transit to mitigate risk of information disclosure and authentication protocol attacks 2. Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks 3. Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks 4. Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 2.SRS D001020097 – 2.8.1Application shall use APIs to communicate between browser application and the backend. SRS D001020023-2.8.1Application shall use APIs to communicate between browser application and the backend. 3. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 4.SRS D001020097 –2.1.7.1.1 Something went wrong with API operation try again / contact API admin. SRS D001020023-2.1.4.1.1 Something went wrong with API operation try again / contact API admin. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
45 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Admin Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. Encrypt the authentication data in storage & transit to mitigate risk of information disclosure and authentication protocol attacks 2. Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks 3. Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks 4. Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 2.SRS D001020097 – 2.8.1Application shall use APIs to communicate between browser application and the backend. SRS D001020023-2.8.1Application shall use APIs to communicate between browser application and the backend. 3. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 4.SRS D001020097 –2.1.7.1.1 Something went wrong with API operation try again / contact API admin. SRS D001020023-2.1.4.1.1 Something went wrong with API operation try again / contact API admin. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
46 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. Encrypt the authentication data in storage & transit to mitigate risk of information disclosure and authentication protocol attacks 2. Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks 3. Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks 4. Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 2.SRS D001020097 – 2.8.1Application shall use APIs to communicate between browser application and the backend. SRS D001020023-2.8.1Application shall use APIs to communicate between browser application and the backend. 3. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 4.SRS D001020097 –2.1.7.1.1 Something went wrong with API operation try again / contact API admin. SRS D001020023-2.1.4.1.1 Something went wrong with API operation try again / contact API admin. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
47 | T03 | Gaining Access ([S]TRID[E]) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.52545 | 3.3734 | 3.9 | Low | 0.2 | 3.5 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
48 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods | 1. NA 2.SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5.NA 6.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
49 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | 1. If devices/apps being accessed by login credentials & MFA. Then, strong password policies & management are required 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 2.SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5.NA 6.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
50 | T05 | Clearing Track (TTP) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
51 | T05 | Clearing Track (TTP) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Physical | Low | Low | None | Unchanged | 0.66712 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
52 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
53 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Low | 0.2 | 3.7 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
54 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | None | Low | Low | Local | Low | High | None | Unchanged | 0.79893 | 0.3916 | 2.5141 | 3.4 | Moderate | 0.5 | 3 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
55 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
56 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A02 | Tablet OS/network details & Tablet Application | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
57 | T05 | Clearing Track (TTP) | V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
58 | T05 | Clearing Track (TTP) | V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
59 | T06 | Elevation of privilege (STRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A04 | Authentication/Authorisation method of all device(s)/app | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | Low | Low | Low | Network | Low | Low | Required | Unchanged | 2.0681 | 0.52545 | 3.3734 | 5.5 | Low | 0.2 | 3.8 | LOW | 1. Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. 2. Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. 3. Ensure default credentials not existing for any assets (such as applications, operating systems, routers, firewalls, wireless access points). | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 2.NA 3. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
60 | T06 | Elevation of privilege (STRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | None | Low | High | Network | Low | High | Required | Unchanged | 0.90061 | 0.6568 | 4.2167 | 5.2 | Moderate | 0.5 | 4.7 | MEDIUM | 1. Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. 2. Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. 3. Ensure default credentials not existing for any assets (such as applications, operating systems, routers, firewalls, wireless access points). | 1. SRS D001020024 -2.1.2.2 If the Admin Pin is valid and then the application shall navigate further. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha. 2.NA 3. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
61 | T07 | Denial of service (STRI(D)E) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Bring down the service availability 2) Blocking the end user usage |  | None | None | High | Network | Low | Low | None | Unchanged | 2.8353 | 0.56 | 3.5952 | 6.5 | Low | 0.2 | 4.2 | MEDIUM | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
62 | T08 | Information disclosure (STR(I)DE) | V16 | Unencrypted data at rest in all possible locations | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | Low | Low | Local | High | High | None | Unchanged | 0.45653 | 0.52545 | 3.3734 | 3.9 | Moderate | 0.5 | 3.7 | LOW | 1. Identification of the sensitive data in storage and encryption of storage subsystem 2. Stateful firewall 3. Hardening of the host system containing sensitive data at rest 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. Use strong encrption algorithm | 1. SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 2. SOM D001020115 - 23. Malware Detection/Protection 3.NA 4. NA 5.SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
63 | T08 | Information disclosure (STR(I)DE) | V17 | Unencrypted data in transit in all flowchannels | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | 1. Use secure tunnel communication channel 2. Configure and upgrade routers for the n/w security 3. Configure firewalls to reject any packets with spoofed addresses. 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. For sensitive data proper encryption mechanism needs to be designed & implemented | 1. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel 2. SOM D001020115 - 16. Transmission confidentiality and integrity 3. SOM D001020115 - 23. Malware Detection/Protection 4. SOM D001020115 - 05. Access control policy and management 5. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
64 | T08 | Information disclosure (STR(I)DE) | V18 | Weak Encryption Implementaion in data at rest and in transit tactical and design wise | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | Low | Low | Local | High | High | None | Unchanged | 0.45653 | 0.52545 | 3.3734 | 3.9 | Moderate | 0.5 | 3.7 | LOW | 1. Implement server-side encryption using Service-Managed keys/recomended practise by azure. 2. Proper way of network access control 3. Encryption for sensitive data in transit, for ex: when files are moved to cloud storage, etc.. 4. Transfer over encrypted tunnel 5. Use strong encryption algorithm | 1. NA 2. SOM D001020115 - 05. Access control policy and management 3. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 4. NA 5. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
65 | T08 | Information disclosure (STR(I)DE) | V18 | Weak Encryption Implementaion in data at rest and in transit tactical and design wise | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | 1. Statefull firewall 2. Configure and upgrade routers for the n/w security 3. Configure firewalls to reject any packets with spoofed addresses. 4. Use secure tunnel communication channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SOM D001020115 - 16. Transmission confidentiality and integrity 3. SOM D001020115 - 23. Malware Detection/Protection 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
66 | T08 | Information disclosure (STR(I)DE) | V19 | Weak Algorthim implementation with respect cipher key size | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | Low | Low | Local | High | High | None | Unchanged | 0.45653 | 0.52545 | 3.3734 | 3.9 | Moderate | 0.5 | 3.7 | LOW | 1. Weak algorithms such as DES, RC4, etc.. should be avoided and usage of strong algorithms such as AES, RSA, etc.. are recomended 2. Typical key lengths are 128 and 256 bits for private keys and 2048 for public keys are recommended. | 1. SRS D001020097 –2.8.1Application shall use APIs to communicate between browser application and the backend. 2. SRS D001020097 –2.8.1Application shall use APIs to communicate between browser application and the backend. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
67 | T08 | Information disclosure (STR(I)DE) | V19 | Weak Algorthim implementation with respect cipher key size | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | 1. Weak algorithms such as DES, RC4, etc.. should be avoided and usage of strong algorithms such as AES, RSA, etc.. are recomended 2. Typical key lengths are 128 and 256 bits for private keys and 2048 for public keys are recommended. | 1. SRS D001020097 –2.8.1Application shall use APIs to communicate between browser application and the backend. 2. SRS D001020097 –2.8.1Application shall use APIs to communicate between browser application and the backend. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
68 | T08 | Information disclosure (STR(I)DE) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | Low | Low | Network | High | High | None | Unchanged | 0.70555 | 0.52545 | 3.3734 | 4.1 | Moderate | 0.5 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.NA 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
69 | T08 | Information disclosure (STR(I)DE) | V14 | Unencrypted Network segment through out the information flow | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | 1. Anonymization/Pseudomyzation of patient details 2. Data encyrption 3. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 4. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 3. Mainitaing Access Logs 4. Maintain Server Security Logs | 1.SRS D001020023- 2.1.6.21The application shall allow to assign and edit patient reference ID to patient. 2. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 3. NA 4. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
70 | T08 | Information disclosure (STR(I)DE) | V05 | Insecure communications in networks (hospital) | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | Low | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | 1. Secure communication with Secure Sockets Layer (SSL) or TLS protocols that provide message confidentiality 2. Secure sensitive data in the channel flow using strong encryption 3. Statefull firewall 4. Proper way of network access control | 1. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel 2. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 3. SOM D001020115 - 23. Malware Detection/Protection 4. SOM D001020115 - 05. Access control policy and management | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
71 | T09 | Data Access (STR[I]DE) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | None | None | High | Network | High | High | None | Unchanged | 0.70555 | 0.56 | 3.5952 | 4.4 | Low | 0.2 | 3.8 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.NA 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
72 | T09 | Data Access (STR[I]DE) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | None | Low | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | Low | 0.2 | 2.7 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.NA 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
73 | T09 | Data Access (STR[I]DE) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A09 | Data at Rest | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | Low | Low | Low | Network | High | High | None | Unchanged | 0.70555 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Statefull firewall 3. Do not store sensitive data in plaintext. 4. Use strong encrption algorithm. 5. Apply salting over sensitive data. | 1. NA 2. SOM D001020115 - 23. Malware Detection/Protection 3. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 4. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 5. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
74 | T09 | Data Access (STR[I]DE) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authentication/Authorisation method of all device(s)/app | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | None | None | Network | High | High | None | Unchanged | 0.70555 | 0.56 | 3.5952 | 4.4 | Moderate | 0.5 | 4 | MEDIUM | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Statefull firewall 3. Do not store sensitive data in plaintext. 4. Use strong encrption algorithm. 5. Apply salting over sensitive data. | 1. NA 2. SOM D001020115 - 23. Malware Detection/Protection 3.SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 4. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 5. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
75 | T09 | Data Access (STR[I]DE) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A10 | Data in Transit | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | High | None | None | Network | High | High | None | Unchanged | 0.70555 | 0.56 | 3.5952 | 4.4 | Low | 0.2 | 3.8 | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Statefull firewall 3 Do not store sensitive data in plaintext. 4. Use strong encrption algorithm. 5. Apply salting over sensitive data. | 1. NA 2. SOM D001020115 - 23. Malware Detection/Protection 3. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 4. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 5. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
76 | T09 | Data Access (STR[I]DE) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A09 | Data at Rest | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | Low | Low | Low | Network | High | High | None | Unchanged | 0.70555 | 0.52545 | 3.3734 | 4.1 | Low | 0.2 | 3.6 | LOW | 1. Strong password strength practices is recommended for admin web app. 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods | 1. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 2. .SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5. NA 6. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
77 | T09 | Data Access (STR[I]DE) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. 3) Information related to authenication/authorisation data (credential/pins/MFA/Biometrics) |  | None | Low | Low | Physical | High | High | None | Unchanged | 0.16601 | 0.3916 | 2.5141 | 2.7 | Moderate | 0.5 | 2.6 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.NA 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
78 | T10 | Open network port exploit (TTP) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | None | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.22 | 1.4124 | 3.1 | Moderate | 0.5 | 2.3 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.NA 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
79 | T10 | Open network port exploit (TTP) | V12 | Unprotected network port(s) on network devices and connection points | A08 | Wireless Network device (Scope of HDO) | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | None | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.22 | 1.4124 | 3.1 | Moderate | 0.5 | 2.3 | LOW | SOM responsibility 1. Statefull firewall and using port security on fiewall 2. Configure router wireless security options based on manufacture manual | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SOM D001020115 -11. Configuration settings | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
80 | T10 | Open network port exploit (TTP) | V14 | Unencrypted Network segment through out the information flow | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | None | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.22 | 1.4124 | 3.1 | Moderate | 0.5 | 2.3 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.NA 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
81 | T10 | Open network port exploit (TTP) | V17 | Unencrypted data in transit in all flowchannels | A10 | Data in Transit | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. |  | None | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.22 | 1.4124 | 3.1 | Moderate | 0.5 | 2.3 | LOW | 1. Use secure tunnel communication channel 2. Configure and upgrade routers for the n/w security 3. Configure firewalls to reject any packets with spoofed addresses. 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. For sensitive data proper encryption mechanism needs to be designed & implemented | 1. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel 2. SOM D001020115 - 16. Transmission confidentiality and integrity 3. SOM D001020115 - 23. Malware Detection/Protection 4. SOM D001020115 - 05. Access control policy and management 5. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
82 | T10 | Open network port exploit (TTP) | V05 | Insecure communications in networks (hospital) | A02 | Tablet OS/network details & Tablet Application | 1) Capture your account’s user ID and credentials. 2) Log the data of online traffic accessed on your tablet or computer. In this way, they can maintain a data of the websites you mostly visit, and plan attack from these websites. 3) Gain access to your computer, its network and data. 4) Launch a spam or malware attack on your device. |  | None | Low | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | Moderate | 0.5 | 3.4 | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2.SRS D001020024 -2.17.6The Application shall support the use of anti-malware mechanism 3.NA 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
83 | T11 | Brute-force Attack (CAPEC-112) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A11 | Smart medic app (Stryker Admin Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | None | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | Moderate | 0.5 | 2.9 | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs | 1. NA 2. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
84 | T11 | Brute-force Attack (CAPEC-112) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | None | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | Moderate | 0.5 | 2.9 | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs | 1. NA 2. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
85 | T11 | Brute-force Attack (CAPEC-112) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A11 | Smart medic app (Stryker Admin Web Application) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | None | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | Moderate | 0.5 | 2.9 | LOW | 1. Strong password strength practices is recommended for admin web app. 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods | 1. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 2. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5. NA 6. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
86 | T11 | Brute-force Attack (CAPEC-112) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | None | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | Moderate | 0.5 | 2.9 | LOW | 1. Strong password strength practices is recommended in azure 2. Require multi-factor authentication 3. limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods | 1. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 2. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha SRS D001020097 – 2.1.2.6 The Application shall be validated by using invisible captcha during login. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. NA 5. NA 6. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
87 | T11 | Brute-force Attack (CAPEC-112) | V18 | Weak Encryption Implementaion in data at rest and in transit tactical and design wise | A09 | Data at Rest | 1) An attacker may attempt to discover a weak encryption by systematically trying every possible combination of decryption key. |  | Low | None | Low | Local | High | High | None | Unchanged | 0.45653 | 0.3916 | 2.5141 | 3 | Moderate | 0.5 | 2.8 | LOW | 1. Implement server-side encryption using Service-Managed keys (provided for azure) 2. Proper way of network access control 3. Encryption for sensitive data in transit, for ex: when files are moved to cloud storage, etc.. 4. Transfer over encrypted tunnel 5. Use strong encryption algorithm | 1. SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 2. SOM D001020115 - 05. Access control policy and management 3.SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel 5.SRS D001020097 – 2.1.2.1On the Admin Login Screen, the Application shall allow entry of Email and Password which has a maximum of 50 chars and has a feature to hide the code in asterisk form. SRS D001020024 – 2.1.2.1 On agreeing to the Terms and Conditions, the application shall allow the user to enter the Admin Pin which has maximum 6 chars and a feature to hide the code. SRS D001020023- 2.1.2.1 Application shall allow the user to enter Hospital Code. The code has a maximum of 9 chars and a feature to hide it using asterisk form. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
88 | T11 | Brute-force Attack (CAPEC-112) | V18 | Weak Encryption Implementaion in data at rest and in transit tactical and design wise | A10 | Data in Transit | 1) An attacker may attempt to discover a weak encryption by systematically trying every possible combination of decryption key. |  | Low | None | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | Moderate | 0.5 | 2.9 | LOW | 1. Statefull firewall 2. Configure and upgrade routers for the n/w security 3. Configure firewalls to reject any packets with spoofed addresses. 4. Use secure tunnel communication channel | 1. SOM D001020115 - 23. Malware Detection/Protection 2. SOM D001020115 - 16. Transmission confidentiality and integrity 3. SOM D001020115 - 23. Malware Detection/Protection 4. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
89 | T12 | Social Engineering (TTP) | V22 | Legacy system identification if any | A11 | Smart medic app (Stryker Admin Web Application) | 1) This threat may hamper digital or physical resources, infractructure and end points 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. |  | None | Low | High | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.6568 | 4.2167 | 4.6 | Moderate | 0.5 | 4.5 | MEDIUM | Stryker IT team responsibility 1. Delete any request for personal information 2. Statefull firewall 3. Disable device network discoverable 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. | (SRS REQUIREMENT FOR STRYKER IT TEAM) | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
90 | T12 | Social Engineering (TTP) | V04 | Checking authentication modes for possible hacks and bypasses | A07 | Interface/API Communication | 1) This threat may hamper digital or physical resources, infractructure and end points through spear phishing mail 2) Get the user (employee/ client/ customer) to download malware, send money or perform actions that are dangerous. 3) Reputational harm 4) Economical harm |  | None | Low | High | Adjacent Network | High | High | Required | Unchanged | 0.37538 | 0.6568 | 4.2167 | 4.6 | Moderate | 0.5 | 4.5 | MEDIUM | Stryker IT team responsibility 1. Set your spam filters setting options to high 2. Restrict the system getting connected to unauthenticated sources 3. Restrict the system in accessing the 3rd party sites, social engineering sites, mails, etc.. 4. Configure firewalls to reject any packets with spoofed addresses. | (SRS REQUIREMENT FOR STRYKER IT TEAM) | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
91 | T13 | Lack of evidence to conclude any malicious attempt/attack (ST[R]IDE) | V28 | Insufficient Logging information | A12 | Smart medic app (Azure Portal Administrator) | 1) Adversary tried to obtain knowledge about system internals 2) Attempt to find attack vectors 3) Executed malicious activities 4) Complete details related to the attacker/malicious activities not recorded |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | Audit log: All the information needed for identifying the threat (malicious) activity and adversary information needed to be logged for determining the attack vector and attack surface. This helps to make the system less vulnerable in future by correcting those issues. | SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
92 | T13 | Lack of evidence to conclude any malicious attempt/attack (ST[R]IDE) | V29 | Insufficient Access permissions for accessing and modifying Log files | A12 | Smart medic app (Azure Portal Administrator) | 1) Adversary tried to obtain knowledge about system internals 2) Attempt to find attack vectors 3) Executed malicious activities 4) Complete details related to the attacker/malicious activities not recorded |  | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | Low | 0.2 | 3.8 | LOW | Audit log & security: 1. All the information needed for identifying the threat (malicious) activity and adversary information needed to be logged for determining the attack vector and attack surface. This helps to make the system less vulnerable in future by correcting those issues. 2. Audit log should be secured from unauthorized access. | 1. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 2.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
93 | T03 | Gaining Access ([S]TRID[E]) | V24 | Error Info containing sensitive data for Failed Authentication attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | 1. Standard error info messages should be used. Information aiding the attacks should be avoided. 2. Apart from user id there should be additional security factor for verification. 3. Limit on the login attempts is mandatory. | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha. SRS D001020097 – 2.1.2.6The Application shall be validated by using invisible captcha during login. 2. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 3. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
94 | T03 | Gaining Access ([S]TRID[E]) | V30 | Improper security (for ex.,Storage & Access) for Key tokens and Certificates | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | 1. If database access using keys/certificates, their generation & storage should be done securely. 2. Apart from user id there should be additional security factor (keys/certificates) for verification. 3. Limit on the login attempts is mandatory. | 1. NA 2. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 3.SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
95 | T03 | Gaining Access ([S]TRID[E]) | V25 | Absence of additional security factor along with user identification | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | 1. Azure portal can be accessed by login credentials & MFA. Hence, strong password policies & management are required 2. If database access using keys/certificates, their generation & storage should be done securely. 3. Apart from user id there should be additional security factor for verification. 4. Limit on the login attempts is mandatory. | 1. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 2. NA 3. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 4. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
96 | T03 | Gaining Access ([S]TRID[E]) | V25 | Absence of additional security factor along with user identification | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | 1. If database access provided & also using keys/certificates, their generation & storage should be done securely. 2. Apart from user id there should be additional security factor (keys/certificates) for verification. 3. Limit on the login attempts is mandatory. | 1. NA 2. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 3. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
97 | T11 | Brute-force Attack (CAPEC-112) | V24 | Error Info containing sensitive data for Failed Authentication attempts | A13 | Azure Cloud DataBase | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | 1. Standard error info messages should be used. Information aiding the attacks should be avoided. 2. Apart from user id there should be additional security factor for verification. 3. Limit on the login attempts is mandatory. | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha. SRS D001020097 – 2.1.2.6The Application shall be validated by using invisible captcha during login. 2. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 3. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
98 | T11 | Brute-force Attack (CAPEC-112) | V26 | Having no limit on the login attempts | A12 | Smart medic app (Azure Portal Administrator) | 1) An attacker may attempt to discover a weak credential by systematically trying every possible combination of letters, numbers, and symbols until it discovers the one correct combination that works. |  | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Moderate | 0.5 | 5.6 | MEDIUM | 1. Standard error info messages should be used. Information aiding the attacks should be avoided. 2. Apart from user id there should be additional security factor for verification. 3. Limit on the login attempts is mandatory. | 1. SRS D001020023-2.1.2.2 If the Hospital Code is valid, then on pressing the PROCEED button, the application shall be validated by the invisible captcha. SRS D001020097 – 2.1.2.6The Application shall be validated by using invisible captcha during login. 2. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. 3. SRS D001020097 – 2.1.2.1.1 Invalid email or password, only 3 attempts left. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
99 | T14 | Unauthorized Alterations (S[T]RIDE) | V31 | Improper/insufficient provisioning of IOT hub | A02 | Tablet OS/network details & Tablet Application | 1. If provisioning got failed/misleaded then the complete functionality gets affected. 2. Proper reason for the provisioning failure needs to addressed |  | None | None | High | Network | High | High | None | Unchanged | 0.70555 | 0.56 | 3.5952 | 4.4 | Low | 0.2 | 3.8 | LOW | 1. All devices which need to be registered & gets communicated in iot hub needs to be identified 2. Proper provisioning needs to be established with secure data transfer during provisioning 3. End-to-end provisioning with iot hub needs to be carried out without getting failed 4. All failure cases during the provisioning need to be documented/addressed | 1. NA 2. NA 3. NA 4. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
100 | T14 | Unauthorized Alterations (S[T]RIDE) | V32 | Unsecured communication with unauthenticated 3rd party devices | A02 | Tablet OS/network details & Tablet Application | If there is no proper authentication between the devices in the smart medic environment then 3rd party devices can easily establish the communication with the stryker devices |  | None | None | High | Network | High | High | None | Unchanged | 0.70555 | 0.56 | 3.5952 | 4.4 | Low | 0.2 | 3.8 | LOW | 1. In complete smart medic environment, only authenticated stryker and HDO devices need to be present 2. Secure communication between the stryker devices needs to be established & documented 3. Handling of the unauthenticated 3rd party devices trying to communicate with stryker devices needs to be taken care | 1. NA 2. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel 3. SRS D001020024 -2.17.2The Application shall provide secure tunnel Communications channel | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | LOW
2 | T01 | Deliver undirected malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
3 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
4 | T01 | Deliver undirected malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
5 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
6 | T01 | Deliver undirected malware (CAPEC-185) | V22 | Legacy system identification if any | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
7 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
8 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
9 | T01 | Deliver undirected malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
10 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
11 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
12 | T01 | Deliver undirected malware (CAPEC-185) | V09 | Lack of plan for periodic Software Vulnerability Management | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
13 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
14 | T01 | Deliver undirected malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
15 | T01 | Deliver undirected malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
16 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in transit in all flowchannels | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
17 | T01 | Deliver undirected malware (CAPEC-185) | V17 | Unencrypted data in transit in all flowchannels | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
18 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
19 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
20 | T01 | Deliver undirected malware (CAPEC-185) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
21 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | MEDIUM | Justification
22 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
23 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
24 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A08 | Wireless Network device (Scope of HDO) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | SOM responsibility 1. Statefull Firewall 2. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present.
25 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
26 | T02 | Deliver directed malware (CAPEC-185) | V13 | Unprotected external USB Port on the tablet/devices. | A11 | Smart medic app (Stryker Admin Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
27 | T02 | Deliver directed malware (CAPEC-185) | V02 | External communications and exposure for communciation channels from and to application and devices like tablet and smartmedic device. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | 1. Only stryker made/authenticated devices should communicate with smart medic device & tablet 2. Asset should be behind stateful firewall 3. Use secure tunnel communications channel
28 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
29 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
30 | T02 | Deliver directed malware (CAPEC-185) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
31 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A03 | Smart medic (Stryker device) System Component | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | MEDIUM | 1. Only Stryker/HDO authenticated devices should communicate with smart medic device & tablet 2. Asset should be behind stateful firewall 3. Use secure tunnel communications channel
32 | T02 | Deliver directed malware (CAPEC-185) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
33 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A11 | Smart medic app (Stryker Admin Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Deployed (V&V) secure system configuration model needs to be mentioned in the installation manual. 2. Deploy system configuration management tool 3. Implement automated configuration monitoring systems 4. Establish internal and external information sources for threat intelligence and vulnerability data, monitoring them regularly and taking appropriate action for high-priority items 5. Use upgraded software, firmware 6 Never create/use credentials with personal details such as date of birth, spouse, or child’s or pet’s name 7. Stateful Firewall
34 | T02 | Deliver directed malware (CAPEC-185) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
35 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | LOW
36 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A02 | Tablet OS/network details & Tablet Application | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
37 | T02 | Deliver directed malware (CAPEC-185) | V16 | Unencrypted data at rest in all possible locations | A11 | Smart medic app (Stryker Admin Web Application) | 1) Malicious utilization of computer resources 2) computing power 3) denial of service attacks, 4) ransomware attack 5) Bitcoin mining, etc |  | LOW | 1. Identification of the sensitive data in storage and encryption of storage subsystem 2. Stateful firewall 3. Hardening of the host system containing sensitive data at rest 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. Use strong encrption algorithm
38 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
39 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A11 | Smart medic app (Stryker Admin Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | 1. Admin application can be accessed by login credentials & MFA. Hence, strong password policies & management are required 2. Data transfer between the admin application and the smart medic components needs to be encrypted & secured. 3. Any vulnerable network ports and connection points should be identified and hardened. 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. Stateful firewall
40 | T03 | Gaining Access ([S]TRID[E]) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel | LOW
41 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | 1.During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Maintain Access Logs
42 | T03 | Gaining Access ([S]TRID[E]) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A07 | Interface/API Communication | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods
43 | T03 | Gaining Access ([S]TRID[E]) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | MEDIUM | 1. If devices/apps being accessed by login credentials & MFA. Then, strong password policies & management are required 2. Require multi-factor authentication 3 Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 5 Maintain Server Security Logs 6. Stronger authentication methods
44 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. Encrypt the authentication data in storage & transit to mitigate risk of information disclosure and authentication protocol attacks 2. Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks 3. Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks 4. Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
45 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A11 | Smart medic app (Stryker Admin Web Application) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. Encrypt the authentication data in storage & transit to mitigate risk of information disclosure and authentication protocol attacks 2. Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks 3. Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks 4. Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
46 | T03 | Gaining Access ([S]TRID[E]) | V04 | Checking authentication modes for possible hacks and bypasses | A12 | Smart medic app (Azure Portal Administrator) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. Encrypt the authentication data in storage & transit to mitigate risk of information disclosure and authentication protocol attacks 2. Encrypt authentication data using non reversible encryption such as using a digest (e.g., HASH) and a seed to prevent dictionary attacks 3. Lock out accounts after reaching a log on failure threshold and mitigate risk of brute force attacks 4. Display generic error messages upon validation of credentials to mitigate risk of account harvesting or enumeration
47 | T03 | Gaining Access ([S]TRID[E]) | V13 | Unprotected external USB Port on the tablet/devices. | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
48 | T04 | Maintaining Access (TTP) | V01 | Devices with default passwords needs to be checked for bruteforce attacks | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. During the access providing, if default password is provided then immediately changing the password is needed. Also ensure: 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods
49 | T04 | Maintaining Access (TTP) | V03 | The password complexity or location vulnerability. Like weak passwords and hardcoded passwords. | A04 | Authentication/Authorisation method of all device(s)/app | 1) Obtain knowledge about system internals 2) Attempt to find attack vectors 3) Possibilities for exploitation of publicly known Vulnerabilities. |  | LOW | 1. If devices/apps being accessed by login credentials & MFA. Then, strong password policies & management are required 2. Require multi-factor authentication 3. Limit authentication attempts (rate Limiting) 4. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 5. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 4. Maintain Access Logs 5. Maintain Server Security Logs 6. Stronger authentication methods
50 | T05 | Clearing Track (TTP) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
51 | T05 | Clearing Track (TTP) | V23 | Outdated - Software/Hardware | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
52 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
53 | T05 | Clearing Track (TTP) | V07 | Lack of configuration controls for IT assets in the informaion system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
54 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
55 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
56 | T05 | Clearing Track (TTP) | V08 | Ineffective patch management of firware, OS and applications thoughout the information system plan | A02 | Tablet OS/network details & Tablet Application | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
57 | T05 | Clearing Track (TTP) | V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | A05 | Device Maintainence tool (Hardware/Software) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
58 | T05 | Clearing Track (TTP) | V10 | The static connection digaram between devices and applications with provision for periodic updation as per changes | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Tampering of forensic data 2) This involves modifying/corrupting/deleting the values of Logs, 3) Modifying registry values 4) Uninstalling all malcious applications/tools 5) Deleting all folders which were created |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
59 | T06 | Elevation of privilege (STRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A04 | Authentication/Authorisation method of all device(s)/app | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | LOW | 1. Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. 2. Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. 3. Ensure default credentials not existing for any assets (such as applications, operating systems, routers, firewalls, wireless access points).
60 | T06 | Elevation of privilege (STRID[E]) | V15 | Controlled Use of Administrative Privileges over the network | A12 | Smart medic app (Azure Portal Administrator) | 1) Gaining access to the portal 2) Accessing confidential data, 3) Lead misuse of confidential data 4) Company defamation |  | MEDIUM | 1. Require that administrators establish multi factor authentication for their administrator and non-administrative accounts. 2. Access to a machine (either remotely or locally) should be blocked for administrator-level accounts. 3. Ensure default credentials not existing for any assets (such as applications, operating systems, routers, firewalls, wireless access points).
61 | T07 | Denial of service (STRI(D)E) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Bring down the service availability 2) Blocking the end user usage |  | MEDIUM | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
62 | T08 | Information disclosure (STR(I)DE) | V16 | Unencrypted data at rest in all possible locations | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Identification of the sensitive data in storage and encryption of storage subsystem 2. Stateful firewall 3. Hardening of the host system containing sensitive data at rest 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. Use strong encrption algorithm
63 | T08 | Information disclosure (STR(I)DE) | V17 | Unencrypted data in transit in all flowchannels | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Use secure tunnel communication channel 2. Configure and upgrade routers for the n/w security 3. Configure firewalls to reject any packets with spoofed addresses. 4. Maintain access control (read/modify) permission list for any sensitive & unencrypted data if present. 5. For sensitive data proper encryption mechanism needs to be designed & implemented
64 | T08 | Information disclosure (STR(I)DE) | V18 | Weak Encryption Implementaion in data at rest and in transit tactical and design wise | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Implement server-side encryption using Service-Managed keys/recomended practise by azure. 2. Proper way of network access control 3. Encryption for sensitive data in transit, for ex: when files are moved to cloud storage, etc.. 4. Transfer over encrypted tunnel 5. Use strong encryption algorithm
65 | T08 | Information disclosure (STR(I)DE) | V18 | Weak Encryption Implementaion in data at rest and in transit tactical and design wise | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Statefull firewall 2. Configure and upgrade routers for the n/w security 3. Configure firewalls to reject any packets with spoofed addresses. 4. Use secure tunnel communication channel
66 | T08 | Information disclosure (STR(I)DE) | V19 | Weak Algorthim implementation with respect cipher key size | A09 | Data at Rest | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Weak algorithms such as DES, RC4, etc.. should be avoided and usage of strong algorithms such as AES, RSA, etc.. are recomended 2. Typical key lengths are 128 and 256 bits for private keys and 2048 for public keys are recommended.
67 | T08 | Information disclosure (STR(I)DE) | V19 | Weak Algorthim implementation with respect cipher key size | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Weak algorithms such as DES, RC4, etc.. should be avoided and usage of strong algorithms such as AES, RSA, etc.. are recomended 2. Typical key lengths are 128 and 256 bits for private keys and 2048 for public keys are recommended.
68 | T08 | Information disclosure (STR(I)DE) | V21 | InSecure Configuration for Software/OS on Mobile Devices, Laptops, Workstations, and Servers | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
69 | T08 | Information disclosure (STR(I)DE) | V14 | Unencrypted Network segment through out the information flow | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Anonymization/Pseudomyzation of patient details 2. Data encyrption 3. Audit log - Maintain Access logs (login (attempted & failed), logoff, id change) 4. Audit log - Maintain security logs (such as change/modification of system configuration settings, services, etc.) 3. Mainitaing Access Logs 4. Maintain Server Security Logs
70 | T08 | Information disclosure (STR(I)DE) | V05 | Insecure communications in networks (hospital) | A10 | Data in Transit | Information of health data can be exploit and disclose with various means like network, tablet etc. . |  | LOW | 1. Secure communication with Secure Sockets Layer (SSL) or TLS protocols that provide message confidentiality 2. Secure sensitive data in the channel flow using strong encryption 3. Statefull firewall 4. Proper way of network access control
71 | T09 | Data Access (STR[I]DE) | V12 | Unprotected network port(s) on network devices and connection points | A01 | Tablet Resources - web cam, microphone, OTG devices, Removable USB, Tablet Application, Network interfaces (Bluetooth, Wifi) | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
72 | T09 | Data Access (STR[I]DE) | V12 | Unprotected network port(s) on network devices and connection points | A02 | Tablet OS/network details & Tablet Application | 1) Allowing application or script to perform abnormal activites on the system. 2) Modifying the data, tampering the confidential data making it unavailable or challenging the integrity of data. |  | LOW | 1. Asset should be behind stateful firewall 2. Anti-virus with updated virus definitions 3. Audit log capturing any abnormal activity identified/reported by the application 4. Use hardened interfaces (n/w) & secure tunnel communications channel
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls