## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Knee Balancer
Scope:
Date: |  | <yyyyy-mm-dd>
Conducted by: |  | <Author Name / Function / Organization> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description | Comments
A01 | Physical Asset | Mobile Device (ipad, iphone V8-13)
A02 | Information asset | Login (4-digit pin) management
A03 | Information asset | Knee balancer application | Contains Knee Planning software which controls the overall features for segmentation, landmark planning, implant size calculation etc.
A04 | Information asset | Patient Identifiable data | DICOM data stored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A05 | Information asset | MAC/iOS operating system | Information about internals of the system (Device identification, software versions, supported protocols, etc.) | Outside of Trust Boundary. May include guidance in SOM for this asset, if appriopriate
A06 | Information asset | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..)
A07 | Physical Asset | MAKO System | Network location where the CT Dicom data is stored | Outside of Trust Boundary. May include guidance in SOM for this asset, if appriopriate
A08 | Information Asset | Application Settings | Application specific configuration settings
A09 | Information Asset | User Settings | User specific configuration settings
A10 | Information Asset | Log Files | Log files containing application errors and debug logs
A11 | Information asset | User License data | Application instance specific license data
A12 | Information asset | Security keys, tokens, certificates | Encryption keys and tokens
A13 | Information asset | Wireless communication interfaces (Wifi, BT,etc..)
A14 | Information asset | Physical network interfaces
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
D05788-1, Ver 1

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user credentials | Yes | n/a
V02 | Improper Authentication (single factor, no captcha, etc..) | Yes | n/a
V03 | Improper Input Validation (allowed range) | Yes | n/a
V04 | Improper Storage of Credentials | Yes | n/a
SBOM/SOUP
V11 | Unpatched operating system | Yes | Security Operations Manual
V12 | Unsafely compiled Application Binaries - App | Yes | n/a
V13 | Ineffective patch management | Yes | Security Operations Manual
V14 | Lack of OS vulnerability management | Yes | Security Operations Manual
V15 | 3rd Party Component Dependency & Vulnerabilities - App | Yes | n/a
V16 | Corrupting the binaries/injecting malicious source code - App | Yes | n/a
V17 | Static Analysis for developed code & SCA for SOUP - App | Yes | n/a
V18 | 3rd party installed applications from app store | Yes | n/a
Application Data Management
V21 | Unencrypted data at rest in all possible locations (Device, Cloud) | Yes | n/a
V22 | Unencrypted data in transit in all flowchannels | Yes | n/a
V23 | Weak Encryption Implementaion for data at rest and in motion tactical and design wise or Inadequate encryption strength | Yes | n/a
V24 | Use of a Broken or Risky Cryptographic Algorithm | Yes | n/a
V25 | Insecure Storage of Sensitive Information | Yes | n/a
V26 | Absence of application (status/behavior) logs and audit logs | Yes | n/a
V27 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging | Yes | n/a
V28 | Information Exposure Through Log Files | Yes | n/a
Application/Communication Management
V31 | Untrusted 3rd party applications installation from appstore | Yes | n/a
V32 | No security applications (antivirus, etc..) running in the ipad/iphone | Yes | n/a
V33 | Unrestricted Wifi communication with open/untrusted sources | Yes | n/a
V34 | Unrestricted BT communication/transfer with multiple sources | Yes | n/a
V35 | Unauthorized device communication to network interfaces | Yes | n/a
InSecure Configurations of Resources
V40 | Lack of efficient configuration change control - App | Yes | n/a
V41 | Malicious Software Download | No | n/a
V42 | Malicious Automated Software Update | No | n/a
V43 | External Control of System or Configuration Setting | Yes | n/a
V44 | Unsecured configuration of wireless devices | Yes | n/a
Cloud Backend
V51 | Unsecurely transmitted data to cloud | No | n/a
V52 | Unsecurely encrypting the database | No | n/a
V53 | Single Factor Authentication to cloud | No | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware | For ipad/iphone, the charging/USB interface is provided & this can be used to connect the OTG devices. | TSA-1 - Individual (Employees, Insider, Trusted Insider/Priveleged Insider) TSA-3 - Skript Kiddies | Yes | n/a
T02 | Deliver directed malware | Threat source delivers malware on a removable media which was designed to exploit a known vulnerability of the Software. Directed attack on the Application using knowledge about the Application. | TSA-2 Organization | Yes | n/a
T03 | Perform wireless perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies | Yes | n/a
T04 | Unauthorized modifications to Knee Balancer Software | Adversary gains unauthorized access to the system and removes or modifies application binaries. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T05 | Unauthorized access to Knee Balancer App | Adversary gains unauthorized access to the system and removes or modifies application configuration. | TSA-3 - Skript Kiddies | Yes
T06 | Unauthorized access/modification to Pre-OP data/Treatment plan | Adversary gains unauthorized access to the treatment design plan of the patient. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes
T07 | Unauthorized access/modification to PHI/Health data | Adversary gains unauthorized access to the system and steals ePHI information | TSA-3 - Skript Kiddies- | Yes
T08 | Unauthorized access to app sensitive data (device's keys/certificates/crypto algorithms) | Adversary gains unauthorized access to the encryption keys and tokens | TSA-3 - Skript Kiddies | Yes
T09 | Brute force login attempts | Adversary gains unauthorized access to the system | TSA-3 - Skript Kiddies | Yes
T10 | Unprotected access to device settings (uninstalling applications, etc..) | Adversary gains unauthorized access to the system and make modifications to the software configuration files | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes
T11 | Unrestricted wireless communication with multiple sources
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V33 | Unrestricted Wifi communication with open/untrusted sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1. Virus Scan 2. Firewall 3. Windows security update install 4. Application should be updated with the latest version of dependencies. | [1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [1.3] Access management - Windows user MSI - SRS-2.22.1, 2.22.2, 2.10.3 SADD- PKG1, PKG2, PL01,ES05, ES06 [2.1] SOM- 13.3 Firewall Implementation [3.1] SOM - 8.3 Operating System Updates (CSUP-2) | [1.1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [1.3.1] PR - PY-SYK-CC-003 Application Dependency Check [2.1.1] SOM- 13.3 Firewall Implementation [3.1.1] SOM - 8.3 Operating System Updates (CSUP-2) | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
2 | T01 | Deliver undirected malware | V34 | Unrestricted BT communication/transfer with multiple sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  |  |  |  | Network |  |  |  |  | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T01 | Deliver undirected malware | V35 | Unauthorized device communication to network interfaces | A14 | Physical network interfaces | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  |  |  |  | Network |  |  |  |  | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T01 | Deliver undirected malware | V15 | 3rd Party Component Dependency & Vulnerabilities - App | A04 | Patient Identifiable data | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | Low | Low | Low | Network | Low | None | None | Unchanged | 3.887 | 0.52545 | 3.3734 | 7.3 | Low | 0.2 | 4.2 | MEDIUM | 1. The 3rd party modules and dependencies should be updated with latest versions. 2. Application upgrade should verify if all required dependencies are updated in system. 3. Application upgrade should detect any upgrade or change in dependecies on user system and show dependency issue message to application user. | [1.1] SDP - D005010020 - 5.2.4 [2.1] SDP - D005010020 - 5.2.4 [3.1] SDP - D005010020 - 5.2.4 | [1.1.1] Pentest Report [2.1.1] Pentest Report [3.1.1] Pentest Report | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T01 | Deliver undirected malware | V18 | 3rd party installed applications from app store | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T02 | Deliver directed malware | V33 | Unrestricted Wifi communication with open/untrusted sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | By knowing the end-to-end functionality of the application, the adversary can exploit the application and try to compromise the system with specific malware attacks |  |  |  |  | Network |  |  |  |  | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T02 | Deliver directed malware | V34 | Unrestricted BT communication/transfer with multiple sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | By knowing the end-to-end functionality of the application, the adversary can exploit the application and try to compromise the system with specific malware attacks |  |  |  |  | Network |  |  |  |  | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T02 | Deliver directed malware | V35 | Unauthorized device communication to network interfaces | A14 | Physical network interfaces | By knowing the end-to-end functionality of the application, the adversary can exploit the application and try to compromise the system with specific malware attacks |  |  |  |  | Network |  |  |  |  | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T03 | Perform wireless perimeter network reconnaissance/scanning. |  | #N/A | A14 | Physical network interfaces | Loss of Stryker Reputation, potential loss of revenue |  | High | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.7323 | 4.7014 | 5.3 | Low | 0.2 | 4.9 | MEDIUM | 1. Firewall 2. The internet communication of Thor application should be secured using SSL or other secure protocols. 3. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. 4. The License should be encrypted for communication over internet. | [1.1] SOM- 13.3 Firewall Implementation [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable as no internet connection is getting used. [4.1]License file encryption is done using AES SRS:2.17.13, 2.17.4 SADD: ES04, DS10, C277 | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] Not applicable as no internet connection is getting used. [4.1.1] License file encryption is done using AES [4.1.2] PR - PY-SYK-CC-010 Use of Secure Crypto | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T03 | Perform wireless perimeter network reconnaissance/scanning. |  | #N/A | A15 | #N/A | Loss of Stryker Reputation, loss of Stryker IP, potential loss of PHI |  | High | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.7323 | 4.7014 | 5.3 | Low | 0.2 | 4.9 | MEDIUM | 1. The internet communication of Thor application should be secured using SSL or other secure protocols. 2. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. 3. The keys and token should be stored using windows credential manager. | [1.1] Not applicable as no internet connection is getting used. [2.1] Not applicable as no internet connection is getting used. [3.1] | [1.1.1] Not applicable as no internet connection is getting used. [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] PR - PY-SYK-CC-001 User Authentication [3.1.2] PR - PY-SYK-CC-010 Use of Secure Crypto | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T04 | Unauthorized modifications to Knee Balancer Software | V12 | Unsafely compiled Application Binaries - App | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T04 | Unauthorized modifications to Knee Balancer Software | V15 | 3rd Party Component Dependency & Vulnerabilities - App | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T04 | Unauthorized modifications to Knee Balancer Software | V16 | Corrupting the binaries/injecting malicious source code - App | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T04 | Unauthorized modifications to Knee Balancer Software | V17 | Static Analysis for developed code & SCA for SOUP - App | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T05 | Unauthorized access to Knee Balancer App | V16 | Corrupting the binaries/injecting malicious source code - App | A04 | Patient Identifiable data | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | Low | Low | Low | Physical | Low | High | Required | Unchanged | 0.21191 | 0.52545 | 3.3734 | 3.6 | Low | 0.2 | 3.5 | LOW | 1. Virus scan 2. Application should verify the checksum of all the DLL files and executable inside application folder before starting main thread. 3. Proper Code review of applciation before releasing. 4. Anti-dynamic instrumentation techniques should be implimented in application. | [1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1] Checksum of DLL: TBD [3.1] Code review: process followed. [4.1] Anti dynamic instrumentation: TBD | [1.1.1] SOM - 8.5 Anti-Malware Software Updates (CSUP-4) [1.2.1] SOM - 13. MALWARE DETECTION/PROTECTION (MLDP-1) [2.1.1] Checksum of DLL: TBD [3.1.1] Code review: process followed. [4.1.1] PR - PY-SYK-CC-007 Dynamic Instrumentation | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T05 | Unauthorized access to Knee Balancer App | V17 | Static Analysis for developed code & SCA for SOUP - App | A04 | Patient Identifiable data | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. Application should not have direct access to hidden functionalities. It is recommended to delete all unused funtionalities from application that could be there while development of application. 2. Code Review. 3. Application Testing | [1.1] Hidden functionalities: Not applicable [2.1] Code review: process followed. [3.1] Application testing : Done | [1.1.1] Hidden functionalities: Not applicable [2.1.1] Code review: process followed. [3.1.1] Application testing : Done | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T05 | Unauthorized access to Knee Balancer App | V31 | Untrusted 3rd party applications installation from appstore | A13 | Wireless communication interfaces (Wifi, BT,etc..) | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. The keys, password and token should be stored in encrypted form. Few of the sensitive data like keys should be stored using windows credential manager. | [1.1] Keys and tokens are kept in application binary, not stored on drive or configs. | [1.1.1] PR-PY-SYK-CC-008 Hard Coded Sensitive Information [1.1.2] PR-PY-SYK-CC-010 Use of Secure Crypto | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T05 | Unauthorized access to Knee Balancer App | V16 | Corrupting the binaries/injecting malicious source code - App | A03 | Knee balancer application | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | Low | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.52545 | 3.3734 | 4 | Low | 0.2 | 3.5 | LOW | 1. All log files should be stored encrypted . 2. The log files should not contain senstive user information and license information. | [1.1][2.1] Audit logs are encrypted SRS: 2.17.10, 2.17.4 SADD: DS13 SDD: ML11 | [1.1.1][2.1.1] PR-PY-SYK-CC-002 Proper Crash Logs are Maintained Locally in Encrypted Form [1.1.2][2.1.2] PR-PY-SYK-CC-010 Use of Secure Crypto | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T05 | Unauthorized access to Knee Balancer App | V25 | Insecure Storage of Sensitive Information | A03 | Knee balancer application | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T06 | Unauthorized access/modification to Pre-OP data/Treatment plan | V01 | Ineffective management of user credentials | A03 | Knee balancer application | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | Low | Low | Local | Low | High | Required | Unchanged | 0.58275 | 0.7323 | 4.7014 | 5.3 | Low | 0.2 | 4.9 | MEDIUM | 1. The keys, password and token should be stored in encrypted form.Few of the sensitive data like keys should be stored using windows credential manager. | [1.1] Keys and tokens are kept in application binary, not stored on drive or configs. | [1.1.1] PR-PY-SYK-CC-008 Hard Coded Sensitive Information [1.1.2] PR - PY-SYK-CC-010 Use of Secure Crypto | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T06 | Unauthorized access/modification to Pre-OP data/Treatment plan | V02 | Improper Authentication (single factor, no captcha, etc..) | A03 | Knee balancer application | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | Moderate | 0.5 | 5.4 | MEDIUM | 1. Proper Access control should be implimented in application directory so that only authorised users can access application and its files. | [1.1] Handled through user id linked license, windows user directory. SRS: 2.17.13, 2.17.4 SADD: DS02, DS04, ES05, ES06 SDD: ML09 | [1.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T06 | Unauthorized access/modification to Pre-OP data/Treatment plan | V03 | Improper Input Validation (allowed range) | A01 | Mobile Device (ipad, iphone V8-13) | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | Low | 0.2 | 5 | MEDIUM | 1. The application should accept inputs from application UI interface only. 2. The application should verify input type before accepting it. Example, Date type of input should be in a specific format and should not conatain any other symbol or any alphabetic letter inside it. | [1.1] Handled in application input widgets. [2.1] SRS:2.5.4 SADD: EL06, EL09, ML02, UI02 SDD: C115, C202, C261, C262 | [1.1.1] PR - PY-SYK-CC-009 Command Line Arguments [2.1.1] PR - PY-SYK-CC-005 Input Validation Check | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T06 | Unauthorized access/modification to Pre-OP data/Treatment plan | V25 | Insecure Storage of Sensitive Information | A01 | Mobile Device (ipad, iphone V8-13) | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | Low | Low | Network | High | High | Required | Unchanged | 0.51463 | 0.7323 | 4.7014 | 5.3 | Low | 0.2 | 4.9 | MEDIUM | 1. Firewall 2. The internet communication of Thor application should be secured using SSL or other secure protocols. 3. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. 4. The License should be encrypted for communication over internet. | [1.1] SOM- 13.3 Firewall Implementation [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable as no internet connection is getting used. [4.1] Not applicable as no internet connection is getting used. | [1.1.1] SOM- 13.3 Firewall Implementation [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] Not applicable as no internet connection is getting used. [4.1.1] Not applicable as no internet connection is getting used. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T07 | Unauthorized access/modification to PHI/Health data | V21 | Unencrypted data at rest in all possible locations (Device, Cloud) | A04 | Patient Identifiable data | PHI/Health data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Moderate | 0.5 | 5 | MEDIUM | 1. Proper Access control should be implimented in application directory so that only authorised users can access application and its files. 2. The license should be verified everytime the application is started. 3. The license time period and authority should not get manipulated at client side by manupulating registery data or saved license data. 4. The license should be encrypted . 5. The treatment plan directory should have proper permissions only to authorized user . | [1.1] SOM-6.3 Audit Log Protection [1.2] SOM-7.2 Privilege and Access (AUTH-2, 3, 4, 5, SGUD-3, 3.1 [2.1] SOM-13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [3.1] SOM-16.1 Password/ID Assignments (PAUT-1) [4.1] SOM-13.2 Other Compensation/Protection Controls (MDLP-3, 4, 5) [4.2] Handled through user id linked license, windows user directory. License data is encrypted. SRS: 2.17.13, 2.17.4 SADD: DS02, DS04, ES05, ES06 SDD: ML09, ES04, DS10, C277 [5.1] SOM-7.2 Privilege and Access (AUTH-2, 3, 4, 5, SGUD-3, 3.1) | [1.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management [2.1.1] PY-SYK-CC-014 Licence Check [3.1.1] PR-PY-SYK-CC-013 Secure Configuration File [4.1.1] PR-PY-SYK-CC-010 Use of Secure Crypto [5.1.1] PR-PY-SYK-CC-006 Secured Application Path and Proper Access Management [5.1.2] PY-SYK-CC-013 Secure Configuration File | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T07 | Unauthorized access/modification to PHI/Health data | V25 | Insecure Storage of Sensitive Information | A04 | Patient Identifiable data | PHI/Health data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | Low | Low | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.7323 | 4.7014 | 5.2 | Low | 0.2 | 4.8 | MEDIUM | 1. The user credentials should not be stored unencrypted. 2. The user credentials /license should not be stored in path where treatment plan data is getting stored. | [1.1] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) | [1.1.1] PR-PY-SYK-CC-001 User Authentication [2.1.1] PR-PY-SYK-CC-001 User Authentication | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T08 | Unauthorized access to app sensitive data (device's keys/certificates/crypto algorithms) | V21 | Unencrypted data at rest in all possible locations (Device, Cloud) | A04 | Patient Identifiable data | If any of the cryptographic security sensitive data such as keys, certificates, crypto algorithms, etc.., got modified then complete security gets compromised |  | High | Low | None | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.6568 | 4.2167 | 4.8 | Low | 0.2 | 4.4 | MEDIUM | 1. The PHI data should be stored encrypted. 2. The PHI data directory should only be accessible to system user using the application. 3. The PHI data should not be stored in local SQLIte databases in plain text. 4. The PHI data should not be stored in Log files. | [1.1]PHI data stored in audit log file and case list is encrypted. [2.1]Directory acces is managed using windows user credential. [3.1]PHI data is not stored in local DB. [4.1]PHI data stored in audit log file and case list is encrypted. SRS: 2.17.10, 2.17.13, 2.17.4 SADD: DS13, DS02, DS04, ES05, ES06 SDD:ML11, C276, C277 | [1.1.1]PR - PY-SYK-CC-012 Secure PHI Data [2.1.1] PR - PY-SYK-CC-006 Secured Application Path and Proper Access Management [3.1.1] PR - PY-SYK-CC-012 Secure PHI Data [4.1.1] PR - PY-SYK-CC-012 Secure PHI Data | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T08 | Unauthorized access to app sensitive data (device's keys/certificates/crypto algorithms) | V25 | Insecure Storage of Sensitive Information | A04 | Patient Identifiable data | If any of the cryptographic security sensitive data such as keys, certificates, crypto algorithms, etc.., got modified then complete security gets compromised |  | High | Low | None | Physical | High | Low | Required | Unchanged | 0.27806 | 0.6568 | 4.2167 | 4.5 | Low | 0.2 | 4.3 | MEDIUM | 1. Firewall 2. The internet communication of Thor application should be secured using SSL or other secure protocols. 3. The application should not trust the system installed Trust certificates instead the application should verify SSL communication. | [1.1] SOM- 13.3 Firewall Implementation [2.1] Not applicable as no internet connection is getting used. [3.1] Not applicable as no internet connection is getting used. | [1.1.1] SOM- 13.3 Firewall Implementation. [2.1.1] Not applicable as no internet connection is getting used. [3.1.1] Not applicable as no internet connection is getting used. | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T09 | Brute force login attempts | V23 | Weak Encryption Implementaion for data at rest and in motion tactical and design wise or Inadequate encryption strength | A12 | Security keys, tokens, certificates | Single factor authentication is not sufficient to surpress the brute force attacks. Minimum 2-factor authentication is needed to provide enough security from automated attacks |  | High | Low | None | Physical | Low | High | Required | Unchanged | 0.21191 | 0.6568 | 4.2167 | 4.5 | Moderate | 0.5 | 4.4 | MEDIUM | 1. The user encrypted license should be stored using windows credential manager. 2. The License data should be encrypted and should not be modified. | [1.1] SOM-16.2 User Account Management (PAUT-2, 3, 4, 5, 6, 7, 8, 9, 10, 11, 12, 13, 14) [2.1] SOM-3.5 System Security Context and Intended Environment (SGUD-4) Access management - Windows user, AES Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 | [1.1.1] PR-PY-SYK-CC-001 User Authentication | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T09 | Brute force login attempts | V25 | Insecure Storage of Sensitive Information | A12 | Security keys, tokens, certificates | Single factor authentication is not sufficient to surpress the brute force attacks. Minimum 2-factor authentication is needed to provide enough security from automated attacks |  | High | Low | None | Physical | High | High | Required | Unchanged | 0.12109 | 0.6568 | 4.2167 | 4.4 | Moderate | 0.5 | 4.3 | MEDIUM | 1. The application should use secure cryptographic algorithms to encrypt user Licence data. | Access management - Windows user, AES 256 Data encryption handler SRS: 2.17.3 SADD: ES05, ES06, DS02, DS04 SDD: C276, C277 [1.1] SOM - 3.5 System Security Context and Intended Environment (SGUD-4) | [1.1.1] PR-PY-SYK-CC-010 Use of Secure Crypto | [8 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T10 | Unprotected access to device settings (uninstalling applications, etc..) | V02 | Improper Authentication (single factor, no captcha, etc..) | A03 | Knee balancer application | Sensitive system settings on the host should be identified & protected from unauthorized access | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T11 | Unrestricted wireless communication with multiple sources | V43 | External Control of System or Configuration Setting | A01 | Mobile Device (ipad, iphone V8-13) | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T11 | Unrestricted wireless communication with multiple sources | V43 | External Control of System or Configuration Setting | A03 | Knee balancer application | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T11 | Unrestricted wireless communication with multiple sources | V44 | Unsecured configuration of wireless devices | A01 | Mobile Device (ipad, iphone V8-13) | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
| T11 | Unrestricted wireless communication with multiple sources | V44 | Unsecured configuration of wireless devices | A03 | Knee balancer application | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers | [9 empty] | #N/A | #N/A | #N/A | #N/A |  | #N/A | #N/A | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[6 empty] | Document references: | [18 empty] | SRS - CT knee SRS: D005010021
SAD - THOR: CT Knee Planning SOFTWARE_ARCHITECTURAL_DESIGN: D005010022
SDD - THOR: CT Knee Planning SOFTWARE_DETAILED_DESIGN: D005
PR- Pentest report
Note on AA5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav for 4th point "Application should be updated with the latest version of dependencies" add a refrence.
Note on AA8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: # Gaurav add the refrences properly. we have added the verififcation. Just go through it and create a cross reference. These points are technical so kindly see to it.
Note on AB8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav give a number to this point
Note on AA14: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: # Gaurav add reference for 3.1
Note on AA19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav Please add reference to point 2.1 and 4.1
Note on AA29: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav add reference for point 2
Note on AA32: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #Gaurav please add proper points to document references
Note on AA33: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: #gaurav Please provide reference for the point 1. The one given here is taking about the access management but it is related to license data. So please check if its applicable here

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V33 | Unrestricted Wifi communication with open/untrusted sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | MEDIUM | 1. Virus Scan 2. Firewall 3. Windows security update install 4. Application should be updated with the latest version of dependencies.
2 | T01 | Deliver undirected malware | V34 | Unrestricted BT communication/transfer with multiple sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host
3 | T01 | Deliver undirected malware | V35 | Unauthorized device communication to network interfaces | A14 | Physical network interfaces | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host
5 | T01 | Deliver undirected malware | V15 | 3rd Party Component Dependency & Vulnerabilities - App | A04 | Patient Identifiable data | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | MEDIUM | 1. The 3rd party modules and dependencies should be updated with latest versions. 2. Application upgrade should verify if all required dependencies are updated in system. 3. Application upgrade should detect any upgrade or change in dependecies on user system and show dependency issue message to application user.
6 | T01 | Deliver undirected malware | V18 | 3rd party installed applications from app store | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls