## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Knee Balancer
Scope:
Date: |  | <2022-02-25>
Conducted by: |  | <SIVA KUMAR /Security / LTTS> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description | Comments
A01 | Physical asset | Mobile Device (ipad, iphone V8-13) | Iphone/ipad is used for running the knee balancer application (currently v8-13 are supported)
A02 | Information asset | Login (4-digit pin) management | 4-digit numeric pin is used for authentication/login
A03 | Information asset | Knee balancer application | Contains Knee Planning software which controls the overall features for segmentation, landmark planning, implant size calculation etc.
A04 | Information asset | Patient Identifiable data | DICOM data stored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A05 | Information asset | MAC/iOS operating system | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A06 | Information asset | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | Knee balancer application is composed of 3rd party applications (binaries, frameworks, 3rd party components, etc..)
A07 | Physical asset | MAKO System | Network location where the CT Dicom data is stored
A08 | Information asset | Application Settings | Application specific configuration settings
A09 | Information asset | User Settings | User specific configuration settings
A10 | Information asset | Log Files | Log files containing application errors and debug logs
A11 | Information asset | User License data | Application instance specific license data
A12 | Information asset | Security keys, tokens, certificates | Keys,tokens and certificates
A13 | Information asset | Wireless communication interfaces (Wifi, BT,etc..) | Mobile interfaces for network communication
A14 | Physical asset | Physical charging ports/interfaces | Charging interface of mobile
A15 | Information asset | Data at Rest | Sensitive Data residing in Mobile
A16 | Information asset | Data in Transit | Sensitive data being transmitted
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user credentials | No | n/a
V02 | Improper Authentication (single factor, no captcha, etc..) | No | n/a
V03 | Improper Input Validation (allowed range) | Yes | n/a
V04 | Improper Storage of Stryker Credentials | Yes | Downloadable only from Stryker store. User is verified, no need of authentication
V05 | Weak/predictable 4-digit pin (0000, 1234, etc..) | No | n/a
V06 | Input from untrusted sources | Yes | n/a
SBOM/SOUP
V11 | Unpatched operating system | Yes | Security Operations Manual
V12 | Unsafely compiled Application Binaries in Application | Yes | n/a
V13 | Ineffective patch management | Yes | Security Operations Manual
V14 | Lack of OS vulnerability management | Yes | Security Operations Manual
V15 | 3rd Party Component Dependency & Vulnerabilities in Application | Yes | n/a
V16 | Corrupting the binaries/injecting malicious source code in Application | Yes | n/a
V17 | No Static Analysis for developed code & SCA for SOUP in Application | Yes | n/a
V18 | 3rd party installed applications from app store | Yes | n/a
Application Data Management
V21 | Unencrypted data at rest in all possible locations (Device, Cloud) | Yes | n/a
V22 | Unencrypted data in transit in all flowchannels | Yes | n/a
V23 | Weak Encryption Implementaion for data at rest and in motion tactical and design wise or Inadequate encryption strength | Yes | n/a
V24 | Use of a Broken or Risky Cryptographic Algorithm | Yes | n/a
V25 | Insecure Storage of Sensitive Information (Application data) | Yes | n/a
V26 | Absence of application (status/behavior) logs and audit logs | Yes | n/a
V27 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging | Yes | n/a
V28 | Information Exposure Through Log Files | Yes | n/a
V29 | Insecure Storage of Sensitive Information (Stryker data - keys, certs, etc.) | Yes | n/a
Application/Communication Management
V31 | Untrusted 3rd party applications installation from appstore | Yes | n/a
V32 | No security applications (antivirus, etc..) running in the ipad/iphone | Yes | n/a
V33 | Unrestricted Wifi communication with open/untrusted sources | Yes | n/a
V34 | Unrestricted BT communication/transfer with multiple sources | Yes | n/a
V35 | Removable device communication through charging interfaces | Yes | n/a
V36 | Unsecured/Unhardened Network Interfaces | Yes | n/a
InSecure Configurations of Resources
V41 | Lack of efficient configuration change control - App | Yes | n/a
V42 | Malicious Software Download | No | n/a
V43 | Malicious Automated Software Update | No | n/a
V44 | Control/modification of System(host) settings or Device Setting | Yes | n/a
V45 | Unsecured configuration of wireless devices | Yes | n/a
Cloud Backend
V51 | Unsecurely transmitted data to cloud | No | n/a
V52 | Unsecurely encrypting the database | No | n/a
V53 | Single Factor Authentication to cloud | No | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware | For ipad/iphone, the charging/USB interface is provided & this can be used to connect the OTG devices. Other Network Interfaces can also be used for injecting the malware. | TSA-1 - Individual (Employees, Insider, Trusted Insider/Priveleged Insider) TSA-3 - Skript Kiddies | Yes | n/a
T02 | Deliver directed malware | Threat source delivers malware on a removable media which was designed to exploit a known vulnerability of the existing Software. Directed attack on the Application using knowledge about the Application. | TSA-2 Organization | Yes | n/a
T03 | Perform wireless perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies | Yes | n/a
T04 | Unauthorized modifications to Knee Balancer Software | Adversary gains unauthorized access to the system and removes or modifies application binaries. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T05 | Unauthorized access to Knee Balancer App | Adversary gains unauthorized access to the system and removes or modifies application configuration and gains access | TSA-3 - Skript Kiddies | Yes | n/a
T06 | Unauthorized access/modification to Treatment data/plan | Adversary gains unauthorized access to the treatment/Knee Gap Analysis plan of the patient. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T07 | Unauthorized access/modification to PHI/Health data | Adversary gains unauthorized access to the system and steals ePHI information | TSA-3 - Skript Kiddies- | Yes | n/a
T08 | Unauthorized access to app sensitive data (device's keys/certificates/crypto algorithms) | Adversary gains unauthorized access to the Crypto keys and associated tokens,certificates | TSA-3 - Skript Kiddies | Yes | n/a
T09 | Brute force login attempts | Adversary gains unauthorized access to the system using different kinds of repeated attacks | TSA-3 - Skript Kiddies | Yes | n/a
T10 | Unprotected access to device settings (uninstalling applications, etc..) | Adversary gains unauthorized access to the system and make modifications to the software apps & configuration files | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T11 | Unrestricted wireless communication with multiple sources | Several wireless interfaces such as WiFi, BT exists in the mobile device and they can be used to connect with untrusted network(s)/source(s) | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T12 | Sensitive information exposure through audit log | Sensitive information such as patient name, PHI data should not be revealed through audit log. Audit (log) data is considered as a sensitive data. It should be secured with proper access control | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T13 | Improper access/modification of audit log | Audit (log) data is considered as a sensitive data. It should be secured with proper access control | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T14 | Data input from untrusted source | Application receiving input health data from another system should verify the source | TSA-1 - Individual (Employees, Insider, Trusted Insider/Priveleged Insider) | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V33 | Unrestricted Wifi communication with open/untrusted sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | None | Low | High | Network | Low | Low | Required | Unchanged | 2.0681 | 0.6568 | 4.2167 | 6.3 | Moderate | 0.5 | 5.3 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
2 | T01 | Deliver undirected malware | V34 | Unrestricted BT communication/transfer with multiple sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | None | Low | High | Network | Low | Low | Required | Unchanged | 2.0681 | 0.6568 | 4.2167 | 6.3 | Moderate | 0.5 | 5.3 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T01 | Deliver undirected malware | V35 | Removable device communication through charging interfaces | A14 | Physical charging ports/interfaces | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | None | Low | High | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.6568 | 4.2167 | 4.8 | Very Low | 0.04 | 4.3 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
4 | T01 | Deliver undirected malware | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T01 | Deliver undirected malware | V18 | 3rd party installed applications from app store | A01 | Mobile Device (ipad, iphone V8-13) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T02 | Deliver directed malware | V33 | Unrestricted Wifi communication with open/untrusted sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | By knowing the end-to-end functionality & Components used in the application, the adversary can exploit the application and try to compromise the system with specific malware attacks. |  | None | Low | High | Network | Low | Low | Required | Unchanged | 2.0681 | 0.6568 | 4.2167 | 6.3 | Moderate | 0.5 | 5.3 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T02 | Deliver directed malware | V34 | Unrestricted BT communication/transfer with multiple sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | By knowing the end-to-end functionality & Components used in the application, the adversary can exploit the application and try to compromise the system with specific malware attacks. |  | None | Low | High | Network | Low | Low | Required | Unchanged | 2.0681 | 0.6568 | 4.2167 | 6.3 | Moderate | 0.5 | 5.3 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T02 | Deliver directed malware | V35 | Removable device communication through charging interfaces | A14 | Physical charging ports/interfaces | By knowing the end-to-end functionality & Components used in the application, the adversary can exploit the application and try to compromise the system with specific malware attacks. |  | None | Low | High | Physical | Low | Low | Required | Unchanged | 0.4866 | 0.6568 | 4.2167 | 4.8 | Very Low | 0.04 | 4.3 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T03 | Perform wireless perimeter network reconnaissance/scanning. | V36 | Unsecured/Unhardened Network Interfaces | A01 | Mobile Device (ipad, iphone V8-13) | Open/Unrestricted Network interfaces are easy to exploit. Secure and hardened network elements protect the system from attacks. Code validation through static analysis required |  | Low | None | High | Network | High | None | None | Unchanged | 2.2212 | 0.6568 | 4.2167 | 6.5 | Moderate | 0.5 | 5.4 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T03 | Perform wireless perimeter network reconnaissance/scanning. | V17 | No Static Analysis for developed code & SCA for SOUP in Application | A03 | Knee balancer application | Open/Unrestricted Network interfaces are easy to exploit. Secure and hardened network elements protect the system from attacks. Code validation through static analysis required |  | Low | None | High | Network | High | None | None | Unchanged | 2.2212 | 0.6568 | 4.2167 | 6.5 | Moderate | 0.5 | 5.4 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T04 | Unauthorized modifications to Knee Balancer Software | V12 | Unsafely compiled Application Binaries in Application | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T04 | Unauthorized modifications to Knee Balancer Software | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T04 | Unauthorized modifications to Knee Balancer Software | V16 | Corrupting the binaries/injecting malicious source code in Application | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
14 | T04 | Unauthorized modifications to Knee Balancer Software | V17 | No Static Analysis for developed code & SCA for SOUP in Application | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T05 | Unauthorized access to Knee Balancer App | V01 | Not in scope | A02 | Login (4-digit pin) management | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. PIN Establishment & Mapping to Proper MPS User 2. Multi Factor Authentication 3. Credential/PIN Secure local storage 4. Remote PIN Reset facility after timeline | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T05 | Unauthorized access to Knee Balancer App | V02 | Not in scope | A02 | Login (4-digit pin) management | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. PIN Establishment & Mapping to Proper MPS User 2. Multi Factor Authentication 3. Credential/PIN Secure local storage 4. Remote PIN Reset facility after timeline | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T05 | Unauthorized access to Knee Balancer App | V04 | Improper Storage of Stryker Credentials | A02 | Login (4-digit pin) management | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. PIN Establishment & Mapping to Proper MPS User 2. Multi Factor Authentication 3. Credential/PIN Secure local storage 4. Remote PIN Reset facility after timeline | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T05 | Unauthorized access to Knee Balancer App | V01 | Not in scope | A03 | Knee balancer application | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. PIN Establishment & Mapping to Proper MPS User 2. Multi Factor Authentication 3. Credential/PIN Secure local storage 4. Remote PIN Reset facility after timeline | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T05 | Unauthorized access to Knee Balancer App | V04 | Improper Storage of Stryker Credentials | A03 | Knee balancer application | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. PIN Establishment & Mapping to Proper MPS User 2. Multi Factor Authentication 3. Credential/PIN Secure local storage 4. Remote PIN Reset facility after timeline | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T06 | Unauthorized access/modification to Treatment data/plan | V21 | Unencrypted data at rest in all possible locations (Device, Cloud) | A03 | Knee balancer application | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | Proper Access control should be implimented in application directory so that only authorised users can access Treatment plan and its files. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
21 | T06 | Unauthorized access/modification to Treatment data/plan | V02 | Not in scope | A03 | Knee balancer application | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | Proper Access control should be implimented in application directory so that only authorised users can access Treatment plan and its files. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
22 | T06 | Unauthorized access/modification to Treatment data/plan | V03 | Improper Input Validation (allowed range) | A01 | Mobile Device (ipad, iphone V8-13) | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. 1. If allowed input range/format not validated then the application functionality gets affected. 2. With insuffient encryption & security this data gets affected. |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | Proper Access control should be implimented in application directory so that only authorised users can access Treatment plan and its files. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
23 | T06 | Unauthorized access/modification to Treatment data/plan | V25 | Insecure Storage of Sensitive Information (Application data) | A01 | Mobile Device (ipad, iphone V8-13) | Pre-op/Treatment plan data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | Proper Access control should be implimented in application directory so that only authorised users can access Treatment plan and its files. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
24 | T07 | Unauthorized access/modification to PHI/Health data | V21 | Unencrypted data at rest in all possible locations (Device, Cloud) | A04 | Patient Identifiable data | PHI/Health data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. The Needed PHI data should be identified, then stored with encryption. 2. The PHI data directory should only be accessible to MPS user using the application. 3. The PHI data should not be stored in local SQLIte databases in plain text. 4. The PHI data should not be stored in Log files. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
25 | T07 | Unauthorized access/modification to PHI/Health data | V25 | Insecure Storage of Sensitive Information (Application data) | A04 | Patient Identifiable data | PHI/Health data has to be secured properly & enforce the encryption to protect from unauthorized access/modifications. With insuffient encryption & security this data gets affected |  | High | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. The Needed PHI data should be identified, then stored with encryption. 2. The PHI data directory should only be accessible to MPS user using the application. 3. The PHI data should not be stored in local SQLIte databases in plain text. 4. The PHI data should not be stored in Log files. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
26 | T08 | Unauthorized access to app sensitive data (device's keys/certificates/crypto algorithms) | V21 | Unencrypted data at rest in all possible locations (Device, Cloud) | A04 | Patient Identifiable data | If any of the cryptographic security sensitive data such as keys, certificates, crypto algorithms, etc.., got modified then complete security gets compromised |  | None | High | Low | Local | Low | High | None | Unchanged | 0.79893 | 0.6568 | 4.2167 | 5.1 | Low | 0.2 | 4.4 | MEDIUM | 1. The keys, certificates and token should be stored in encrypted form. 2. Few of the sensitive data like the above should be stored using crypto. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
27 | T08 | Unauthorized access to app sensitive data (device's keys/certificates/crypto algorithms) | V25 | Insecure Storage of Sensitive Information (Application data) | A04 | Patient Identifiable data | If any of the cryptographic security sensitive data such as keys, certificates, crypto algorithms, etc.., got modified then complete security gets compromised |  | None | High | Low | Local | Low | High | None | Unchanged | 0.79893 | 0.6568 | 4.2167 | 5.1 | Low | 0.2 | 4.4 | MEDIUM | The keys, password and token should be stored in encrypted form.Few of the sensitive data like keys should be stored using windows credential manager. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
28 | T09 | Brute force login attempts | V02 | Not in scope | A12 | Security keys, tokens, certificates | Single factor authentication is not sufficient to surpress the brute force attacks. Minimum 2-factor authentication is needed to provide enough security from automated attacks |  | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Low | 0.2 | 4.6 | MEDIUM | 1. PIN Establishment & Mapping to Proper MPS User 2. Multi Factor Authentication 3. Credential/PIN Secure local storage 4. Remote PIN Reset facility after timeline/no.of unsuccessful attempts | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
29 | T09 | Brute force login attempts | V05 | Not in scope | A12 | Security keys, tokens, certificates | Single factor authentication is not sufficient to surpress the brute force attacks. Minimum 2-factor authentication is needed to provide enough security from automated attacks |  | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Low | 0.2 | 4.6 | MEDIUM | 1. PIN Establishment & Mapping to Proper MPS User 2. Multi Factor Authentication 3. Credential/PIN Secure local storage 4. Remote PIN Reset facility after timeline/no.of unsuccessful attempts | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
30 | T10 | Unprotected access to device settings (uninstalling applications, etc..) | V43 | Not in scope | A03 | Knee balancer application | Sensitive system settings on the host should be identified & protected from unauthorized access |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Device settings should be accessed with proper authentication 2. Modification should be enabled with proper authentication and limited range of options | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
31 | T11 | Unrestricted wireless communication with multiple sources | V43 | Not in scope | A01 | Mobile Device (ipad, iphone V8-13) | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Wireless communication with other sources should be ensured with any certificate/key/secured mechanism verification 2. Wireless interfaces (Wifi, BT) should be hardened with essential functionality enabled | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
32 | T11 | Unrestricted wireless communication with multiple sources | V43 | Not in scope | A03 | Knee balancer application | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Wireless communication with other sources should be ensured with any certificate/key/secured mechanism verification 2. Wireless interfaces (Wifi, BT) should be hardened with essential functionality enabled | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
33 | T11 | Unrestricted wireless communication with multiple sources | V44 | Control/modification of System(host) settings or Device Setting | A01 | Mobile Device (ipad, iphone V8-13) | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Wireless communication with other sources should be ensured with any certificate/key/secured mechanism verification 2. Wireless interfaces (Wifi, BT) should be hardened with essential functionality enabled | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
34 | T11 | Unrestricted wireless communication with multiple sources | V44 | Control/modification of System(host) settings or Device Setting | A03 | Knee balancer application | For wifi access only the HDO/Hospital environment should allowed to connect with. BT should only be connected on purpose for file transfers |  | None | Low | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Wireless communication with other sources should be ensured with any certificate/key/secured mechanism verification 2. Wireless interfaces (Wifi, BT) should be hardened with essential functionality enabled | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
35 | T12 | Sensitive information exposure through audit log | V26 | Absence of application (status/behavior) logs and audit logs | A03 | Knee balancer application | Sensitive information should not be present in the audit log. Audit log data should be maintained to identify security breaches. With no proper authentication to access the audit logs the content can be modified/deleted. |  | High | None | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. All log files should be stored encrypted . 2. The log files should not contain senstive user information and license information. 3. The path should not be exposed to the user | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
36 | T13 | Improper access/modification of audit log | V27 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging | A03 | Knee balancer application | Audit log data should be maintained to identify security breaches. With no proper authentication to access the audit logs the content can be modified/deleted. |  | None | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. All log files should be stored encrypted . 2. The log files should not contain senstive user information and license information. 3. The path should not be exposed to the user | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V33 | Unrestricted Wifi communication with open/untrusted sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities.
2 | T01 | Deliver undirected malware | V34 | Unrestricted BT communication/transfer with multiple sources | A13 | Wireless communication interfaces (Wifi, BT,etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities.
3 | T01 | Deliver undirected malware | V35 | Removable device communication through charging interfaces | A14 | Physical charging ports/interfaces | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities.
4 | T01 | Deliver undirected malware | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A06 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities.
5 | T01 | Deliver undirected malware | V18 | 3rd party installed applications from app store | A01 | Mobile Device (ipad, iphone V8-13) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | MEDIUM | 1. Restriction of wireless connectivity 2. Hospital Firewall 3. Secure AccessPoints & device H/W interfaces 4. Antivirus and any OS recommended security tools 5. The 3rd party modules and dependencies should be updated with latest versions. 6. Static code analysis for finding the Vunerabilities.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls