## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Knee Balancer
Scope:
Date: |  | <2022-06-08>
Conducted by: |  | <SIVA KUMAR /Security / LTTS> <Author Name / Function / Organization>
ID # | Asset Type (Information/Physical) | Asset | Asset Description | Comments
A01 | Physical asset | Mobile Device (iPad, iPhone V8-13) | iPhone/iPad is used for running the knee balancer application (currently V8-13 are supported)
A02 | Information asset | User identity | User tagged with Azure AD credentials for validation
A03 | Information asset | Knee balancer application | Knee balancer application is composed of 3rd party applications (binaries, frameworks, 3rd party components, etc..)
A04 | Information asset | MAC/iOS operating system | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A05 | Information asset | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | Knee balancer application is composed of 3rd party applications (binaries, frameworks, 3rd party components, etc..)
A06 | Information asset | MAKO Screen | It provides input for the Knee Balancer application to generate solutions using algorithms
A07 | Information asset | Log Files | Log files supporting the application
A08 | Information asset | Security keys, tokens, certificates | Keys, tokens and certificates
A09 | Information asset | Data at Rest | Data used by the application
A10 | Information asset | Data in Transit | Sensitive data being transmitted
A11 | Information asset | Cloud interface & Storage | For storing blobs
A12 | Information asset | Stryker IP data | Algorithms used to generate the solutions
A13 | Information asset | Stryker App store | For authorized MPS users, Knee Balancer application can be downloaded from stryker app store
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user related data (creds) | No | Knee balancer application leveraging the mobile login credentials and for upload feature Microsoft authentication gets leveraged
V02 | Improper Authentication in the application (single factor, no captcha, etc..) | No | Knee balancer application leveraging the mobile login credentials and for upload feature Microsoft authentication gets leveraged
V03 | Improper Input Validation (allowed range) | Yes | n/a
V04 | Improper Storage of Stryker IP data | Yes | n/a
V05 | Device deviated from the normal functionality (jailbroken) | Yes | n/a
V06 | Irregular upload of case data to cloud storage | Yes | n/a
V07 | Open to download (All stryker users) without restriction | Yes | n/a
SBOM/SOUP
V11 | Unpatched operating system | No | Security Operations Manual
V12 | Unsafely compiled Application Binaries in Application | Yes | n/a
V13 | Ineffective patch management | No | Security Operations Manual
V14 | Lack of OS vulnerability management | No | Security Operations Manual
V15 | 3rd Party Component Dependency & Vulnerabilities in Application | Yes | n/a
V16 | Unrestricted access to the stored application resources | Yes | n/a
V17 | No Static Analysis for developed code & SCA for SOUP in Application | Yes | n/a
Application Data Management
V21 | Unencrypted sensitive data at rest in all possible locations (Device, Cloud) | Yes | n/a
V22 | Unencrypted sensitive data in transit in all flow channels | Yes | n/a
V23 | Weak Encryption Implementation for data at rest and in motion tactical and design wise or Inadequate encryption strength | Yes | n/a
V24 | Use of a Broken or Risky Cryptographic Algorithm | Yes | n/a
V25 | Information Exposure Through Log Files | Yes | n/a
V26 | Absence of application (status/behavior) logs and audit logs | Yes | n/a
V27 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging | Yes | n/a
V28 | Insecure Storage of Sensitive Information (Application data) | Yes | n/a
V29 | Insecure Storage of Sensitive Information (Stryker data - keys, certs, etc.) | Yes | n/a
Application/Communication Management
V31 | Untrusted 3rd party applications installation from appstore | No | Security Operations Manual
V32 | No security applications (antivirus, etc..) running in the ipad/iphone | No | Security Operations Manual
V33 | Unrestricted Wifi communication with open/untrusted sources (app update) | No | Security Operations Manual
V34 | Unrestricted BT communication/transfer with multiple sources (app update) | No | Security Operations Manual
InSecure Configurations of Resources
V41 | Malicious Software Download | No | Security Operations Manual
V42 | Malicious Automated Software Update | No | Security Operations Manual
Blob Storage
V51 | Generating SAS with indefinite time period | Yes | n/a
V52 | Improper key management during SAS generation | Yes | n/a
V53 | Compromised SAS URL's for blob storage | Yes | n/a
V54 | Exhaustion of the confined storage limit | Yes | n/a
V55 | Improper authorization (RBAC) during blob storage access | Yes | n/a
V56 | Usage of storage account key to access the blob storage | Yes | n/a
V57 | Sharing the SAS URL in unsecured channel | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware | For iPad/iPhone, the charging/USB interface is provided & this can be used to connect the OTG devices. Other Network Interfaces can also be used for injecting the malware. | TSA-1 - Individual (Employees, Insider, Trusted Insider/Priveleged Insider) TSA-3 - Skript Kiddies | No | The trust boundary of this risk analysis is only permitted to application and hence not considering the mobile device
T02 | Deliver directed malware | Threat source delivers malware on a removable media which was designed to exploit a known vulnerability of the existing Software. Directed attack on the Application using knowledge about the Application. | TSA-2 Organization | No | The trust boundary of this risk analysis is only permitted to application and hence not considering the mobile device
T03 | Perform wireless perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies | No | The trust boundary of this risk analysis is only permitted to application and hence not considering the mobile device. Application has no web interface functionality
T04 | Unauthorized modifications to Knee Balancer Software | Adversary gains unauthorized access to the system and removes or modifies application binaries. | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T05 | Unauthorized access to Knee Balancer App | Adversary gains unauthorized access to the system and removes or modifies application configuration and gains access | TSA-3 - Skript Kiddies | Yes | n/a
T06 | Unauthorized access to app sensitive data (device's keys/certificates/crypto algorithms) | Adversary gains unauthorized access to the Crypto keys and associated tokens, certificates | TSA-3 - Skript Kiddies | Yes | n/a
T07 | Sensitive information exposure through audit log | Sensitive information such as patient name, PHI data should not be revealed through audit log. Audit (log) data is considered as a sensitive data. It should be secured with proper access control | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T08 | Improper access/modification of audit log | Audit (log) data is considered as a sensitive data. It should be secured with proper access control | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T09 | Corrupting the binaries/injecting malicious source code in Application | Adversary identifies the binaries/source files supporting the application and tries to tamper for making the app unavailable | TSA-3 - Skript Kiddies | Yes | n/a
T10 | Improper access/modification of Stryker IP data | Adversary gains unauthorized access to the system and can remove/modify Stryker IP data | TSA-2 Organization TSA-3 - Skript Kiddies | Yes | n/a
T11 | Improper validation/assessment of 3rd party components (libs, fw, etc..) | All 3rd party components (libs, fw, etc…) should be validated properly related to the vulnerabilities and also their history of updates | TSA-2 Organization | Yes | n/a
T12 | Application running on the abnormal device (jailbroken) | Adversary might gain access to the jailbroken device easily and can harms the application | TSA-2 Organization | Yes | n/a
T13 | Missing data with factory reset | As upload to the blob waits for the device to be online and with factory reset before upload leads to data unavailability | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
T14 | Exposed SAS (token/URL/Key) to unauthorized parties | SAS (token/URL/key) needs to be confidential and should not be revealed to 3rd parties | TSA-3 - Skript Kiddies | Yes | n/a
T15 | Unavailability of uploading to Blob storage | New data/files can't be stored or moved to the blob storage | TSA-3 - Skript Kiddies | Yes | n/a
T16 | Reveal/exposure of sensitive information (Stryker IP data) | Exposure of sensitive information (stryker IP data) to unauthorized users | TSA-1 - Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Not in scope | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | None | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.22 | 1.4124 | 3.3 | Low | 0.2 | 1.8 | LOW | Not in scope | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
2 | T01 | Not in scope | V31 | Not in scope | A01 | Mobile Device (iPad, iPhone V8-13) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | None | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.22 | 1.4124 | 3.3 | Low | 0.2 | 1.8 | LOW | Not in scope | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
3 | T03 | Not in scope | V17 | No Static Analysis for developed code & SCA for SOUP in Application | A03 | Knee balancer application | Open/Unrestricted Network interfaces are easy to exploit. Secure and hardened network elements protect the system from attacks. Code validation through static analysis required |  | Low | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.3916 | 2.5141 | 4.8 | Moderate | 0.5 | 3.7 | LOW | Not in scope | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
4 | T04 | Unauthorized modifications to Knee Balancer Software | V12 | Unsafely compiled Application Binaries in Application | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. In the build configuration, enable binary protection with safe compilation by ensuring stack protection, PIE support, Automatic Reference Counting (ARC) and setting byte-code minification. 2. Reference to the binary protection followed in the compilation process with design document. | 1 &2. Xcode Configurtions are updated in Design Doument. Refer Section -7 Document ID :SGTC-QFM-DLC-001-06_SOFTWARE_DETAILED_DESIGN | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
5 | T04 | Unauthorized modifications to Knee Balancer Software | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. The 3rd party components (SOUP items) and dependencies should be identified and listed in the design document. 2. SCA tools (such as blackduck) should be used for evaluating the risk from 3rd party components with security testing. | 1. Soup Item are updated in Design Document And Soup Document. Refer SGTC-QFM-DLC-001-06_SOFTWARE_DETAILED_DESIGN & SGTC-QFM-DLC-001-30_SOFTWARE_SOUP_DESCRIPTION Documents. 2. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
6 | T04 | Unauthorized modifications to Knee Balancer Software | V16 | Unrestricted access to the stored application resources | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | 1. Intentional/Unintentional copying/deletion of the resources should be validated. 2. Application resources should not be visible (or) accessed directly through any channel. 3. Proper security recommendations should be followed by the MPS users. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
7 | T04 | Unauthorized modifications to Knee Balancer Software | V17 | No Static Analysis for developed code & SCA for SOUP in Application | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Moderate | 0.5 | 5.2 | MEDIUM | Performing static code analysis (such as sonarqube) for identification of deviation from the coding standards and vulnerabilities on developed code. | Generated Analysis Report from xcode and followed ios swift languauge standards. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
8 | T05 | Unauthorized access to Knee Balancer App | V02 | Not in scope | A02 | User identity | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | Low | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.3916 | 2.5141 | 4.4 | Moderate | 0.5 | 3.5 | LOW | Not in scope | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
9 | T05 | Unauthorized access to Knee Balancer App | V01 | Not in scope | A03 | Knee balancer application | If the application access permissions are not properly enforced during the starting of the application, then its become difficult for the application to get protected from the adversary attacks |  | Low | None | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.3916 | 2.5141 | 4.4 | Moderate | 0.5 | 3.5 | LOW | Not in scope | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
10 | T07 | Sensitive information exposure through audit log | V26 | Absence of application (status/behavior) logs and audit logs | A03 | Knee balancer application | Sensitive information should not be present in the audit log. Audit log data should be maintained to identify security breaches. With no proper authentication to access the audit logs the content can be modified/deleted. |  | High | None | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Log files should not store any kind of debugging or detailed information. 2. The log files should not contain sensitive information (such as keys/certs,etc..). 3. The path should not be exposed directly/indirectly. | 1.Log files we are not storing any Detailed information, 2. In Log Files we are not maintaining sensitive information. 3. we are not exposing the Path(Directly/Indirectly) for log files | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
11 | T08 | Improper access/modification of audit log | V27 | Log Files: Audit Log Manipulation, Log Injection-Tampering-Forging | A03 | Knee balancer application | Audit log data should be maintained to identify security breaches. With no proper authentication to access the audit logs the content can be modified/deleted. |  | None | High | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Critical error/failure logs (unavailable, not working) in the application source code should be logged into audit file. 2. The path should not be exposed directly/indirectly. 3. Intentional/accidental, modification/deletion of audit log should not be allowed. | 1.Maintaining Audit file.2.Path Not Exposing.3.No one can able to delete. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
12 | T09 | Corrupting the binaries/injecting malicious source code in Application | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A03 | Knee balancer application | If the software components are not protected properly or exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. The 3rd party components (SOUP items) and dependencies should be identified and listed in the design document. 2. SCA tools (such as blackduck) should be used for evaluating the risk from 3rd party components with security testing. | 1. Soup Item are updated in Design Document And Soup Document. Refer SGTC-QFM-DLC-001-06_SOFTWARE_DETAILED_DESIGN & SGTC-QFM-DLC-001-30_SOFTWARE_SOUP_DESCRIPTION Documents. 2. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
13 | T10 | Improper access/modification of Stryker IP data | V04 | Improper Storage of Stryker IP data | A12 | Stryker IP data | Adversary gains unauthorized access to the system and can remove/modify Stryker IP data |  | None | High | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.6568 | 4.2167 | 6.1 | Low | 0.2 | 4.6 | MEDIUM | 1. Path should not be exposed directly/indirectly. 2. Copying of the binaries should not be allowed. 3. If we could do reverse engineering, no information should revealed by the binary. | 1.Path Not Exposing. 2. Followed 3.Followed | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
14 | T11 | Improper validation/assessment of 3rd party components (libs, fw, etc..) | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A03 | Knee balancer application | All 3rd party components (libs, fw, etc…) should be validated properly related to the vulnerabilities and also their history of updates |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. The 3rd party components (SOUP items) and dependencies should be identified and listed in the design document. 2. SCA tools (such as blackduck) should be used for evaluating the risk from 3rd party components with security testing. | 1. Soup Item are updated in Design Document And Soup Document. Refer SGTC-QFM-DLC-001-06_SOFTWARE_DETAILED_DESIGN & SGTC-QFM-DLC-001-30_SOFTWARE_SOUP_DESCRIPTION Documents. 2. NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
15 | T12 | Application running on the abnormal device (jailbroken) | V05 | Device deviated from the normal functionality (jailbroken) | A03 | Knee balancer application | App may gets effected if its running on a jailbroken device. Also, adversary can get easy access to the jailbroken devices. |  | None | None | High | Local | Low | None | None | Unchanged | 2.5151 | 0.56 | 3.5952 | 6.2 | Low | 0.2 | 4.1 | MEDIUM | 1. During application starting, there should be a provision to check the device status (normal/abnormal). 2. Deviation from any normal condition should identified by the device and inform the user it is vulnerable. | 1&2 TBD | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
16 | T13 | Missing data with factory reset | V06 | Irregular upload of case data to cloud storage | A01 | Mobile Device (iPad, iPhone V8-13) | As upload to the blob waits for the device to be online and with factory reset before upload leads to data unavailability |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Once the case data is ready to upload, notify the user regarding the normal life time of this case data. 2. The special cases where the data can be lost needs to be mentioned in the SOM (such as factory reset, app uninstallation). 3. TBD(Incase of any upload failure consider extending the case data life time) | 1.Popup will display every 12 hours if pastcase data is available. 2. For App uninstallation the data will lost which is not uploaded to cloud. 3.TBD | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
17 | T14 | Exposed SAS (token/URL/Key) to unauthorized parties | V53 | Compromised SAS URL's for blob storage | A11 | Cloud interface & Storage | SAS (token/URL/key) needs to be confidential and should not be revealed to 3rd parties |  | High | None | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Generate the SAS with limited time period. 2. Generate the SAS for every upload and never store the SAS in source code 3. Based on the application id SAS should be generated. 4. Keys should never be used for uploading the blob. Only generated SAS token from authorized source needs to be used. | 1.SAS time period is 2.Generating the SAS for every upload from source code. 3.Based on Resource group Name and stroage Account Name and Subscription Id. 4. Keys are not using for upload, it is used for generating SAS to upload. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
18 | T14 | Exposed SAS (token/URL/Key) to unauthorized parties | V57 | Sharing the SAS URL in unsecured channel | A11 | Cloud interface & Storage | SAS (token/URL/key) needs to be confidential and should not be revealed to 3rd parties |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Generate the SAS with limited time period. 2. Generate the new SAS for every upload and never store the SAS in source code 3. Based on the application id SAS should be generated. | 1.SAS time period is 7 days from start time 2.Generating the SAS for every upload from source code. 3.Based on Resource group Name and stroage Account Name and Subscription Id. 4. Keys are not using for upload, it is used for generating SAS to upload. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
19 | T15 | Unavailability of uploading to Blob storage | V54 | Exhaustion of the confined storage limit | A11 | Cloud interface & Storage | New data/files can't be stored or moved to the blob storage |  | None | None | High | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Notify the admin related to the upload failure with azure resources. 2. Possible solutions from the admin needs to be addressed (such as increasing/deleting the storage limit) | 1.TBD 2. we are deleting the data for every 90 days automatically. | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
20 | T16 | Reveal/exposure of sensitive information (Stryker IP data) | V07 | Open to download (All stryker users) without restriction | A13 | Stryker App store | Exposure of Stryker IP data to unauthorized users |  | High | None | None | Local | Low | Low | None | Unchanged | 1.8346 | 0.56 | 3.5952 | 5.5 | Low | 0.2 | 4 | MEDIUM | 1. Knee balancer application should be available for download only to the authorized MPS users. 2. with anykind of reverse engineering, no information should revealed from the binary. | 1. Kneebalancer Application should be able to install for only authorized users from Comp Portal app 2.NA | [9 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Not in scope | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | As Knee Balancer is running on a mobile device, it is important to secure the wireless/Network interfaces on the host such as BT, Wifi & charging interface. Without restrictions, any device can be connected with the host & hence affect the application |  | LOW | Not in scope
2 | T01 | Not in scope | V31 | Not in scope | A01 | Mobile Device (iPad, iPhone V8-13) | As Knee Balancer is running on a mobile device, it is important to secure the wireless interfaces on the host such as BT, Wifi. Without restrictions any device can be connected with the host |  | LOW | Not in scope
3 | T03 | Not in scope | V17 | No Static Analysis for developed code & SCA for SOUP in Application | A03 | Knee balancer application | Open/Unrestricted Network interfaces are easy to exploit. Secure and hardened network elements protect the system from attacks. Code validation through static analysis required |  | LOW | Not in scope
4 | T04 | Unauthorized modifications to Knee Balancer Software | V12 | Unsafely compiled Application Binaries in Application | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | MEDIUM | 1. In the build configuration, enable binary protection with safe compilation by ensuring stack protection, PIE support, Automatic Reference Counting (ARC) and setting byte-code minification. 2. Reference to the binary protection followed in the compilation process with design document.
5 | T04 | Unauthorized modifications to Knee Balancer Software | V15 | 3rd Party Component Dependency & Vulnerabilities in Application | A05 | Knee balancer supporting software components (binaries, frameworks, 3rd party components, etc..) | If the software components are not protected properly/exposed for access inside the mobile host with unrestricted/unauthorized access, then the modification/tampering of the software can lead to the unavailability of the application |  | MEDIUM | 1. The 3rd party components (SOUP items) and dependencies should be identified and listed in the design document. 2. SCA tools (such as blackduck) should be used for evaluating the risk from 3rd party components with security testing.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Priveleged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls