## Header (Optional)
Product Security Risk Table
Ortho Guidance Precision Knee System
Ortho Guidance Precision Knee Software
Document Control
| Author: | Nitin Sharma
| Document No.: | D0000055902
| Revision No.: | AC.11
| Project: | Precision Knee 6.0 | DC-0000002518
| Project Lead: | Raman Bhardwaj
Change Records
| Rev. No. | Comment | Date | Author
| AA | Initial revision | See onePLM | Nitin Sharma
| AB | Updated, System & Assests, Vulberability and Threats Updated Safety impact and security risk level as per DR4 expectations | See onePLM | Nitin Sharma
| AC | Updated security risk control measures and summary | See onePLM | Nitin Sharma
Note on B17: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## Document References (Optional)
Product Security Risk Table
References
ID | Title | Rev.* | Doc. No.
DI_PLATFORM | DI - Flex - Platform (0000013017 - Flex Platform) [DI_PLATFORM] | Auto | D0000063236NAV
DIOV_OPK | DIOV - Ortho Guidance Precision Knee System - Ortho Guidance Precision Knee Software (DC-0000002518 Precision Knee 6.0) [DIOV_OPK] | Auto | D0000039078
SRS | SRS - Ortho Guidance Precision Knee System - Ortho Guidance Precision Knee Software (DC-0000002518 Precision Knee 6.0) [SRS] | Auto | D0000039056
RARC | Risk Assessment - Ortho Guidance Precision Knee System - Ortho Guidance Precision Knee Software (DC-0000002518 Precision Knee 6.0) [RARC] | Auto | D0000039062
SRT | Product Security Risk Table (0000013017 - Flex Platform) [SRT] | Auto | D0000043989
DS_LS | DS - FP8000 Camera - FP8000 Localization Service (0000012975 - FP8000 Camera) [DS_LS] | Auto | D0000065533NAV
RARC_PLATFORM | Risk Assessment - Flex - Platform (0000013017 - Flex Platform) [RARC_PLATFORM] | Auto | D0000069270NAV
DI_CAMERA | DI - FP8000 Localization System - (0000012975 - FP8000 Camera) [DI_CAMERA] | Auto | D0000062256NAV
SDD_LS | SDD - FP8000 Camera - FP8000 Localization Service (0000012975 - FP8000 Camera) [SDD_LS] | Auto | D0000065242NAV
SDD_TOOL | SDD - EL100038 Firmware, EL100043 Firmware, EL300407 Firmware (0000010131 - Nav Handpieces + Trackers for Neuro) [SDD_TOOL] | Auto | D0000062198NAV
SDD_API | SDD - FP8000 Camera - Application Programming Interface (0000012975 - FP8000 Camera) [SDD_API] | Auto | D0000065269NAV
DIOV_PLATFORM | DIOV - Flex Platform - Flex Platform (0000013017 - Flex Platform) [DIOV_PLATFORM] | Auto | D0000065005NAV
DIOV_EMBOX | DIOV - Q Guidance System - EM Tracking system - Flex EM Box (DC-0000002585 EM Box) [DIOV_EMBOX] | Auto | D0000065004NAV
DIOV_EMSTYLET | DIOV - EM Tracking - EM Single Use Sterile Instruments (DC-0000002586 EM Stylet) [DIOV_EMSTYLET] | Auto | D0000072106NAV
DI_SPIDER | DI - Spider - Spider (DC-0000002512 Spider) [DI_SPIDER] | Auto | D0000024194
DIOV_SPIDER | DIOV - Spider - Spider (DC-0000002512 Spider) [DIOV_SPIDER] | Auto | D0000059611
SDS_APPMAN | DS - Flex - Application Manager (0000013017 - Flex Platform) [SDS_APPMAN] | Auto | D0000023166
DIOV_FP8000_CAMERA | DIOV - FP8000 Camera - SubSystem (0000012975 - FP8000 Camera) [DIOV_FP8000_CAMERA] | Auto | D0000064363NAV
Note on A3: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## System & Asset Identification
System & Asset Identification
Medical Device / System: |  | Stryker Flex Cart Navigation System
Scope: |  | Ortho Guidance Precision Knee Software
Date: |  | Refer to OnePLM
Conducted by: |  | Nitin Sharma, Senior Staff Engineer, SGTC R&D Akshay Shukla, Senior Software Quality Engineer, SGTC QA/RA
ID # | Asset Type (Information/Physical) | Asset | Asset Description
A01 | Physical Asset | System resources | Contains software module which controls the overall functionality of Precision Knee System i.e. intra-operative planning and guidance.
A02 | Information Asset | Admin Password / Credentials / System Configuration / Certificates | Admin Passwords/Credentials allow modifying config files, accessing ePHI, software binaries, computer configuration files, log files, and computer resources as well as installing applications, security certificates, OS updates.
A03 | Information Asset | User Password / Credentials / System Configuration | User Passwords/Credentials provide access to ePHI, Software Binaries, read access to config files.
A05 | Information asset | Patient health information at rest | Datastored on hard disk after import. Contains patient identity, DOB, Age/Gender, image data
A07 | Information asset | Computer/OS network identification | Information about internals of the system (Device identification, software versions, supported protocols, etc.)
A08 | Information Asset | Intellectual Property | Algorithms and other core-elements of interest to competitors.
A09 | Physical Asset | Navigation Accuracy | Navigation Accuracy of FP8000 Camera, critical components for essential performance
A10 | Information Asset | SSH credentials for system admin access | Only using an 2048bit SSH key allows full access to all system functions. That SSH key is not provided to the customer and can be used by Stryker internals only.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Vulnerability Identification
Vulnerability Identification
Vuln. ID | Vulnerability Description | Applicable (Yes/No) | Rationale (if Vulnerability not applicable)
V01 | Ineffective management of user credentials | Yes | n/a
V02 | Ineffective management of admin credentials | Yes | n/a
V03 | Uneducated/ Malicious User | Yes | n/a
V04 | Insecure Boot / Boot for external media | Yes | n/a
SBOM
V11 | Unpatched COTS operating system | Yes | n/a
V12 | Potentially outdated SW-Library SOUPs | Yes | n/a
V13 | Defect or bug in Stryker software | Yes | n/a
Access points
V21 | Unprotected network port | Yes | n/a
V22 | Unprotected external USB Port | Yes | n/a
V24 | Unprotected hardware | Yes | n/a
Data
V31 | Unencrypted ePHI at rest | Yes | n/a
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Threat Assessment
Threat Assessment
# | Threat Event | Description | Threat Source | In Scope (Yes/No) | Rationale (if out of scope)
T01 | Deliver undirected malware | Threat source delivers malware by providing removable media prepared with malware. e.g. Removable media is left on a parking lot and picked up by hospital staff. USB stick finds its way to the Navigation System. Malware exploits a known vulnerability in order to gain access to restricted data or to manipulate data. Undirected attack on computer systems. | TSA-3 - Skript Kiddies, TSA-1 Individual | Yes | n/a
T02 | Deliver directed malware | Threat source delivers malicious software on a removable media which was designed to exploit a known vulnerability of the Navigation System. Directed attack on the Navigation System using knowledge of the Navigation System. | TSA-2 Organization | Yes | n/a
T03 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-3 - Skript Kiddies and TSA-1 Individual | Yes | n/a
T04 | Intercept network communication | Adversary manipulates the network environment to intercept network communication to gain ePHI, passwords, credit card information etc. or to manipulate DICOM data etc. (man-in-the middle attack) | TSA-2 Organization | No | For Imageless applications, data is never transferred over network.
T05 | Conduct scavenging of ePHI at rest | Adversary gains unauthorized access to the system and steals ePHI information | TSA-1 - Individual | Yes | n/a
T06 | Theft of system or hard drives | A thief steals the navigation system and/or the hard drives as loot, which can lead to unavailability of the navigation system and/or leak of Protected Health Information (PHI). | TSA-1 - Individual | Yes | n/a
T07 | Data theft via physical media | The malicious or inadvertent copy of ePHI on portable media (USB Sticks, DVD drives) from the navigation system. | TSA-2 Organization | Yes | n/a
T08 | Network-based denial of service (DoS) attack | A threat actor seeks to make a machine or network resource unavailable to its intended users by overloading the network resource, the network or target services with malicious data traffic. | TSA-3 - Skript Kiddies TSN-4 Malfunctioning Software / OS TSN-3 Malfunctioning device | Yes | n/a
T09 | Exploit physical access of authorized staff to gain access to organizational facilities. | Adversary follows (“tailgates”) authorized individuals into secure/controlled locations with the goal of gaining access to facilities, circumventing physical security checks. | TSA-1 Individual | No | HDO responsible for physical security controls
T10 | Mis-configuration by user | A user unknowingly or maliciously mis-configures the system so that it malfunctions or no longer functions at all. | TSN-1 Accidental | Yes | n/a
T11 | Remote exploit | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. This is likely inadvertent, targetting general vulnerability in popular operating systems. After successful exploit, continues as malware threat. | TSA-1 Individual | Yes | n/a
T12 | Improper disposal of hard disks | A user disposes the entire system or the hard disks (for instance after a service issue or at the end of the service life of the product). The hard disks still contain patient data. | TSN-1 Accidental | Yes | n/a
T13 | Man-in-the middle attack / intercept Navigation communication | Adversary manipulates the localization communication such as infrared - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T14 | Conduct scavenging of IP at rest | Adversary gains unauthorized access to the system and steals IP off the system | TSA-1 Individual | No | HDO responsible for physical security controls
T15 | Frequency Jamming or Destruction of Wireless Communication | Adversary manipulates the wireless network environment in order to render the service unavailable. | TSA-3 - Skript Kiddies | No | Ortho Imageless application will not be impacted due to unavailability of wireless network
T16 | Physical Manipulation of Hardware | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. | TSA-2 Organization or TSA-1 Individual | Yes | n/a
T17 | Gather information using open source discovery of organizational information | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. | TSA-3 - Skript Kiddies | Yes | n/a
T18 | Craft counterfeit certificates. | Adversary counterfeits or compromises a certificate authority, so that malware or connections will appear legitimate. | TSA-1 Individual | No | Device is a pass through of data in transit from end point devices. The certificate validation is the responsibility of the end points.
T19 | Install persistent and targeted sniffers on organizational information systems and networks. | Adversary places within internal organizational information systems or networks software designed to (over a continuous period of time) collect (sniff) network traffic. | TSA-2 Organization | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T20 | Insert malicious scanning devices (e.g., wireless sniffers) inside facilities. | Adversary uses postal service or other commercial delivery services to deliver to organizational mailrooms a device that is able to scan wireless communications accessible from within the mailrooms and then wirelessly transmit information back to adversary. | TSA-1 Individual | No | Device operates within the HDO trust boundary, overall network security is responsibility of the customer.
T21 | Power Failure at primary facility | Power Failure at primary facility makes facility including the navigation sytem inoperable | TSN-7 Power Failure | Yes | n/a
T22 | Theft of SSH key for root access of Navigation Systems | Adverary uses SSH root key to get unauthorized access to the full system. | TSA-1 Individual | No | SSH root key is stored safely at Stryker facility and never delivered to the HDO / customer.
T23 | Manipulation of tool firmware / memory | Adversary manipulates the tool related calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | No | For Imageless applications, no tool is calibrated.
T24 | Manipulation of navigation camera firmware / memory | Adversary manipulates the navgation camera related calibration data - thereby influencing the navigation accuracy | TSA-1 Individual | Yes | n/a
T25 | Manipulation of EM box firmware / memory | Adversary manipulates the EM box firmware - thereby influencing the navigation accuracy | TSA-1 Individual | No | EM is not used by Ortho Imageless applications
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Security Risk Assess
Security Risk Assessment | [6 empty] | 9814270399
[5 empty] | Adverse Impact |  |  |  | Pre-Implementation of Security Controls | [15 empty] | Security Controls/Mitigations |  |  |  |  | Post-Implementation of Security Controls
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Asset | Impact Description | Safety Impact (Risk ID# or N/A) | Confidentiality | Integrity | Availability | Attack Vector | Attack Complexity | Privileges Required | User Interaction | Scope | Exploitability Sub Score | ISC Base | Impact Sub Score | CVSS v3.0 Base Score | Threat Event Initiation | Threat Event Initiation Score | Overall Risk Score | Security Risk Level | Security Risk Control Measures | Implementation of Risk Control Measures | Verification of Risk Control Measures (Effectiveness) | Implementation of Risk Control Measures for Spider | Verification of Risk Control Measures (Effectiveness)3 | ConfidentialityP | IntegrityP | AvailabilityP | Attack VectorP | Attack ComplexityP | Privileges RequiredP | User InteractionP | ScopeP | Exploitability Sub ScoreP | ISC BaseP | Impact Sub ScoreP | CVSS v3.0 Base ScoreP | Overall Risk ScoreP | Security Risk LevelP | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V11 | Unpatched COTS operating system | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). | [RARC.RI040] | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | Moderate | 0.5 | 4.1 | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C5: It shall not be possible to manipulate system files or install malicious software in these sections | [DI_PLATFORM.SR512022] Local User cannot execute commands on system level [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512024] Operating System is protected against unauthorized manipulation | [DIOV_PLATFORM] | [DI_SPIDER.SR58020: Local command input on operating system level is not possible] [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58040: Operating System is protected against unauthorized manipulation] | [DIOV_SPIDER] | Low | None | Low | Local | High | High | Required | Unchanged | 0.333 | 0.3916 | 2.5141 | 2.9 | 2.7 | LOW | N/A
2 | T02 | Deliver directed malware | V24 | Unprotected hardware | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. | [RARC.RI040] | Low | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Moderate | 0.5 | 5.3 | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C3: BIOS password protection C4: BIOS settings are maintained after BIOS battery power loss. C5: It shall not be possible to manipulate system files or install malicious software in these sections | [DI_PLATFORM.SR512022] Local User cannot execute commands on system level [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512026] Bootup is protected against manipulation] [DI_PLATFORM.SR512024] Operating System is protected against unautherized manipulation | [DIOV_PLATFORM] | [DI_SPIDER.SR58020: Local command input on operating system level is not possible] [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58060: Bootup is protected against manipulation] [DI_SPIDER.SR58040: Operating System is protected against unauthorized manipulation] | [DIOV_SPIDER] | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | 3.9 | LOW | Very low prevalence of the navigation system makes it an extremely unlikely target for any directed attacks. An attacker with access to the physically secure OR environment does have easier targets.
3 | T10 | Mis-configuration by user | V03 | Uneducated/ Malicious User | A02 | Admin Password / Credentials / System Configuration / Certificates | If a user misconfigures the system, this might lead to unavailability of the navigation system or make the system less performant for the user. | [RARC.RI280] | None | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.3916 | 2.5141 | 3.9 | Low | 0.2 | 2.8 | LOW | C6: Password protection for user log in C7: Standard user has read only rights on config files C12: Use of data Encryption while storing data C35: Separate login account for user, where system files can not be manipulated (write protected). | [DI_PLATFORM.SR512023] Principle of least privileges. No Root user and other OS users have defined rights [SRS.F1230: Encryption of patient data at rest] [SRS.F1200: System access in an authorized way only] | [DIOV_OPK] [DIOV_PLATFORM] | [DI_SPIDER.SR58030: Principle of least privileges] [SRS.F1230: Encryption of patient data at rest] [SRS.F1200: System access in an authorized way only] | [DIOV_OPK] [DIOV_SPIDER] | None | None | None | Local | Low | Low | Required | Unchanged | 1.3382 | 0 | 0 | 0 | 0 | NONE | N/A
4 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities | [RARC.RI280] | Low | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.3916 | 2.5141 | 4.8 | Moderate | 0.5 | 3.7 | LOW | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512025] Remote Access is secured [DI_PLATFORM.SR512040] Segregated FP8000 ethernet network. | [DIOV_PLATFORM] | [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58050: Remote access is secured] [DI_SPIDER.SR58090: Segregated FP8000 ethernet network.] | [DIOV_SPIDER] | Low | None | Low | Network | High | None | None | Unchanged | 2.2212 | 0.3916 | 2.5141 | 4.8 | 3.7 | LOW | N/A
5 | T05 | Conduct scavenging of ePHI at rest | V01 | Ineffective management of user credentials | A05 | Patient health information at rest | ePHI stored on Navigation System is exposed to third parties. | [RARC.RI190] | High | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Low | 0.2 | 4.6 | MEDIUM | C6: Password protection for user log in C7: Standard user has read only rights on config files C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C17: The system shall provides role based controlled access to health data and functions C12: Use of data Encryption while storing data C14: The system ensures that data in logfile is masked with all available patient information C15: Disposing temporary files | [DI_PLATFORM.SR512023] Principle of least privileges. No Root user and other OS users have defined rights [DI_PLATFORM.SR512025] Remote Access is secured [SRS.F1230: Encryption of patient data at rest] [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked | [DIOV_OPK] [DIOV_PLATFORM] | [DI_SPIDER.SR58030: Principle of least privileges] [DI_SPIDER.SR58050: Remote access is secured] [SRS.F1230: Encryption of patient data at rest] [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked | [DIOV_OPK] [DIOV_SPIDER] | Low | None | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.3916 | 2.5141 | 4.2 | 2.9 | LOW | N/A
6 | T02 | Deliver directed malware | v02 | Ineffective management of admin credentials | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs malware in order to gain admin credentials | [RARC.RI040] | Low | High | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Moderate | 0.5 | 5.3 | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C3: BIOS password protection C4: BIOS settings are maintained after BIOS battery power loss. C5: It shall not be possible to manipulate system files or install malicious software in these sections | [DI_PLATFORM.SR512022] Local User cannot execute commands on system level [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512026] Bootup is protected against manipulation] [DI_PLATFORM.SR512024] Operating System is protected against unautherized manipulation Root access is not available locally. (ssh only): Thus Admin credentials cannot be stolen by means of keylogger. Root access is protected by using 2048 bit key stored within Stryker facility which is not acceccible externally. | [DIOV_PLATFORM] | [DI_SPIDER.SR58020: Local command input on operating system level is not possible] [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58060: Bootup is protected against manipulation] [DI_SPIDER.SR58040: Operating System is protected against unauthorized manipulation] Root access is not available locally. (ssh only): Thus Admin credentials cannot be stolen by means of keylogger. Root access is protected by using 2048 bit key stored within Stryker facility which is not acceccible externally. | [DIOV_SPIDER] | None | Low | None | Local | High | Low | None | Unchanged | 1.0483 | 0.22 | 1.4124 | 2.5 | 2 | LOW | N/A
7 | T13 | Man-in-the middle attack / intercept Navigation communication | V24 | Unprotected hardware | A09 | Navigation Accuracy | Man-in-middle attack on localization communication such as infrared - Adversary manipulates navigation accuracy | [RARC.RI190] | Low | High | High | Network | High | High | None | Unchanged | 0.70555 | 0.84899 | 5.4505 | 6.2 | Very Low | 0.04 | 5.5 | MEDIUM | C18: Physical lock of Navigation System service backdoor C19: Navigation camera can be opened only with special tools. All externally accessible screws are Torx C20: Communication between navigation camera and tools via infrared is secured by CRC checks C21: Integrity of tool geometry and plausibility of localization data is checked C13: Localization information can be granted exclusivly C22: User must be trained to use landmark check regularly to check navigation accuracy | [DI_PLATFORM.SR512030] Physical Protection by locking access to internal components such as Hard drives [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. [DI_CAMERA.SR1540] Calibration data is secured by CRC method [DI_CAMERA.SR1530], [SDD_TOOL] Instrument data is secured by CRC method [DI_CAMERA.SR1530] Data transmission from FP8000 camera device to FP8000 Library is secured by CRC method [DS_LS.SSR700] Integrity of all tool geometry is checked by OLS when a new tool gets registered [DS_LS.SSR645], [DS_LS.SSR650] Software detects localization errors by comparing current measured relative fiducial positions of a tracker object with initial relative fiducial positions and disables localization. [SDD_LS] Plausibility check after calculation of centroid position. Calculations of coordinates is terminated if boundaries of plausibility check are not met [DS_LS.SSR150] Exclusive Configuration Mode: An LS cosumer can grant exclusive rights so that no other consumer can change OLS configuration that might infringe tracking [SRS.F2700: Register Patient Dialog] The software application prompts that surgeon has to verify the quality of the registration with a landmark check | [DIOV_OPK] [DIOV_PLATFORM] [DIOV_FP8000_CAMERA] | [DI_SPIDER.SR58080: The system shall provide the possibility to physically lock the access to the internal components of the platform, including PC and hard drives.] [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. [DI_CAMERA.SR1540] Calibration data is secured by CRC method [DI_CAMERA.SR1530], [SDD_TOOL] Instrument data is secured by CRC method [DI_CAMERA.SR1530] Data transmission from FP8000 camera device to FP8000 Library is secured by CRC method [DS_LS.SSR700] Integrity of all tool geometry is checked by OLS when a new tool gets registered [DS_LS.SSR645], [DS_LS.SSR650] Software detects localization errors by comparing current measured relative fiducial positions of a tracker object with initial relative fiducial positions and disables localization. [SDD_LS] Plausibility check after calculation of centroid position. Calculations of coordinates is terminated if boundaries of plausibility check are not met [DS_LS.SSR150] Exclusive Configuration Mode: An LS cosumer can grant exclusive rights so that no other consumer can change OLS configuration that might infringe tracking [SRS.F2700: Register Patient Dialog] The software application prompts that surgeon has to verify the quality of the registration with a landmark check | [DIOV_OPK] [DIOV_SPIDER] [DIOV_FP8000_CAMERA] | Low | Low | Low | Network | High | High | None | Unchanged | 0.70555 | 0.52545 | 3.3734 | 4.1 | 3.5 | LOW | Very low prevalence of the navigation system makes it an extremely unlikely target for any directed attacks. In the cases of EM and IR Communication the technology is based on close proximity (near field), thus making a man-in-the middle attack very unlikely and increases the attack complexity.
8 | T08 | Network-based denial of service (DoS) attack | V21 | Unprotected network port | A01 | System resources | The navigation system is attacked using a DoS attack. In the worst case the system is unavailable to the user. | [RARC.RI280] | None | None | High | Network | Low | High | None | Unchanged | 1.2347 | 0.56 | 3.5952 | 4.9 | Moderate | 0.5 | 4.3 | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in | [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512025] Remote Access is secured | [DIOV_PLATFORM] | [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58050: Remote access is secured] | [DIOV_SPIDER] | None | None | Low | Network | Low | High | None | Unchanged | 1.2347 | 0.22 | 1.4124 | 2.7 | 2.1 | LOW | N/A
9 | T11 | Remote exploit | V21 | Unprotected network port | A09 | Navigation Accuracy | A threat actor uses the hospital network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system in order to manipulate navigation accuracy. | [RARC.RI190] | None | High | Low | Network | High | Low | None | Changed | 1.7769 | 0.6568 | 4.7173 | 7.1 | Low | 0.2 | 5.5 | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512025] Remote Access is secured [DI_PLATFORM.SR512040] Segregated FP8000 ethernet network. | [DIOV_PLATFORM] | [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58050: Remote access is secured] [DI_SPIDER.SR58090: Segregated FP8000 ethernet network.] | [DIOV_SPIDER] | None | Low | Low | Network | High | High | None | Unchanged | 0.70555 | 0.3916 | 2.5141 | 3.3 | 2.7 | LOW | N/A
10 | T11 | Remote exploit | V21 | Unprotected network port | A01 | System resources | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. In the worst case the system is unavailable to the user. | [RARC.RI280] | None | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.6568 | 4.2167 | 5.9 | Moderate | 0.5 | 5.1 | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in | [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512025] Remote Access is secured | [DIOV_PLATFORM] | [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58050: Remote access is secured] | [DIOV_SPIDER] | None | Low | Low | Network | High | None | None | Unchanged | 2.2212 | 0.3916 | 2.5141 | 4.8 | 3.7 | LOW | N/A
11 | T11 | Remote exploit | V11 | Unpatched COTS operating system | A02 | Admin Password / Credentials / System Configuration / Certificates | If a remote exploit (e.g. "worm") is able to circumvent the operating system defenses to gain local access, then several other attack scenarios are enabled. In the worst case the system is unavailable to the user. | [RARC.RI280] | High | High | Low | Network | Low | High | None | Unchanged | 1.2347 | 0.84899 | 5.4505 | 6.7 | Moderate | 0.5 | 6.1 | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C23: No automatic software update of the Navigation System | [DI_PLATFORM.SR512022] Local User cannot execute commands on system level [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512025] Remote Access is secured [DI_PLATFORM.SR512050] Prevention of automatic OS updates. [DI_PLATFORM.SR61041] Installation of Stryker Software (including OS updates) | [DIOV_PLATFORM] | [DI_SPIDER.SR58020: Local command input on operating system level is not possible] [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58050: Remote access is secured] [DI_SPIDER.SR58100: Prevent automatic update] [SDS_APPMAN.TR37090] Installation of Stryker Software (including OS updates) | [DIOV_SPIDER] | Low | Low | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.52545 | 3.3734 | 5 | 4.2 | MEDIUM | The Area of Attack, i.e. the Vulnarability of security related exploits in the OS is significantly reduced by the fact that a stripped down Linux Distribution is used. The likelyhood of exploits emerging during the product lifecycle is as opposed to a more widely spread COTS such as a Windows OS for example. Furthermore, the Security Operation Manual provided to the users contains information regarding updates and security patches.
12 | T12 | Improper disposal of hard disks | V24 | Unprotected hardware | A05 | Patient health information at rest | ePHI might be exposed to third parties if the storage medium of the navigation unit is improperly disposed of. | N/A | High | None | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | C18: Physical lock of Navigation System service backdoor C12: Use of data Encryption while storing data C15: Disposing temporary files C14: The system ensures that data in logfile is masked with all available patient information | [DI_PLATFORM.SR512030: The system shall provide the possibility to physically lock the access to the internal components of the flex platform, including PC and hard drives.] Physical Protection by locking access to internal components such as Hard drives [SRS.F1230: Encryption of patient data at rest] [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked | [DIOV_OPK] [DIOV_PLATFORM] | [DI_SPIDER.SR58080: The system shall provide the possibility to physically lock the access to the internal components of the platform, including PC and hard drives.] [SRS.F1230: Encryption of patient data at rest] [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked | [DIOV_OPK] [DIOV_SPIDER] | Low | None | None | Physical | Low | Low | None | Unchanged | 0.66712 | 0.22 | 1.4124 | 2.1 | 1.6 | LOW | N/A
13 | T06 | Theft of system or hard drives | V24 | Unprotected hardware | A01 | System resources | When parts of the Navigation system are stolen the navigation system might be rendered unavailable for use. | N/A | High | none | High | Physical | low | low | none | unchanged | 0.66712 | 0.8064 | 5.1771 | 5.9 | Low | 0.2 | 5.4 | MEDIUM | C18: Physical lock of Navigation System service backdoor. C19: Navigation camera can be opened only with special tools. | [DI_PLATFORM.SR512030] Physical Protection by locking access to internal components such as Hard drives [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. | [DIOV_PLATFORM] | [DI_SPIDER.SR58080: The system shall provide the possibility to physically lock the access to the internal components of the platform, including PC and hard drives.] [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. | [DIOV_SPIDER] | Low | None | Low | Physical | Low | High | none | unchanged | 0.29052 | 0.3916 | 2.5141 | 2.9 | 2.6 | LOW | N/A
14 | T07 | Data theft via physical media | V22 | Unprotected external USB Port | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by inserting physical media such as USB Sticks or External Drives. | N/A | High | none | none | Physical | low | low | none | unchanged | 0.66712 | 0.56 | 3.5952 | 4.3 | Low | 0.2 | 3.8 | LOW | C12: Use of data Encryption while storing data C15: Disposing temporary files C14: The system ensures that data in logfile is masked with all available patient information | [SRS.F1230: Encryption of patient data at rest] [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked | [DIOV_OPK] | [SRS.F1230: Encryption of patient data at rest] [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked | [DIOV_OPK] | Low | None | None | Physical | Low | Low | none | unchanged | 0.66712 | 0.22 | 1.4124 | 2.1 | 1.6 | LOW | N/A
15 | T16 | Physical Manipulation of Hardware | V24 | Unprotected hardware | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. Once defenses are circumvented in order to gain local access, several other Assets such as user credentials or PHI may be threatened. | [RARC.RI190] | Low | High | none | Physical | Low | High | none | Unchanged | 0.29052 | 0.6568 | 4.2167 | 4.6 | Low | 0.2 | 4.3 | MEDIUM | C18: Physical lock of Navigation System service backdoor. C19: Navigation camera can be opened only with special tools. | [DI_PLATFORM.SR512030] Physical Protection by locking access to internal components such as Hard drives [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. | [DIOV_PLATFORM] | [DI_SPIDER.SR58080: The system shall provide the possibility to physically lock the access to the internal components of the platform, including PC and hard drives.] [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. | [DIOV_SPIDER] | Low | Low | None | Physical | High | High | none | Unchanged | 0.16601 | 0.3916 | 2.5141 | 2.7 | 2.6 | LOW | N/A
16.A | T21 | Power Failure at primary facility | V24 | Unprotected hardware | A01 | System resources | Adversary attacks power supply of the hospital or OR member cuts of power supply of navigation system by acccident. In both cases navigation system goes off to missing power. | [RARC.RI280] | None | None | High | Local | Low | None | None | Unchanged | 2.5151 | 0.56 | 3.5952 | 6.2 | Very Low | 0.04 | 3.7 | LOW | C25: Powersafe Unit shall be part of the system such that in a power failure user can shift to alternate power or save the existing data and have a secured exit from the application | [DI_PLATFORM.SR46060] It is possible to unplug the system without power loss by using the integrated PowerBox. | [DIOV_PLATFORM] | N/A | NA | None | None | Low | Local | Low | None | None | Unchanged | 2.5151 | 0.22 | 1.4124 | 4 | 1.6 | LOW | N/A
16.B | T21 | Power Failure at primary facility | V24 | Unprotected hardware | A01 | System resources | Adversary attacks power supply of the hospital or OR member cuts of power supply of navigation system by acccident. In both cases navigation system goes off to missing power. | [RARC.RI280] | None | None | High | Local | Low | None | None | Unchanged | 2.5151 | 0.56 | 3.5952 | 6.2 | Very Low | 0.04 | 3.7 | LOW | None | N/A | N/A | None | None | None | None | Low | Local | Low | None | None | Unchanged | 2.5151 | 0.22 | 1.4124 | 4 | 1.6 | LOW | Complete power failure is highly unikely in a medical OR facility. Navigation system can not continue if there is no power at the OR facility.
17 | T17 | Gather information using open source discovery of organizational information | V21 | Unprotected network port | A01 | System resources | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. That information is used for an attack via network. | [RARC.RI040] | Low | Low | High | Network | High | Low | None | Unchanged | 1.6201 | 0.7323 | 4.7014 | 6.4 | Low | 0.2 | 5.1 | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512025] Remote Access is secured [DI_PLATFORM.SR512040] Segregated FP8000 ethernet network. | [DIOV_PLATFORM] | [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58040: Operating System is protected against unauthorized manipulation] [DI_SPIDER.SR58090: Segregated FP8000 ethernet network.] | [DIOV_SPIDER] | Low | Low | Low | Network | High | Low | None | Unchanged | 1.6201 | 0.52545 | 3.3734 | 5 | 3.7 | LOW | N/A
18 | T01 | Deliver undirected malware | V13 | Defect or bug in Stryker software | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) due to discovered defect or bug in Stryker software. | [RARC.RI040] | Low | Low | High | Local | Low | Low | Required | Unchanged | 1.3382 | 0.7323 | 4.7014 | 6.1 | Very Low | 0.04 | 4.8 | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled | [DI_PLATFORM.SR512022] Local User cannot execute commands on system level [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access | [DIOV_PLATFORM] | [DI_SPIDER.SR58020: Local command input on operating system level is not possible] [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] | [DIOV_SPIDER] | Low | Low | Low | Local | Low | Low | Required | Unchanged | 1.3382 | 0.52545 | 3.3734 | 4.8 | 3.5 | LOW | N/A
19 | T05 | Conduct scavenging of ePHI at rest | V04 | Insecure Boot / Boot for external media | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by booting the system with an manipulated / self owned operating system to access the harddrive data and bypass the user role system provided by the original Stryker operating system. | [RARC.RI190] | High | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.7323 | 4.7014 | 6.6 | Low | 0.2 | 5.1 | MEDIUM | C3: Boot menu is protected by password C5: It shall not be possible to manipulate system files or install malicious software in these sections | [DI_PLATFORM.SR512026] Bootup is protected against manipulation] [DI_PLATFORM.SR512024] Operating System is protected against unautherized manipulation | [DIOV_PLATFORM] | [DI_SPIDER.SR58060: Bootup is protected against manipulation] [DI_SPIDER.SR58040: Operating System is protected against unauthorized manipulation] | [DIOV_SPIDER] | Low | Low | Low | Local | Low | Low | None | Unchanged | 1.8346 | 0.52545 | 3.3734 | 5.3 | 3.8 | LOW | N/A
20 | T24 | Manipulation of navigation camera firmware / memory | V24 | Unprotected hardware | A09 | Navigation Accuracy | The threat actor manipulates the memory / firmware of navigation camera with the goal to bring the calibration off / manipulate the firmware to reduce navigation accuracy. | [RARC.RI190] | High | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.7323 | 4.7014 | 5.8 | Low | 0.2 | 5 | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access [DI_PLATFORM.SR512025] Remote Access is secured [DI_PLATFORM.SR512040] Segregated FP8000 ethernet network. | [DIOV_PLATFORM] | [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [DI_SPIDER.SR58050: Remote access is secured] [DI_SPIDER.SR58090: Segregated FP8000 ethernet network.] | [DIOV_SPIDER] | Low | Low | Low | Local | High | Low | None | Unchanged | 1.0483 | 0.52545 | 3.3734 | 4.5 | 3.6 | LOW | N/A
21 | T16 | Physical Manipulation of Hardware | V24 | Unprotected hardware | A09 | Navigation Accuracy | Adversary with direct mal-intent gains local access to the navigation flex cart platform and manipulates hardware such as tools and navigation camera. | [RARC.RI190] | Low | High | None | Physical | Low | High | None | Unchanged | 0.29052 | 0.6568 | 4.2167 | 4.6 | Low | 0.2 | 4.3 | MEDIUM | C18: Physical lock of Navigation System service backdoor C19: Navigation camera can be opened only with special tools C21: Integrity of tool geometry and plausibility of localization data is checked C27: Firmware and hardware version is stored on navigation camera. | [DI_PLATFORM.SR512030] Physical Protection by locking access to internal components such as Hard drives [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. [DS_LS.SSR700] Integrity of all tool geometry is checked by OLS when a new tool gets registered [DS_LS.SSR645], [DS_LS.SSR650] Software detects localization errors by comparing current measured relative fiducial positions of a tracker object with initial relative fiducial positions and disables localization. [SDD_LS] Plausibility check after calculation of centroid position. Calculations of coordinates is terminated if boundaries of plausibility check are not met [SDD_API] chapter 3.1 general properties | [DIOV_PLATFORM] | [DI_SPIDER.SR58080: The system shall provide the possibility to physically lock the access to the internal components of the platform, including PC and hard drives.] [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. [DS_LS.SSR700] Integrity of all tool geometry is checked by OLS when a new tool gets registered [DS_LS.SSR645], [DS_LS.SSR650] Software detects localization errors by comparing current measured relative fiducial positions of a tracker object with initial relative fiducial positions and disables localization. [SDD_LS] Plausibility check after calculation of centroid position. Calculations of coordinates is terminated if boundaries of plausibility check are not met [SDD_API] chapter 3.1 general properties | [DIOV_SPIDER] | Low | Low | None | Physical | High | High | none | Unchanged | 0.16601 | 0.3916 | 2.5141 | 2.7 | 2.6 | LOW | N/A

## Security Risk Control Measures
Security Risk Control Measures
Security Risk Control Measures
ID | Risk Control Measure | Category (optional) | Reference to more detailed specification | Reference to more detailed specification2
C1 | Only Stryker approved applications can be executed by the "Navigation User" | Platform | [DI_PLATFORM.SR512022] Local User cannot execute commands on system level | [DI_SPIDER.SR58020] Local User cannot execute commands on system level
C2 | A firewall shall be installed and activated. Unused ports and services disabled | Platform | [DI_PLATFORM.SR512021] I/O ports are secured against unauthorised access | [DI_SPIDER.SR58010] I/O ports are secured against unauthorised access
C3 | BIOS password protection | Platform | [DI_PLATFORM.SR512026] Bootup is protected against manipulation] | [DI_SPIDER.SR58060] Bootup is protected against manipulation]
C4 | BIOS settings are maintained after BIOS battery power loss. | Platform | n/a | n/a
C5 | It shall not be possible to manipulate system files or install malicious software in these sections | Platform | [DI_PLATFORM.SR512024] Operating System is protected against unauthorized manipulation | [DI_SPIDER.SR58040] Operating System is protected against unauthorized manipulation
C6 | Password protection for user log in | Platform | [DI_PLATFORM.SR512023] Principle of least privileges. No Root user and other OS users have defined rights | [DI_SPIDER.SR58030] Principle of least privileges. No Root user and other OS users have defined rights
C7 | Standard user has read only rights on config files | Platform | [DI_PLATFORM.SR512023] Principle of least privileges. No Root user and other OS users have defined rights | [DI_SPIDER.SR58030] Principle of least privileges. No Root user and other OS users have defined rights
C8 | Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in | Platform | [DI_PLATFORM.SR512025] Remote Access is secured | [DI_SPIDER.SR58050] Remote Access is secured
C9 | Protect FP8000 camera from being accessed from the hospital network | Platform | [DI_PLATFORM.SR512040] Segregated FP8000 ethernet network. | [DI_SPIDER.SR58090] Segregated FP8000 ethernet network.
C12 | Use of data Encryption while storing data | Application | [SRS.F1230: Encryption of patient data at rest] | [SRS.F1230: Encryption of patient data at rest]
C14 | The system ensures that data in logfile is masked with all available patient information and stored in an encrypted patient folder. | Application | [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked | [SRS.F2230: Log file] All patient information in logs like patient name anonymised or masked
C13 | Localization information can be granted exclusively | Localization Service | [DS_LS.SSR150] Exclusive Configuration Mode: An LS consumer can grant exclusive rights so that no other consumer can change OLS configuration that might infringe tracking | [DS_LS.SSR150] Exclusive Configuration Mode: An LS consumer can grant exclusive rights so that no other consumer can change OLS configuration that might infringe tracking
C18 | Physical lock of Navigation System service backdoor | Platform | [DI_PLATFORM.SR512030] Physical Protection by locking access to internal components such as Hard drives | [DI_SPIDER.SR58080] Physical Protection by locking access to internal components such as Hard drives
C19 | Navigation camera can be opened only with special tools | Camera | [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx. | [drawing KE308326] Device can be opened only with special tools. All externally accessible screws are Torx.
C20 | Communication between navigation camera and tools via infrared is secured by CRC checks | Camera | [DI_CAMERA.SR1540] Calibration data is secured by CRC method [DI_CAMERA.SR1530], [SDD_TOOL] Instrument data is secured by CRC method [DI_CAMERA.SR1530] Data transmission from FP8000 camera device to FP8000 Library is secured by CRC method | [DI_CAMERA.SR1540] Calibration data is secured by CRC method [DI_CAMERA.SR1530], [SDD_TOOL] Instrument data is secured by CRC method [DI_CAMERA.SR1530] Data transmission from FP8000 camera device to FP8000 Library is secured by CRC method
C21 | Integrity of tool geometry and plausibility of localization data is checked | LS | [DS_LS.SSR700] Integrity of all tool geometry is checked by OLS when a new tool gets registered [DS_LS.SSR645], [DS_LS.SSR650] Software detects localization errors by comparing current measured relative fiducial positions of a tracker object with initial relative fiducial positions and disables localization. [SDD_LS] Plausibility check after calculation of centroid position. Calculations of coordinates is terminated if boundaries of plausibility check are not met | [DS_LS.SSR700] Integrity of all tool geometry is checked by OLS when a new tool gets registered [DS_LS.SSR645], [DS_LS.SSR650] Software detects localization errors by comparing current measured relative fiducial positions of a tracker object with initial relative fiducial positions and disables localization. [SDD_LS] Plausibility check after calculation of centroid position. Calculations of coordinates is terminated if boundaries of plausibility check are not met
C22 | User must be trained to use landmark check regularly to check navigation accuracy | Application | [SRS.F2700: Register Patient Dialog] The software application prompts that surgeon has to verify the quality of the registration with a landmark check | [SRS.F2700: Register Patient Dialog] The software application prompts that surgeon has to verify the quality of the registration with a landmark check
C23 | No automatic software update of the Navigation System | Platform | [DI_PLATFORM.SR512050] Prevention of automatic OS updates. [DI_PLATFORM.SR61041] Installation of Stryker Software (including OS updates) | [DI_SPIDER.SR58100] Prevention of automatic OS updates. [SDS_APPMAN.TR37090: Installation of Applications]
C25 | Powersafe Unit shall be part of the system such that a temporary power failure does not interrupt the functioning of the navigation system for 6 minutes | Platform | [DI_PLATFORM.SR46060] It is possible to unplug the system without power loss by using the integrated PowerBox. | None
C27 | Firmware and hardware version is stored on navigation camera | Camera | [SDD_API] chapter 3.1 general properties | [SDD_API] chapter 3.1 general properties
C35 | Separate login account for user, where system files can not be manipulated (write protected). | Application | [SRS.F1200: System access in an authorized way only] | [SRS.F1200: System access in an authorized way only]

## Summary
Security Risk Assessment Summary
ID # | T ID | Threat Event(s) | V ID | Vulnerabilities | A ID | Assets | Impact Description | Safety Impact (Risk ID# or N/A) | Pre-Controls Risk Level | Security Risk Control Measures | Post-Controls Risk Level | Residual Security Risk Acceptability Justification
1 | T01 | Deliver undirected malware | V11 | Unpatched COTS operating system | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ). | [RARC.RI040] | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C5: It shall not be possible to manipulate system files or install malicious software in these sections | LOW | N/A
2 | T02 | Deliver directed malware | V24 | Unprotected hardware | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) by exploiting vulnarabilities such as bypassing the Operating System. | [RARC.RI040] | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C3: BIOS password protection C4: BIOS settings are maintained after BIOS battery power loss. C5: It shall not be possible to manipulate system files or install malicious software in these sections | LOW | Very low prevalence of the navigation system makes it an extremely unlikely target for any directed attacks. An attacker with access to the physically secure OR environment does have easier targets.
3 | T10 | Mis-configuration by user | V03 | Uneducated/ Malicious User | A02 | Admin Password / Credentials / System Configuration / Certificates | If a user misconfigures the system, this might lead to unavailability of the navigation system or make the system less performant for the user. | [RARC.RI280] | LOW | C6: Password protection for user log in C7: Standard user has read only rights on config files C12: Use of data Encryption while storing data C35: Separate login account for user, where system files can not be manipulated (write protected). | NONE | N/A
4 | T03 | Perform perimeter network reconnaissance/scanning. | V21 | Unprotected network port | A07 | Computer/OS network identification | Obtain knowledge about system internals in an attempt to find attack vectors and possibilities for exploitation of publicly known Vulnerabilities | [RARC.RI280] | LOW | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | LOW | N/A
5 | T05 | Conduct scavenging of ePHI at rest | V01 | Ineffective management of user credentials | A05 | Patient health information at rest | ePHI stored on Navigation System is exposed to third parties. | [RARC.RI190] | MEDIUM | C6: Password protection for user log in C7: Standard user has read only rights on config files C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C17: The system shall provides role based controlled access to health data and functions C12: Use of data Encryption while storing data C14: The system ensures that data in logfile is masked with all available patient information C15: Disposing temporary files | LOW | N/A
6 | T02 | Deliver directed malware | v02 | Ineffective management of admin credentials | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs malware in order to gain admin credentials | [RARC.RI040] | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C3: BIOS password protection C4: BIOS settings are maintained after BIOS battery power loss. C5: It shall not be possible to manipulate system files or install malicious software in these sections | LOW | N/A
7 | T13 | Man-in-the middle attack / intercept Navigation communication | V24 | Unprotected hardware | A09 | Navigation Accuracy | Man-in-middle attack on localization communication such as infrared - Adversary manipulates navigation accuracy | [RARC.RI190] | MEDIUM | C18: Physical lock of Navigation System service backdoor C19: Navigation camera can be opened only with special tools. All externally accessible screws are Torx C20: Communication between navigation camera and tools via infrared is secured by CRC checks C21: Integrity of tool geometry and plausibility of localization data is checked C13: Localization information can be granted exclusivly C22: User must be trained to use landmark check regularly to check navigation accuracy | LOW | Very low prevalence of the navigation system makes it an extremely unlikely target for any directed attacks. In the cases of EM and IR Communication the technology is based on close proximity (near field), thus making a man-in-the middle attack very unlikely and increases the attack complexity.
8 | T08 | Network-based denial of service (DoS) attack | V21 | Unprotected network port | A01 | System resources | The navigation system is attacked using a DoS attack. In the worst case the system is unavailable to the user. | [RARC.RI280] | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in | LOW | N/A
9 | T11 | Remote exploit | V21 | Unprotected network port | A09 | Navigation Accuracy | A threat actor uses the hospital network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system in order to manipulate navigation accuracy. | [RARC.RI190] | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | LOW | N/A
10 | T11 | Remote exploit | V21 | Unprotected network port | A01 | System resources | A threat actor uses the network connection of the device to target a vulnerability in the operating system and/or software to gain access to the navigation system. In the worst case the system is unavailable to the user. | [RARC.RI280] | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in | LOW | N/A
11 | T11 | Remote exploit | V11 | Unpatched COTS operating system | A02 | Admin Password / Credentials / System Configuration / Certificates | If a remote exploit (e.g. "worm") is able to circumvent the operating system defenses to gain local access, then several other attack scenarios are enabled. In the worst case the system is unavailable to the user. | [RARC.RI280] | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C23: No automatic software update of the Navigation System | MEDIUM | The Area of Attack, i.e. the Vulnarability of security related exploits in the OS is significantly reduced by the fact that a stripped down Linux Distribution is used. The likelyhood of exploits emerging during the product lifecycle is as opposed to a more widely spread COTS such as a Windows OS for example. Furthermore, the Security Operation Manual provided to the users contains information regarding updates and security patches.
12 | T12 | Improper disposal of hard disks | V24 | Unprotected hardware | A05 | Patient health information at rest | ePHI might be exposed to third parties if the storage medium of the navigation unit is improperly disposed of. | N/A | LOW | C18: Physical lock of Navigation System service backdoor C12: Use of data Encryption while storing data C15: Disposing temporary files C14: The system ensures that data in logfile is masked with all available patient information | LOW | N/A
13 | T06 | Theft of system or hard drives | V24 | Unprotected hardware | A01 | System resources | When parts of the Navigation system are stolen the navigation system might be rendered unavailable for use. | N/A | MEDIUM | C18: Physical lock of Navigation System service backdoor. C19: Navigation camera can be opened only with special tools. | LOW | N/A
14 | T07 | Data theft via physical media | V22 | Unprotected external USB Port | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by inserting physical media such as USB Sticks or External Drives. | N/A | LOW | C12: Use of data Encryption while storing data C15: Disposing temporary files C14: The system ensures that data in logfile is masked with all available patient information | LOW | N/A
15 | T16 | Physical Manipulation of Hardware | V24 | Unprotected hardware | A02 | Admin Password / Credentials / System Configuration / Certificates | Adversary with direct mal-intent gains local access to the navigation flex cart platform and installs hardware in order to spy and manipulate. Installed hardware may include HDMI, USB Splitters or even Keyloggers. Once defenses are circumvented in order to gain local access, several other Assets such as user credentials or PHI may be threatened. | [RARC.RI190] | MEDIUM | C18: Physical lock of Navigation System service backdoor. C19: Navigation camera can be opened only with special tools. | LOW | N/A
16.A | T21 | Power Failure at primary facility | V24 | Unprotected hardware | A01 | System resources | Adversary attacks power supply of the hospital or OR member cuts of power supply of navigation system by acccident. In both cases navigation system goes off to missing power. | [RARC.RI280] | LOW | C25: Powersafe Unit shall be part of the system such that in a power failure user can shift to alternate power or save the existing data and have a secured exit from the application | LOW | N/A
16.B | T21 | Power Failure at primary facility | V24 | Unprotected hardware | A01 | System resources | Adversary attacks power supply of the hospital or OR member cuts of power supply of navigation system by acccident. In both cases navigation system goes off to missing power. | [RARC.RI280] | LOW | None | LOW | Complete power failure is highly unikely in a medical OR facility. Navigation system can not continue if there is no power at the OR facility.
17 | T17 | Gather information using open source discovery of organizational information | V21 | Unprotected network port | A01 | System resources | Adversary mines publiclly accessible information togather information about organizational information systems, business processes, users or personnel, or external relationships that the adversary can subsequently employ in support of an attack. That information is used for an attack via network. | [RARC.RI040] | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | LOW | N/A
18 | T01 | Deliver undirected malware | V13 | Defect or bug in Stryker software | A01 | System resources | Malicious utilization of computer resources and computing power, incl. denial of service attacks, ransomware deployment, Bitcoin mining, etc., ) due to discovered defect or bug in Stryker software. | [RARC.RI040] | MEDIUM | C1: Only Stryker approved applications can be executed by the "Navigation User" C2: A firewall shall be installed and activated. Unused ports and services disabled | LOW | N/A
19 | T05 | Conduct scavenging of ePHI at rest | V04 | Insecure Boot / Boot for external media | A05 | Patient health information at rest | The threat actor actively steals data off the navigation unit by booting the system with an manipulated / self owned operating system to access the harddrive data and bypass the user role system provided by the original Stryker operating system. | [RARC.RI190] | MEDIUM | C3: Boot menu is protected by password C5: It shall not be possible to manipulate system files or install malicious software in these sections | LOW | N/A
20 | T24 | Manipulation of navigation camera firmware / memory | V24 | Unprotected hardware | A09 | Navigation Accuracy | The threat actor manipulates the memory / firmware of navigation camera with the goal to bring the calibration off / manipulate the firmware to reduce navigation accuracy. | [RARC.RI190] | MEDIUM | C2: A firewall shall be installed and activated. Unused ports and services disabled C8: Only ‘view-only’ remote access to the navigation system when a “Navigation User” is logged in C9: Protect FP8000 camera from being accessed from the hospital network | LOW | N/A
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - CVSSv3.0
Common Vulnerability Scoring System (CVSS v3.0)
Exploitability Metrics | [12 empty] | Likelihood of Attack Initiation
Attack Vector |  |  | Attack Complexity |  |  | Privelege Required |  |  |  | User Interaction |  |  |  | Rating | Score
Metric | Value | Code | Metric | Value | Code | Metric | Value |  | Code | Metric | Value | Code |  | Very Low | 0.04
Network | 0.85 | N | Low | 0.77 | L | None | 0.85 | 0.85 | N | None | 0.85 | N |  | Low | 0.2
Adjacent Network | 0.62 | A | High | 0.44 | H | Low | 0.62 | 0.68 | L | Required | 0.62 | R |  | Moderate | 0.5
Local | 0.55 | L |  |  |  | High | 0.27 | 0.5 | H |  |  |  |  | High | 0.8
Physical | 0.2 | P | [11 empty] | Very High | 1
Technical Impact Metrics | [12 empty] | In Scope | Yes
Confidentiality, Integrity, Availability Impact | [13 empty] | No
Metric | Value | Code
None | 0 | N | ISCBase = 1 - [(1−ImpactConf) × (1−ImpactInteg) × (1−ImpactAvail)]
Low | 0.22 | L
High | 0.56 | H
Scope
Unchanged | An exploited vulnerability can only affect resources managed by the same authority. In this case the vulnerable component and the impacted component are the same. | [10 empty] | U
Changed | An exploited vulnerability can affect resources beyond the authorization privileges intended by the vulnerable component. In this case the vulnerable component and the impacted component are different. | [10 empty] | C
ASSUMPTIONS:
Base metrics | For the purposes of the medical device only Base metrics are considered. The Base metric group represents the intrinsic characteristics of a vulnerability that are constant over time and across user environments. It is composed of two sets of metrics: the Exploitability metrics and the Impact metrics. The Exploitability metrics reflect the ease and technical means by which the vulnerability can be exploited. That is, they represent characteristics of the thing that is vulnerable, which we refer to formally as the vulnerable component. On the other hand, the Impact metrics reflect the direct consequence of a successful exploit, and represent the consequence to the thing that suffers the impact, which we refer to formally as the impacted component. The document only considers the mandatory base metric since the device is typically utilized in tightly controlled user environments such as hospitals and this is already a consideration of this assessment document. The changing charecteristics of vulnerabilities will be assessed seperately through the software development lifecycle
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## Reference - Threat Source
Threat Sources
Adversarial Threat |  |  | Non-Adverserial Threat
ID# | Threat Source | In Scope (Y/N) | ID# | Source | In Scope (Y/N)
TSA-1 | Individual (Disgruntled/Ex-Employees, Outsider, Insider, Trusted Insider, Privileged Insider) | Y | TSN-1 | Accidental (Priveleged User/Administrator, inexperienced user, inexperienced installer, inexperienced maintainer, unintentional misuse) | Y
TSA-2 | Organization (Competitor, Supplier, Partner, Customer, Researcher) | Y | TSN-2 | Researchers (Professional Security, Academic) | Y
TSA-3 | Script Kiddies | Y | TSN-3 | Vulnerable systems/devices connected to device (e.g., via RS-232, USB, or other connections) | Y
TSA-4 | Political Activists (Hactivists, Anonymous, Wikileaks) | N | TSN-4 | Incompatible Software (OS, Networking, Applications) | Y
TSA-5 | Organized Crime (Cyber Terrorists) | N | TSN-5 | Environmental Impact (IT equipment, Temperature/Humidity Controls, RF Interference) | Y
TSA-6 | Nation States | N | TSN-6 | Natural/Man-Made Disaster (Fire, Flood/Tsunami, Windstorm/Tornado, Earthquake, Bombing, Telecommunications/Power Failure) | N
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## OLD - Threat Assessment (hidden)
Adversarial Threat Events
# | Threat Event | Description | Threat Source | Relevance of Threat Event (L/M/H) | Liklihood of Threat Event Initiation (L/M/H) | Liklihood of Threat Event Resulting in Adverse Impacts (L/M/H) | Overall Likelihood of Threat Event
TEA-1 | Perform perimeter network reconnaissance/scanning. | Adversary uses commercial or free software to scan organizational perimeters to obtain a better understanding of the information technology infrastructure and improve the ability to launch successful attacks. | TSA-1 | Possible | Very Low | Very Low | Very Low

## OLD - Risk Controls (hidden)
Feature/Function | Risk Controls