## Header (Optional)
Product Security Standard Assessment
Ortho Guidance Precision Knee System
Ortho Guidance Precision Knee Software
Document Control
| Author: | Nitin Sharma
| Document No.: | D0000055901
| Revision No.: | AB.6
| Project: | Precision Knee 6.0 | DC-0000002518
| Project Lead: | Raman Bhardwaj
Change Records
| Rev. No. | Comment | Date | Author
| AA | Initial revision | Refer to onePLM | Nitin Sharma
| AB | Added reference to SRS requirements for Disk SPace Check, Log Files not containing patient name, unique folder name to store patient data. Per D0000003423 Product Security Development Life Cycle Appendix B, NIST SP 800-53 controls can be mapped to IEC 62443-3-3 controls. Therefore, the team has selected to use the IEC 62443-3-3 standard for the product security standard assessment | Refer to onePLM | Refer to onePLM
Note on B19: The use of this table is optional. This sheet can be removed from the form if not required.

## Document References (Optional)
Product Security Standard Assessment
References
ID | Title | Rev.* | Doc. No.
DI_SPIDER | DI - Spider - Spider (DC-0000002512 Spider) [DI_SPIDER] | Auto | D0000024194
DI | DI - Flex - Platform (0000013017 - Flex Platform) [DI] | Auto | D0000063236NAV
SDS_OS | DS - Flex - Operating System (0000013017 - Flex Platform) [SDS_OS] | Auto | D0000071649NAV
SRS_OPK | SRS - Ortho Guidance Precision Knee System - Ortho Guidance Precision Knee Software (DC-0000002518 Precision Knee 6.0) [SRS_OPK] | Auto | D0000039056
SDS_APPM | DS - Flex - Application Manager (0000013017 - Flex Platform) [SDS_APPM] | Auto | D0000023166
SADD_OS | SADD - Flex Platform - Operating System (0000013017 - Flex Platform) [SADD_OS] | Auto | D0000069594NAV
DI_CAMERA | DI - FP8000 Localization System - (0000012975 - FP8000 Camera) [DI_CAMERA] | Auto | D0000062256NAV
Note on A9: This table is only required if OfficeTrace is used. This sheet can be removed from the form if not required.

## NIST (hidden)
Product Security Standard Assessment - NIST SP 800-53 Controls
INSTRUCTIONS FOR USE:
1. Based upon identified User Needs and System Architecture, determine which, if any, of the 19 Security Capabilities are applicalbe to this product design. (Ref IEC/TR 80001-2-2)
A. If it is determined at a capability is not in scope for this product, provide a rationale of why it is not in scope.
2. Determine the System Security level for the product under development and provide rationale for this classification. (Ref FIPS 199 Table 1 if utilzing NIST SP 800-53 ) (Note: Currently Medical Devices are considered High System Security Level)
3. For those Security Capabilities identified as in scope in step 1, and the System Security levels determined in step 2, evaluate and identify which Security Controls will be utilized in the design. Include additional controls/requirements based upon the determined System Security Level. (Ref NIST SP 800-53 Table D-2,)
A. If a particular Security code that is specified per step 3 will not be utilized, provide a rationale.
4. If PHI and/or PII will be part of the product system design, then additional Security Controls need to be evaluated and added if applicable (Ref ID 200 & 210)
|  | Pick your target Security Level for System: | [8 empty] | Low
|  |  | SL | SSL | [6 empty] | Moderate
|  | Confidentiality | High | High | [6 empty] | High
|  | Integrity | Moderate
|  | Availability | Low
No
| IEC 80001-2-2 |  |  | IEC 80001-2-8
|  |  |  | System Security level: SP 800-53 Low Mod High | [5 empty] | System Security level Rationale
ID | Security Capability (Applicable = Design Inputs) | In scope of the product? | Rationale (if not utilizing capability) | Security Controls | System Security Level Baseline Low Moderate High |  |  | Additional Control/Requirement Enhancements based upon System Security Level | Rationale for not utilizing control if identified per Security Level | Traceability reference (SRS,SDD, etc)
10 | ALOF - System shall provide automatic logoff after a period of inactivity | Yes |  | AC-1 Access control policy and management | Yes | Yes | Yes
|  |  |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-7 Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-11 Session lock | No | Yes | Yes
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AC-24 Access control decisions | No | No | No
|  |  |  | CM-4 Security impact analysis | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IA-4 Identifier management | Yes | Yes | Yes
|  |  |  | IA-11 Re-authentication | No | No | No
20 | AUDT - System shall provide audit controls documenting who is doing what with HEALTH DATA | Yes |  | AC-21 Information sharing | No | No | No
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AU-1 Audit and accountability policy and procedures | Yes | Yes | Yes
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | AU-3 Content of audit records | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | AU-4 Audit storage capacity | Yes | Yes | Yes
|  |  |  | AU-5 Response to audit processing failures | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | AU-6 Audit review, analysis and reporting | Yes | Yes | Yes
|  |  |  | AU-7 Audit reduction and report generation | No | Yes | Yes
|  |  |  | AU-8 Time stamps | Yes | Yes | Yes
|  |  |  | AU-9 Protection of audit information | Yes | Yes | Yes
|  |  |  | AU-10 Non-repudiation | No | No | Yes
|  |  |  | AU-11 Audit record retention | Yes | Yes | Yes
|  |  |  | AU-12 Audit generation | Yes | Yes | Yes
|  |  |  | AU-13 Monitoring for information disclosure | No | Yes | Yes
|  |  |  | AU-14 Session audit | No | Yes | Yes
|  |  |  | AU-15 Alternate audit capacity | No | No | No
|  |  |  | AU-16 Cross-organizational auditing | No | No | No
30 | AUTH - The system shall provide role based controlled access to HEALTH DATA and functions - access to be provided only as necessary to perform the tasks required consistent with intended use | Yes |  | AC-1: Access Control Policy and Management | Yes | Yes | Yes
|  |  |  | AC-2: Account Management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-3: Access enforcement | Yes | Yes | Yes
|  |  |  | AC-5: Separation of duties | No | Yes | Yes
|  |  |  | AC-6: Least Privilege | No | Yes + (1) (2) | Yes + (1) (2) | Mod (1) (2), High (1) (2)
|  |  |  | AC-7: Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-17: Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | AC-18 Wireless access | Yes | Yes + (1) | Yes + (1) (2) (4) (5) | Mod (1) , High (1) (2) (4) (5)
|  |  |  | AC-19 Access control for mobile devices | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | AC-21 Information sharing | No | No | No
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AC-24 Access control decisions | No | No | No
|  |  |  | PL-4 Rules of behavior | Yes | Yes | Yes
40 | CNFS - System shall provide the Hospital system administrator the ability to configure the products security capabilities | Yes |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-5 Separation of duties | No | Yes | Yes
|  |  |  | AC-6 Least privilege | No | Yes + (1) (2) | Yes + (1) (2) | Mod (1) (2), High (1) (2)
|  |  |  | CM-1 Configuration management policy and procedures | Yes | Yes | Yes
|  |  |  | CM-2 Baseline configuration | Yes | Yes + (1) (3) (4) | Yes + (1) (2) (3) (5) (6) | Mod (1) (3) (4), High (1) (2) (3) (5) (6)
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-4 Security impact analysis | Yes | Yes | Yes + (1) | High (1)
|  |  |  | CM-5 Access restrictions for change | Yes | Yes | Yes + (1) (2) (3) | High (1) (2) (3)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | CM-7 Least functionality | Yes | Yes | Yes
|  |  |  | CM-9 Configuration management plan | No | Yes | Yes
|  |  |  | SA-10 Developer configuration management | No | Yes | Yes
50 | CSUP - The system shall provide the ability of on-site service staff, remote service staff, or authorized customer staff to install/upgrade device's security patches. | Yes |  | AC-17 Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | CM-2 Baseline configuration | Yes | Yes + (1) (3) (4) | Yes + (1) (2) (3) (5) (6) | Mod (1) (3) (4), High (1) (2) (3) (5) (6)
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-4 Security impact analysis | Yes | Yes | Yes + (1) | High (1)
|  |  |  | CM-5 Access restrictions for change | Yes | Yes | Yes + (1) (2) (3) | High (1) (2) (3)
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
|  |  |  | IA-9 Service identification and authentication | No | No | No
|  |  |  | MA-1 System maintenance policy and procedures | Yes | Yes | Yes
|  |  |  | MA-2 Controlled maintenance | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | MA-3 Maintenance tools | No | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MA-4 Nonlocal maintenance | Yes | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MA-5 Maintenance personnel | Yes | Yes | Yes
|  |  |  | MA-6 Timely maintenance | No | Yes | Yes
|  |  |  | MP-1 Media protection policy and procedures | Yes | Yes | Yes
|  |  |  | SA-8 Security engineering principles | No | Yes | Yes
|  |  |  | SA-11 Developer security testing and evaluation | No | Yes | Yes
|  |  |  | SA-14 Criticality analysis | No | No | No
|  |  |  | SI-11 Error handling | No | Yes | Yes
60 | DIDT - The system shall provide the ability to directly remove information that allows identification of a person. | Yes |  | AC-8 System use notification | Yes | Yes | Yes
|  |  |  | AC-21 Information sharing | No | No | No
|  |  |  | AC-23 Data mining protection | No | No | No
|  |  |  | AR-7 Privacy-enhanced system design and development | No | No | No
|  |  |  | AT-1 Security assurance and training policy and protection | Yes | Yes | Yes
|  |  |  | AU-3 Content of audit records | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | AU-9 Protection of audit information | Yes | Yes | Yes
|  |  |  | AU-11 Audit record retention | Yes | Yes | Yes
|  |  |  | DM-1 Minimization of personally identifiable information | No | No | No
|  |  |  | DM-2 Data retention and disposal | No | No | No
70 | DTBK - The system shall have an integral data backup capability to recover after damage or destruction of device data, hardware or software. | Yes |  | AU-9 Protection of audit information | Yes | Yes | Yes
|  |  |  | CM-1 Configuration management policy and procedure | Yes | Yes | Yes
|  |  |  | CM-2 Baseline configuration | Yes | Yes + (1) (3) (4) | Yes + (1) (2) (3) (5) (6) | Mod (1) (3) (4), High (1) (2) (3) (5) (6)
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-5 Access restrictions for changes | Yes | Yes | Yes + (1) (2) (3) | High (1) (2) (3)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | CP-1 Contingency planning policy and procedures | Yes | Yes | Yes
|  |  |  | CP-2 Contingency plan | Yes | Yes + (1) | Yes + (1) (2) (3) | Mod (1), High (1) (2) (3)
|  |  |  | CP-3 Contingency training | Yes | Yes | Yes + (1) | High (1)
|  |  |  | CP-4 Contingency plan testing | Yes | Yes + (1) | Yes + (1) (2) (4) | Mod (1), High (1) (2) (4)
|  |  |  | CP-6 Alternate storage site | No | Yes + (1) (3) | Yes + (1) (2) (3) | Mod (1) (3), High (1) (2) (3)
|  |  |  | CP-7 Alternate processing site | No | Yes + (1) (2) (3) (5) | Yes + (1) (2) (3) (4) (5) | Mod (1) (2) (3) (5), High (1) (2) (3) (4) (5)
|  |  |  | CP-8 Telecommunications services | No | Yes + (1) (2) | Yes + (1) (2) (3) (4) | Mod (1) (2), High (1) (2) (3) (4)
|  |  |  | CP-9 Information system backup | Yes | Yes + (1) | Yes + (1) (2) (3) | Mod (1), High (1) (2) (3)
|  |  |  | CP-10 Information system recovery and reconstitution | Yes | Yes + (2) (3) | Yes + (2) (3) (4) | Mod (2) (3), High (2) (3) (4)
|  |  |  | CP-13 Alternative security mechanisms | No | No | No
|  |  |  | IR-1 Incident response policy and procedures | Yes | Yes | Yes
|  |  |  | IR-2 Incident response training | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | IR-3 Incident response testing | No | Yes | Yes + (1) | High (1)
|  |  |  | IR-4 Incident handling | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-5 Incident monitoring | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IR-6 Incident reporting | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-7 Incident response assistance | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-8 Incident response plan | Yes | Yes | Yes
|  |  |  | IR-9 Information spillage response | No | No | No
|  |  |  | IR-10 Integrated information security analysis team | No | No | No
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | PM-9 RISK MANAGEMENT strategy | Yes | Yes | Yes
80 | EMRG - The system shall provide users the ability to access private data in case of an emergency that requires immediate access to private data. | Yes |  | AC-1 Access control policy and management | Yes | Yes | Yes
|  |  |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-14 Permitted actions without identification or authentication | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
|  |  |  | RA-5 Vulnerability scanning | Yes | Yes + (1) | Yes + (1) (2) (3) (4) (5) (7) | Mod (1), High (1) (2) (3) (4) (5) (7)
90 | IGAU - The device shall ensure the integrity of stored data with implicit/explicit error detection/correction technology | Yes |  | SA-13 Trustworthiness | No | No | Yes
|  |  |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SC-13 Cryptographic protection | Yes | Yes | Yes
|  |  |  | SC-17 Public key infrastructure certificates | No | Yes | Yes
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | SI-7 Software and information integrity | No | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | SI-10 Information input validation | No | Yes | Yes
100 | MLDP - The system shall support the use of anti-malware mechanism | Yes |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | IR-1 Incident response policy and procedures | Yes | Yes | Yes
|  |  |  | IR-2 Incident response training | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | IR-3 Incident response testing | No | Yes | Yes + (1) | High (1)
|  |  |  | IR-4 Incident handling | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-5 Incident monitoring | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IR-6 Incident reporting | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-7 Incident response assistance | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-8 Incident response plan | Yes | Yes | Yes
|  |  |  | MA-3 Maintenance tools | No | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MP-2 Media access | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | RA-5 Vulnerability scanning | Yes | Yes + (1) | Yes + (1) (2) (3) (4) (5) (7) | Mod (1), High (1) (2) (3) (4) (5) (7)
|  |  |  | SA-4 Acquisition PROCESS | Yes | Yes + (1) (4) | Yes + (1) (2) (4) | Mod (1) (4), High (1) (2) (4)
|  |  |  | SA-8 Security engineering principles | No | Yes | Yes
|  |  |  | SA-12 Supply chain protection | No | No | Yes
|  |  |  | SA-13 Trustworthiness | No | No | Yes
|  |  |  | SC-7 Boundary protection | Yes | Yes + (1) (2) (3) (4) (5) (7) | Yes + (1) (2) (3) (4) (5) (6) (7) (8) | Mod (1) (2) (3) (4) (5) (7), High (1) (2) (3) (4) (5) (6) (7) (8)
|  |  |  | SC-26 Honeypots | No | No | No
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
|  |  |  | SC-30 Concealment and misdirection | No | No | No
|  |  |  | SC-34 Non-modifiable executable programs | No | No | No
|  |  |  | SC-35 Honey clients | No | No | No
|  |  |  | SC-37 Out-of-band channels | No | No | No
|  |  |  | SC-44 Detonation chambers | No | No | No
|  |  |  | SI-2 Flaw remediation | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | SI-4 Information system monitoring | No | Yes + (2) (4) (5) (6) | Yes + (2) (4) (5) (6) | Mod (2) (4) (5) (6), High (2) (4) (5) (6)
|  |  |  | SI-7 Software and information integrity | No | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | SI-15 Information output filtering | No | No | No
110 | NAUT - The system shall provide node authentication to ensure that only known devices can connect with each other and share data | Yes |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-7 Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-14 Permitted actions without identification or authentication | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | AC-17 Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | AC-18 Wireless access | Yes | Yes + (1) | Yes + (1) (2) (4) (5) | Mod (1) , High (1) (2) (4) (5)
|  |  |  | AC-19 Access control for mobile devices | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | AU-10 Non-repudiation | No | No | Yes
|  |  |  | CM-1 Configuration management policy and procedures | Yes | Yes | Yes
|  |  |  | CM-3 Configuration change control | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
120 | PAUT - The system shall provide the ability to authenticate users | Yes |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AC-7 Unsuccessful logon attempts | Yes | Yes | Yes
|  |  |  | AC-14 Permitted actions without identification or authentication | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | AC-17 Remote access | Yes | Yes + (1) (2) (3) (4) (5) (7) (8) | Yes + (1) (2) (3) (4) (5) (7) (8) | Mod (1) (2) (3) (4) (5) (7) (8), High (1) (2) (3) (4) (5) (7) (8)
|  |  |  | AC-18 Wireless access | Yes | Yes + (1) | Yes + (1) (2) (4) (5) | Mod (1) , High (1) (2) (4) (5)
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | AU-10 Non-repudiation | No | No | Yes
|  |  |  | CM-1 Configuration management policy and procedures | Yes | Yes | Yes
|  |  |  | IA-1 Identification and authentication policy and procedures | Yes | Yes | Yes
|  |  |  | IA-2 Identification and authentication (organizational users) | Yes + (1) | Yes + (1) (2) (3) (8) | Yes + (1) (2) (3) (4) (8) (9) | Low (1), Mod (1) (2) (3) (8), High (1) (2) (3) (4) (8) (9)
|  |  |  | IA-4 Identifier management | Yes | Yes | Yes
|  |  |  | IA-5 Authenticator management | Yes + (1) | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Low (1), Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | IA-7 Cryptographic module authentication | Yes | Yes | Yes
|  |  |  | IA-8 Identification and authentication (non-organizational users) | Yes | Yes | Yes
|  |  |  | IA-10 Adaptive identification and authentication | No | No | No
|  |  |  | IA-11 Re-authentication | No | No | No
|  |  |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
130 | PLOK - The system components that maintain private data shall be physically secure | Yes |  | AC-1: Access Control Policy and Management | Yes | Yes | Yes
|  |  |  | AU-2 Audit events | Yes | Yes + (3) (4) | Yes + (3) (4) | Mod (3) (4), High (3) (4)
|  |  |  | CA-7 Continuous monitoring | Yes | Yes | Yes
CP-6 Alternate storage site
|  |  |  | MP-2 Media access | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | MP-4 Media | No | Yes | Yes
|  |  |  | MP-7 Media use | No | No | No
|  |  |  | PE-1 Physical and environmental protection policy and procedures | Yes | Yes | Yes
|  |  |  | PE-2 Physical access authorizations | Yes | Yes | Yes
|  |  |  | PE-3 Physical access control | Yes | Yes | Yes + (1) | High (1)
|  |  |  | PE-4 Access control for transmission medium | No | Yes | Yes
|  |  |  | PE-5 Access control for output devices | No | Yes | Yes
|  |  |  | PE-6 Monitoring physical access | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | PE-9 Power equipment and power cabling | No | Yes | Yes
|  |  |  | PE-18 Location of information system components | No | Yes | Yes + (1) | High (1)
|  |  |  | PL-2 System security plan | Yes | Yes | Yes
|  |  |  | RA-5 Vulnerability scanning | Yes | Yes + (1) | Yes + (1) (2) (3) (4) (5) (7) | Mod (1), High (1) (2) (3) (4) (5) (7)
|  |  |  | SC-8 Transmission confidentiality and integrity | No | Yes + (1) | Yes + (1) | Mod (1), High (1)
140 | RDMP - The system software development plan shall address security support of 3rd party components throughout system life cycle | Yes |  | MA-1 System maintenance policy and procedures | Yes | Yes | Yes
|  |  |  | MA-2 Controlled maintenance | Yes | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | MA-3 Maintenance tools | No | Yes + (1) (2) | Yes + (1) (2) (3) | Mod (1) (2), High (1) (2) (3)
|  |  |  | MA-6 Timely maintenance | No | Yes | Yes
|  |  |  | MP-1 Media protection policy and procedures | Yes | Yes | Yes
|  |  |  | MP-8 Media downgrading | No | No | No
|  |  |  | SA-1 System and services acquisition policy and procedures | Yes | Yes | Yes
|  |  |  | SA-3 System development life cycle | Yes | Yes | Yes
|  |  |  | SA-4 Acquisition PROCESS | Yes | Yes + (1) (4) | Yes + (1) (2) (4) | Mod (1) (4), High (1) (2) (4)
|  |  |  | SA-5 Information system documentation | Yes | Yes + (1) (3) | Yes + (1) (2) (3) | Mod (1) (3), High (1) (2) (3)
|  |  |  | SA-8 Security engineering principles | No | Yes | Yes
|  |  |  | SA-9 External information system services | Yes | Yes | Yes
|  |  |  | SA-10 Developer configuration management | No | Yes | Yes
|  |  |  | SA-11 Developer security testing and evaluation | No | Yes | Yes
|  |  |  | SA-12 Supply chain protection | No | No | Yes
|  |  |  | SA-15 Development PROCESS, standards and tools | No | No | No
|  |  |  | SA-16 Developer-provided training | No | No | No
|  |  |  | SA-17 Developer security architecture and design | No | No | No
|  |  |  | SA-21 Developer screening | No | No | No
150 | SAHD - System shall provide hardening features to guard against cyber attacks and malware | Yes |  | AC-19 Access control for mobile devices | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | CM-6 Configuration settings | Yes | Yes + (3) | Yes + (1) (2) (3) | Mod (3), High (1) (2) (3)
|  |  |  | CM-7 Least functionality | Yes | Yes | Yes
|  |  |  | SA-14 Criticality analysis | No | No | No
|  |  |  | SA-17 Developer security architecture and design | No | No | No
|  |  |  | SA-18 Tamper resistance and detection | No | No | No
|  |  |  | SC-25 Thin nodes | No | No | No
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
|  |  |  | SC-29 Heterogeneity | No | No | No
|  |  |  | SC-30 Concealment and misdirection | No | No | No
|  |  |  | SC-31 Covert channel analysis | No | No | No
|  |  |  | SC-35 Honey clients | No | No | No
|  |  |  | SC-40 Wireless link protection | No | No | No
|  |  |  | SC-41 Port and I/O device access | No | No | No
|  |  |  | SC-42 Sensor capability and data | No | No | No
|  |  |  | SC-43 Usage restrictions | No | No | No
|  |  |  | SI-11 Error handling | No | Yes | Yes
160 | SGUD - The system service/user manuals shall provide security guidance for operators and administrators | Yes |  | AC-1 Access control policy and management | Yes | Yes | Yes
|  |  |  | AC-2 Account management | Yes | Yes + (1) (2) (3) (4) | Yes + (1) (2) (3) (4) | Mod (1) (2) (3) (4), High (1) (2) (3) (4)
|  |  |  | AT-1 Security awareness and training policy and procedures | Yes | Yes | Yes
|  |  |  | AT-2 Security awareness training | Yes | Yes | Yes
|  |  |  | AT-3 Security training | Yes | Yes | Yes
|  |  |  | CP-1 Contingency planning policy and procedures | Yes | Yes | Yes
|  |  |  | CP-2 Contingency plan | Yes | Yes + (1) | Yes + (1) (2) (3) | Mod (1), High (1) (2) (3)
|  |  |  | CP-3 Contingency training | Yes | Yes | Yes + (1) | High (1)
|  |  |  | IR-1 Incident response policy and procedures | Yes | Yes | Yes
|  |  |  | IR-2 Incident response training | Yes | Yes | Yes + (1) (2) | High (1) (2)
|  |  |  | IR-7 Incident response assistance | Yes | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | IR-8 Incident response plan | Yes | Yes | Yes
|  |  |  | PL-1 Security planning policy and procedures | Yes | Yes | Yes
|  |  |  | PL-2 System security plan | Yes | Yes | Yes
|  |  |  | PL-4 Rules of behavior | Yes | Yes | Yes
|  |  |  | PL-7 Security concept of operations | No | No | No
|  |  |  | PL-8 Information security architecture | No | No | No
|  |  |  | PS-1 Personnel security policy and procedures | Yes | Yes | Yes
|  |  |  | SA-4 Acquisition PROCESS | Yes | Yes + (1) (4) | Yes + (1) (2) (4) | Mod (1) (4), High (1) (2) (4)
|  |  |  | SA-5 Information system documentation | Yes | Yes + (1) (3) | Yes + (1) (2) (3) | Mod (1) (3), High (1) (2) (3)
|  |  |  | SA-16 Developer-provided training | No | No | No
|  |  |  | SC-1 System and communications protection policy and procedures | Yes | Yes | Yes
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | SI-2 Flaw remediation | Yes | Yes + (2) | Yes + (1) (2) | Mod (2), High (1) (2)
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
|  |  |  | SI-4 Information system monitoring | No | Yes + (2) (4) (5) (6) | Yes + (2) (4) (5) (6) | Mod (2) (4) (5) (6), High (2) (4) (5) (6)
|  |  |  | SI-5 Security alerts, advisories, and directives | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SI-6 Security functionality VERIFICATION | No | No | Yes
|  |  |  | SI-7 Software and information integrity | No | Yes + (1) | Yes + (1) (2) | Mod (1), High (1) (2)
|  |  |  | SI-8 Spam protection | No | Yes | Yes + (1) | High (1)
|  |  |  | SI-10 Information input validation | No | Yes | Yes
|  |  |  | SI-11 Error handling | No | Yes | Yes
|  |  |  | SI-12 Information handling and retention | Yes | Yes | Yes
|  |  |  | SI-17 Fail-safe procedures | No | No | No
|  |  |  | PM-1 Information security program plan | Yes | Yes | Yes
|  |  |  | PM-9 RISK MANAGEMENT strategy | Yes | Yes | Yes
|  |  |  | PM-12 Insider threat program | Yes | Yes | Yes
|  |  |  | PM-14 Testing, training and monitoring | Yes | Yes | Yes
|  |  |  | PM-15 Contacts with security groups and associations | Yes | Yes | Yes
|  |  |  | PM-16 Threat awareness program | Yes | Yes | Yes
170 | STCF - The system shall establish technical controls to mitigate the potential for compromise to the integrity and confidentiality of health data stored on product or removable media | Yes |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SC-13 Cryptographic protection | Yes | Yes | Yes
|  |  |  | SC-17 Public key infrastructure certificates | No | Yes | Yes
|  |  |  | SC-28 Protection of information at rest | No | Yes | Yes
180 | TXCF - The system shall ensure confidentiality of transmitted Health Data | Yes |  | PE-4 Access control for transmission medium | No | Yes | Yes
|  |  |  | SC-1 System and communications protection policy and procedures | Yes | Yes | Yes
|  |  |  | SC-8 Transmission confidentiality and integrity | No | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | SC-12 Cryptographic key establishment and management | Yes | Yes | Yes + (1) | High (1)
|  |  |  | SC-13 Cryptographic protection | Yes | Yes | Yes
190 | TXIG - The system shall ensure integrity of transmitted Health Data | Yes |  | PE-4 Access control for transmission medium | No | Yes | Yes
|  |  |  | SC-1 System and communications protection policy and procedures | Yes | Yes | Yes
|  |  |  | SC-8 Transmission confidentiality and integrity | No | Yes + (1) | Yes + (1) | Mod (1), High (1)
|  |  |  | SI-1 System and information integrity policy and procedures | Yes | Yes | Yes
|  |  |  | SI-3 Malicious code protection | Yes | Yes + (1) (2) (3) | Yes + (1) (2) (3) | Mod (1) (2) (3), High (1) (2) (3)
200 | PII | Yes |  | AP-1 Authority to Collect
AP-2 Purpose Specification
AR-1 Governance and Privacy Program
AR-2 Privacy Impact and Risk Assessment
AR-3 Privacy Requirements for Contractors and Service Providers
AR-4 Privacy Monitoring and Auditing
AR-5 Privacy Awareness and Training
AR-6 Privacy Reporting
AR-7 Privacy-enhanced system design and development
AR-8 Accounting of Disclosures
DI-1 Data Quality
DI-2 Data Integrity and Data Integrity Board
DM-1 Minimization of personally identifiable information
DM-2 Data retention and disposal
DM-3 Minimization of PII Used in Testing, Training, and IP
IP-1 Consent
IP-2 Individual Access
IP-3 Redress
IP-4 Complaint Management
SE-1 Inventory of Personally Idntifiable Information
SE-2 Privacy Incident Response
TR-1 Privacy Notice
TR-2 System of Records Notices and Privacy Act
TR-3 Dissemination of Privacy Program Information
UL-1 Internal Use
UL-2 Information Sharing with Third Parties
210 | PHI | Yes |  | 164.312(a)(2)(i) Unique User Identification (R): Assign a unique name and/or number for identifying and tracking user identity.
AC-2 Account management
AC-3: Access enforcement
IA-2 Identification and authentication (organizational users)
IA-3 Device Identification and Authentication
IA-4 Identifier management
164.312(a)(2)(ii) Emergency Access Procedure (R): Establish (and implement as needed) procedures for obtaining necessary electronic protected health information during an emergency.
AC-2 Account management
AC-3: Access enforcement
CP-2 Contingency Plan
164.312(a)(2)(iii) Automatic Logoff (A): Implement electronic procedures that terminate an electronic session after a predetermined time of inactivity.
AC-11 Session lock
AC-12 Session Termination
164.312(a)(2)(iv) Encryption and Decryption (A): Implement a mechanism to encrypt and decrypt electronic protected health information.
AC-3: Access enforcement
SC-13 Cryptographic protection
164.312(b) Audit Controls: Implement hardware, software, and/or procedural mechanisms that record and examine activity in information systems that contain or use electronic protected health information.
AU-1 Audit and accountability policy and procedures
AU-2 Audit events
AU-3 Content of audit records
AU-4 Audit storage capacity
AU-6 Audit review, analysis and reporting
AU-7 Audit reduction and report generation
164.312(c)(1) Integrity: Implement policies and procedures to protect electronic protected health information from improper alteration or destruction.
CP-9 Information system backup
MP-2 Media access
MP-5 Media Transport
SC-8 Transmission confidentiality and integrity
SI-1 System and information integrity policy and procedures
SI-7 Software and information integrity
164.312(c)(2) Mechanism to Authenticate Electronic Protected Health Information (A): Implement electronic mechanisms to corroborate that electronic protected health information has not been altered or destroyed in an unauthorized manner.
SC-8 Transmission confidentiality and integrity
SI-7 Software and information integrity
164.312(d) Person or Entity Authentication: Implement procedures to verify that a person or entity seeking access to electronic protected health information is the one claimed.
IA-2 Identification and authentication (organizational users)
IA-3 Device Identification and Authentication
IA-4 Identifier management
164.312 (e)(1) Transmission Security: Implement technical security measures to guard against unauthorized access to electronic protected health information that is being transmitted over an electronic communications network.
SC-9 Transmission Confidentiality
164.312(e)(2)(i) Integrity Controls (A): Implement security measures to ensure that electronically transmitted electronic protected health information is not improperly modified without detection until disposed of.
SC-8 Transmission confidentiality and integrity
SI-7 Software and information integrity
164.312(e)(2)(ii) Encryption (A): Implement a mechanism to encrypt electronic protected health information whenever deemed appropriate.
SC-9 Transmission Confidentiality
SC-12 Cryptographic key establishment and management
SC-13 Cryptographic protection
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.

## IEC
Product Security Standard Assessment - ISO IEC 62443-3-3 Controls
INSTRUCTIONS FOR USE:
1. Based upon identified User Needs and System Architecture, determine which, if any, of the 19 Security Capabilities are applicalbe to this product design. (Ref IEC/TR 80001-2-2)
A. If it is determined at a capability is not in scope for this product, provide a rationale of why it is not in scope.
2. Determine the System Security level for the product under development and provide rationale for this classification. (Ref IEC 62443-3-3 Annex A.)
3. For those Security Capabilities identified as in scope in step 1, and the System Security levels determined in step 2, evaluate and identify which Security Controls will be utilized in the design. Iinclude additional controls/requirements based upon the determined System Security Level. (Ref IEC 62443-3-3 Table B.1)
A. If a particular Security code that is specified per step 3 will not be utilized, provide a rationale.
4. If PHI and/or PII will be part of the product system design, then additional Security Controls need to be evaluated and added if applicable (Ref ID 200 & 210)
Pick your target Security Level for each chapter here:
|  | Chapter | Target Security Level | Security Level Rationale: |  |  |  | Scope
|  | 1 | SL-3 | Identification and authentication control: The Navigation System is used in an access controlled OR environment by trained personnel. This already adds an additional layer of proctection for physical access. As the attack could also be initiated via network the Navigation System needs to be protected against targeted attacks by entities using sophisticated means with specific skills and moderate motivation. |  |  |  | Identify and authenticate all users (humans, software processes and devices) before allowing them to access to the control system.
|  | 2 | SL-2 | Use Control: The Navigation System is used in an access controlled OR environment and all user are authenticated by the system before alllowing them to access the system. The Hospital IT is managing the user accounts on the Navigation System and only a limited group of trusted users has access to the system. Therefore a protection against circumvention by entities using simple means with low resources, generic skills and low motivation is sufficient. |  |  |  | Enforce the assigned privileges of an authenticated user (human, software process or device) to perform the requested action on the IACS and monitor the use of these privileges.
|  | 3 | SL-2 | The Navigation System is not a life sustaining device. During surgery the surgeon can check the system integrity/accuracy by touching known landmarks on the patient (landmark check). Therefore a protection by someone using simple means with low resources, generic skills and low motivation is sufficient. |  |  |  | Ensure the integrity of the IACS to prevent unauthorized manipulation.
|  | 4 | SL-2 | The Navigation System need to be protected againgst targeted attacks with simple means. Only a limited amount of patient data are stored on the systems. Therefore a protection against attackers with low motivation and resources is sufficient. |  |  |  | Ensure the confidentiality of information on communication channels and in data repositories to prevent unauthorized disclosure.
|  | 5 | SL-3 | Navigation System can be connected to hospital network infrastructure. A protection against attackers with moderate motivation and ressources is sufficient as the sytem is only connected to the internet via a secured hospital network. |  |  |  | Segment the control system via zones and conduits to limit the unnecessary flow of data.
|  | 6 | SL-1 | Navigation System is a sub system in a controlled hospital network. Hospital is responsible for the network integrity and therefore it is sufficient to be able to collect forensic data from the Navigation System when queried. No automatic reporting required. |  |  |  | Respond to security violations by notifying the proper authority, reporting needed evidence of the violation and taking timely corrective action when incidents are discovered.
|  | 7 | SL-1 | The Navigation System is not a life sustaining device. A denial of service would lead to an extended treatment time. The surgery can be conducted without using Navigation. Therefore it is sufficient that the system runs reliably under normal production conditions and prevents DoS situations caused by the casual or coincidental actions of an entity. |  |  |  | Ensure the availability of the control system against the degradation or denial of essential services.
| IEC 80001-2-2 |  |  | IEC 80001-2-8
|  |  |  | System Security level: ISO IEC 62443-3-3 SL-1 SL-2 SL-3 SL-4 | [8 empty] | System Security level Rationale | System Security level Rationale
ID | Security Capability (Applicable = Design Inputs) | In scope of the product? | Rationale (if not utilizing capability) | Security Controls | Chapter | Target Level | System Security Level Baseline SL-1 SL-2 SL-3 SL-4 |  |  |  | Additional Control/Requirement Enhancements based upon System Security Level | Rationale for not utilizing control if identified per Security Level | Q Traceability reference (SRS,SDD, etc) | Spider Traceability reference (SRS,SDD, etc) | Requirements
10 | ALOF - System shall provide automatic logoff after a period of inactivity | Yes |  | SR 2.5 Session Lock | 2 | SL-2 | Yes | Yes | Yes | Yes | n/a | Session lock not applicable to user accounts. Navigation System needs to be available during surgery to not interupt the surgical workflow even after a longer time of inactivity. | [SDS_APPM.TR37010: Session Lock] | [SDS_APPM.TR37010: Session Lock] | SR 2.5: The control system shall provide the capability to prevent further access by initiating a session lock after a configurable time period of inactivity or by manual initiation. The session lock shall remain in effect until the human user who owns the session or another authorized human user re-establishes access using appropriate identification and authentication procedures.
|  |  |  | SR 2.6 Remote session termination | 2 | SL-2 | No | Yes | Yes | Yes | n/a | No remote session available for user accounts. | n/a | n/a | SR 2.6: The control system shall provide the capability to terminate a remote session either automatically after a configurable time period of inactivity or manually by the user who initiated the session.
20 | AUDT - System shall provide audit controls documenting who is doing what with HEALTH DATA | No | Health data is stored encrypted on the Navigation System. The Navigation System manages only a limited amount of patient data. It does not distribute health data to other medical devices. | SR 2.8 Auditable events | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.9 Audit storage capacity | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 2.10 Response to audit processing failures | 2 | SL-2 | Yes | Yes | Yes | Yes
|  |  |  | SR 2.11 Timestamps | 2 | SL-2 | No | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2)
|  |  |  | SR 2.12 Non-repudiation | 2 | SL-2 | No | No | Yes | Yes + (1) | SL-4 (1)
|  |  |  | SR 3.9 Protection of audit information | 3 | SL-2 | No | Yes | Yes | Yes + (1) | SL-4 (1)
|  |  |  | SR 6.1 Audit reduction and report generation | 6 | SL-1 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1)
|  |  |  | SR 6.2 Continuous monitoring | 6 | SL-1 | No | Yes | Yes | Yes
30 | AUTH - The system shall provide role based controlled access to HEALTH DATA and functions - access to be provided only as necessary to perform the tasks required consistent with intended use | Yes |  | SR 1.3 Account management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | n/a | Navigation System provides the capability to support the management of all accounts by authorized users. [DI.SR61010: User management] [SDS_APPM.TR31010: Creation of user] [SDS_APPM.TR31020: User attributes] [SDS_APPM.TR31030: Deletion of users] [SDS_APPM.TR31040: Edit user] Unified account management: [SDS_APPM.TR32040: LDAP Configuration] [DI.SR61020: Network configuration] | Navigation System provides the capability to support the management of all accounts by authorized users. [SDS_APPM.TR31010: Creation of user] [SDS_APPM.TR31020: User attributes] [SDS_APPM.TR31030: Deletion of users] [SDS_APPM.TR31040: Edit user] Unified account management: [SDS_APPM.TR32040: LDAP Configuration] [SDS_APPM.TR32030: Network Configuration] | SR 1.3: The control system shall provide the capability to support the management of all accounts by authorized users, including adding, activating, modifying, disabling and removing accounts. RE 1: The control system shall provide the capability to support unified account management.
|  |  |  | SR 2.1 Authorization enforcement | 2 | SL-2 | Yes | Yes + (1) | Yes + (1) (2) (3) | Yes + (1) (2) (3) (4) | SL-2 (1) (2), SL-3 (1) (2) (3), SL-4 (1) (2) (3) (4) | n/a | Each user account has dedicated folders to store personal assets. Patient data is stored on a shared drive. Admin priviledges are required to modify system settings and the mapping of permissions to roles. [DI.SR512023: Principle of least privileges] [DI.SR512022: Local command input on operating system level is not possible] [SDS_OS.TR31460: Application user with limited rights] [SDS_OS.TR31450: No local login possible] [SDS_OS.TR31500: Application user with limited rights] [SDS_OS.TR31490: Limited user rights for Application Manager] | Each user account has dedicated folders to store personal assets. Patient data is stored on a shared drive. Admin priviledges are required to modify system settings and the mapping of permissions to roles. [DI_SPIDER.SR58030: Principle of least privileges] [DI_SPIDER.SR58020: Local command input on operating system level is not possible] [SDS_OS.TR31460: Application user with limited rights] [SDS_OS.TR31450: No local login possible] [SDS_OS.TR31500: Application user with limited rights] [SDS_OS.TR31490: Limited user rights for Application Manager] | SR 2.1 RE 1 : On all interfaces, the control system shall provide the capability to enforce authorizations assigned to all users (humans, software processes and devices) for controlling use of the control system to support segregation of duties and least privilege.
40 | CNFS - System shall provide the Hospital system administrator the ability to configure the products security capabilities | Yes |  | SR 1.3 Account management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See Above. | See Above. | See Above. | SR 1.3: The control system shall provide the capability to support the management of all accounts by authorized users, including adding, activating, modifying, disabling and removing accounts. RE 1: The control system shall provide the capability to support unified account management.
|  |  |  | SR 7.6 Network and security configuration settings | 7 | SL-1 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | n/a | Navigation System shall provide the capability to be configured according to recommended network and security configurations as described in Security Operations Manual. [SDS_OS.TR31470: Ethernet port protection] [DI.SR512021: I/O port usage is secured against unauthorised access] | Navigation System shall provide the capability to be configured according to recommended network and security configurations as described in Security Operations Manual. [SDS_OS.TR31470: Ethernet port protection] [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] | SR 7.6: The control system shall provide the capability to be configured according to recommended network and security configurations as described in guidelines provided by the control system supplier. The control system shall provide an interface to the currently deployed network and security configuration settings.
50 | CSUP - The system shall provide the ability of on-site service staff, remote service staff, or authorized customer staff to install/upgrade device's security patches. | Yes |  | No applicable SECURITY CONTROLS | 0 | #N/A | N/A | N/A | N/A | N/A | n/a | n/a | n/a | n/a
60 | DIDT - The system shall provide the ability to directly remove information that allows identification of a person. | Yes |  | SR 4.2 Information persistence | 4 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | Navigation System is normally rebooted every day or even after each case. Only one Application is running at once. Therefore a purge of the random access memory is not necessary. | Application Manager offers a data purging tool to remove patient data and other sensitive information (All user and runtime data) from the system. Secure data deletion is a requirement demanded by some customers. [SDS_APPM.TR35030: Reset to factory default] | Application Manager offers a data purging tool to remove patient data and other sensitive information (All user and runtime data) from the system. Secure data deletion is a requirement demanded by some customers. [SDS_APPM.TR35030: Reset to factory default] | SR 4.2: The control system shall provide the capability to purge all information for which explicit read authorization is supported from components to be released from active service and/or decommissioned. RE1: The control system shall provide the capability to prevent unauthorized and unintended information transfer via volatile shared memory resources
70 | DTBK - The system shall have an integral data backup capability to recover after damage or destruction of device data, hardware or software. | Yes |  | SR 2.8 Auditable events | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | Request errors, potential reconnaissance activity and audit log event are not available. Navigation System is not a critical infrastructure component. No extended audit functionality required. | - Access control: Login of different users is logged using the Appman. - Operating system events: Yes - Control system events: Log4Cpp application events (Application Manager and Navigation Application). - Backup events: Logged by Navigation Application. - Restore events: Logged by Navigation Application. - Configuration changes: Not implemened for Navigation Application config files but for other parts of the system. [DI.SR512023: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] [SADD_OS: SADD - Flex Platform - Operating System] | - Access control: Login of different users is logged using the Appman. - Operating system events: Yes - Control system events: Log4Cpp application events (Application Manager and Navigation Application). - Backup events: Logged by Navigation Application. - Restore events: Logged by Navigation Application. - Configuration changes: Not implemened for Navigation Application config files but for other parts of the system. [DI_SPIDER.SR58030: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] [SADD_OS: SADD - Flex Platform - Operating System] | SR 2.8 The control system shall provide the capability to generate audit records relevant to security for the following categories: access control, request errors, operating system events, control system events, backup and restore events, configuration changes, potential reconnaissance activity and audit log events. Individual audit records shall include the timestamp, source (originating device, software process or human user account), category, type, event ID and event result.
|  |  |  | SR 3.6 Deterministic output | 3 | SL-2 | Yes | Yes | Yes | Yes | n/a | User can check system functionality by doing a landmark check . No automatic fallback to predetermined state required. | All user and runtime data are cleared from the system by factory reset. Predetermined state can be restored. [SDS_APPM.TR35030: Reset to factory default] | All user and runtime data are cleared from the system by factory reset. Predetermined state can be restored. [SDS_APPM.TR35030: Reset to factory default] | SR 3.6: The control system shall provide the capability to set outputs to a predetermined state if normal operation cannot be maintained as a result of an attack.
|  |  |  | SR 7.3 Control system backup | 7 | SL-1 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) | Operating system backups are not necessary. The operating system can be installed from scratch if needed. | System recovery is done by Sales Rep with OS and application Installation media. PHI is not stored on the System partition. Patient data can be restored if regular backups are created. In worst case the loss of patient data is not considered being a critical event. Predetermined state can be maintained by factory reset. [SDS_APPM.TR35030: Reset to factory default] [SRS_OPK.F2270: Patient Backup/export] [SRS_OPK.F2280: Patient Restore/import] | System recovery is done by Sales Rep with OS and application Installation media. PHI is not stored on the System partition. Patient data can be restored if regular backups are created. In worst case the loss of patient data is not considered being a critical event. Predetermined state can be maintained by factory reset. [SDS_APPM.TR35030: Reset to factory default] [SRS_OPK.F2270: Patient Backup/export] [SRS_OPK.F2280: Patient Restore/import] | SR 7.3: The identity and location of critical files and the ability to conduct backups of user-level and system-level information (including system state information) shall be supported by the control system without affecting normal plant operations.
|  |  |  | SR 7.4 Control system recovery and reconstitution | 7 | SL-1 | Yes | Yes | Yes | Yes | n/a | Only critical files for the functionallity of the Navigation Application are restored. | Manual backup procedure for patient data is available in Navigation Application and described in IFU. Files critical for the surgery such as registration, are backuped for the duration of the surgery. By setting sensitive operating system parts to read-only it is not possible to manipulate system files or install malicious software in these sections. Predetermined state can be restored. [DI.SR512024: Operating System is protected against unauthorized manipulation] [SRS_OPK.F2200: Patient data file] [SRS_OPK.F2260: Patient folder] [SRS_OPK.F2270: Patient Backup/export] [SRS_OPK.F2280: Patient Restore/import] | Manual backup procedure for patient data is available in Navigation Application and described in IFU. Files critical for the surgery such as registration, are backuped for the duration of the surgery. By setting sensitive operating system parts to read-only it is not possible to manipulate system files or install malicious software in these sections. Predetermined state can be restored. [DI_SPIDER.SR58040: Operating System is protected against unauthorized manipulation] [SRS_OPK.F2200: Patient data file] [SRS_OPK.F2260: Patient folder] [SRS_OPK.F2270: Patient Backup/export] [SRS_OPK.F2280: Patient Restore/import] | SR 7.4: The control system shall provide the capability to recover and reconstitute to a known secure state after a disruption or failure.
80 | EMRG - The system shall provide users the ability to access private data in case of an emergency that requires immediate access to private data. | No | No emergency relevant data like e.g. blood type, medication is saved on the Navigation System. | SR 1.4 Identifier Management | 1 | SL-3 | Yes | Yes | Yes | Yes |  |  | n/a | n/a
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) |  | n/a | n/a
|  |  |  | SR 2.8 Auditable events | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) |  | n/a | n/a
90 | IGAU - The device shall ensure the integrity of stored data with implicit/explicit error detection/correction technology | Yes |  | SR 3.1 Communication Integrity | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | n/a | Communication to ConnectedCare server is protected by encryption Communication to active Navigation Tools is protected by CRC check. [DI_CAMERA.SR1540: Data protection] [DI_CAMERA.SR1530: Communication security] A secure shell encrypts the communication between host and client. [DI.SR512025: Remote access is secured] | Communication to ConnectedCare server is protected by encryption Communication to active Navigation Tools is protected by CRC check. [DI_CAMERA.SR1540: Data protection] [DI_CAMERA.SR1530: Communication security] A secure shell encrypts the communication between host and client. [DI_SPIDER.SR58050: Remote access is secured] | SR 3.1: The control system shall provide the capability to protect the integrity of transmitted information.
|  |  |  | SR 3.3 Security functionality VERIFICATION | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2) | Security functions cannot be deactivated by the user or admin. Therefore no verification of the security functions required. | n/a | n/a | SR 3.3 The control system shall provide the capability to support verification of the intended operation of security functions and report when anomalies are discovered during FAT, SAT and scheduled maintenance. These security functions shall include all those necessary to support the security requirements specified in this standard.
|  |  |  | SR 3.4 Software and information integrity | 3 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | Critical software configuration files and binaries cannot be accessed or modified by the user due to restricted permissions. Therefore no modification detection and reporting is required. | Restrictive user permissions [DI.SR512023: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] | Restrictive user permissions [DI_SPIDER.SR58030: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] | SR 3.4: The control system shall provide the capability to detect, record, report and protect against unauthorized changes to software and information at rest.
|  |  |  | SR 3.5 Input validation | 3 | SL-2 | Yes | Yes | Yes | Yes | n/a | No external input is imported for intended use of the navigation system. | n/a | n/a | SR 3.5: The control system shall validate the syntax and content of any input which is used as an industrial process control input or input that directly impacts the action of the control system.
|  |  |  | SR 3.7 Error handling | 3 | SL-2 | No | Yes | Yes | Yes | n/a | n/a | Error conditions relevant for the Navigation System (such as hard disk full, or navigation inaccuracy) can be identified and proper error handling is in place (e.g. hard disk check, landmark check). This shall be done in a manner which does not provide information that could be exploited by adversaries. [SRS_OPK.F3530: Confirm Verification Point Button (Femur)] [SRS_OPK.F3560: Confirm Verification Point Button (Tibia)] [SRS_OPK.F2010: Disk Space Check] | Error conditions relevant for the Navigation System (such as hard disk full, or navigation inaccuracy) can be identified and proper error handling is in place (e.g. hard disk check, landmark check). This shall be done in a manner which does not provide information that could be exploited by adversaries. [SRS_OPK.F3530: Confirm Verification Point Button (Femur)] [SRS_OPK.F3560: Confirm Verification Point Button (Tibia)] [SRS_OPK.F2010: Disk Space Check] | SR 3.7: The control system shall identify and handle error conditions in a manner such that effective remediation can occur. This shall be done in a manner which does not provide information that could be exploited by adversaries to attack the IACS unless revealing this information is necessary for the timely troubleshooting of problems.
100 | MLDP - The system shall support the use of anti-malware mechanism | Yes |  | SR 1.2 Software PROCESS and device identification and authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | USB storage Devices cannot be authenticated digitally. | For Remote access, Connected Care provides client/server authorization. Bootup ends in Application Manager, no other applications permitted. [DI.SR512026: Bootup is protected against manipulation] | For Remote access, Connected Care provides client/server authorization. Bootup ends in Application Manager, no other applications permitted. [DI_SPIDER.SR58060: Bootup is protected against manipulation] | SR 1.2: The control system shall provide the capability to identify and authenticate all software processes and devices. This capability shall enforce such identification and authentication on all interfaces which provide access to the control system to support least privilege in accordance with applicable security policies and procedures. RE 1: The control system shall provide the capability to uniquely identify and authenticate all software processes and devices.
|  |  |  | SR 2.3 Use Control for portable and mobile devices | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | USB storage devices are required to create/import backups. | I/O ports, only dedicated devices are allowed. [DI.SR512021: I/O port usage is secured against unauthorised access] [SDS_OS.TR31200: USB ports with restricted access] Data transfer from USB storage device is restricted. As user, USB storage device can only be accessed via Navigation Application for encrypted backup / restore | I/O ports, only dedicated devices are allowed. [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [SDS_OS.TR31200: USB ports with restricted access] Data transfer from USB storage device is restricted. As user, USB storage device can only be accessed via Navigation Application for encrypted backup / restore | SR 2.3: The control system shall provide the capability to automatically enforce configurable usage restrictions that include: a) preventing the use of portable and mobile devices; b) requiring context specific authorization; and c) restricting code and data transfer to/from portable and mobile devices.
|  |  |  | SR 3.2 Malicious Code protection | 3 | SL-2 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) | n/a | Users can only execute Stryker approved processes: Bootup ends in Application Manager, no other applications permitted. [DI.SR512026: Bootup is protected against manipulation] Update of OS: [SDS_OS.TR31380: In-Field update] | Users can only execute Stryker approved processes: Bootup ends in Application Manager, no other applications permitted. [DI_SPIDER.SR58060: Bootup is protected against manipulation] Update of OS: [SDS_OS.TR31380: In-Field update] | SR 3.2: The control system shall provide the capability to employ protection mechanisms to prevent, detect, report and mitigate the effects of malicious code or unauthorized software. The control system shall provide the capability to update the protection mechanisms. RE1: The control system shall provide the capability to employ malicious code protection mechanisms at all entry and exit points.
|  |  |  | SR 3.3 Security functionality VERIFICATION | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2) | See above. | See above. | See above. | SR 3.3: The control system shall provide the capability to support verification of the intended operation of security functions and report when anomalies are discovered during FAT, SAT and scheduled maintenance. These security functions shall include all those necessary to support the security requirements specified in this standard.
|  |  |  | SR 5.3 General purpose person-to-person communication restrictions | 5 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | n/a | No installed mail system or browser. Bootup ends in Application Manager, no other applications permitted. [DI.SR512026: Bootup is protected against manipulation] | No installed mail system or browser. Bootup ends in Application Manager, no other applications permitted. [DI_SPIDER.SR58060: Bootup is protected against manipulation] | SR 5.3: The control system shall provide the capability to prevent general purpose person-to-person messages from being received from users or systems external to the control system. RE1: The control system shall provide the capability to prevent both transmission and receipt of general purpose person-to-person messages.
|  |  |  | SR 6.2 Continuous monitoring | 6 | SL-1 | No | Yes | Yes | Yes | n/a | n/a | n/a | n/a | n/a
110 | NAUT - The system shall provide node authentication to ensure that only known devices can connect with each other and share data | Yes |  | SR 1.2 Software PROCESS and device identification and authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.2: The control system shall provide the capability to identify and authenticate all software processes and devices. This capability shall enforce such identification and authentication on all interfaces which provide access to the control system to support least privilege in accordance with applicable security policies and procedures. RE 1: The control system shall provide the capability to uniquely identify and authenticate all software processes and devices.
|  |  |  | SR 1.3 Account Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.3: The control system shall provide the capability to support the management of all accounts by authorized users, including adding, activating, modifying, disabling and removing accounts. RE 1: The control system shall provide the capability to support unified account management.
|  |  |  | SR 1.4 Identifier Management | 1 | SL-3 | Yes | Yes | Yes | Yes | n/a | Managing user accounts with the username as the identifier is sufficient because only a limited number of accounts need to be managed. | Navigation System provides the capability to support the management of all accounts by authorized users. The username is the managed identifier. [SDS_APPM.TR31010: Creation of user] [SDS_APPM.TR31020: User attributes] [SDS_APPM.TR31030: Deletion of users] [SDS_APPM.TR31040: Edit user] | Navigation System provides the capability to support the management of all accounts by authorized users. The username is the managed identifier. [SDS_APPM.TR31010: Creation of user] [SDS_APPM.TR31020: User attributes] [SDS_APPM.TR31030: Deletion of users] [SDS_APPM.TR31040: Edit user] | SR 1.4: The control system shall provide the capability to support the management of identifiers by user, group, role or control system interface.
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | RE1 not implemented. Authenfication via hardware mechanism will reduce usability for the user. | Authenticators are: - User Passwords for physical access. Full password management possible. [SDS_APPM.TR31020: User attributes] [SDS_APPM.TR31040: Edit user] Connected Care provides certificate infrastructure. | Authenticators are: - User Passwords for physical access. Full password management possible. [SDS_APPM.TR31020: User attributes] [SDS_APPM.TR31040: Edit user] Connected Care provides certificate infrastructure. | SR 1.5: The control system shall provide the capability to: h) initialize authenticator content; i) change all default authenticators upon control system installation; j) change/refresh all authenticators; and k) protect all authenticators from unauthorized disclosure and modification when stored and transmitted. RE1: For software process and device users, the control system shall provide the capability to protect the relevant authenticators via hardware mechanisms.
|  |  |  | SR 1.6 Wireless access management | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1) | n/a | Wireless communication use the same user authentication methods as the wired communication via ethernet. [DI.SR512027: WiFi protection] Wireless communication is deactivated by default. Communication can only be setup with an admin account by the hospital IT or service technicians. [SDS_OS.TR31240: WLAN] | Wireless communication use the same user authentication methods as the wired communication via ethernet. [DI_SPIDER.SR58070: WiFi protection] Wireless communication is deactivated by default. Communication can only be setup with an admin account by the hospital IT or service technicians. [SDS_OS.TR31240: WLAN] | SR 1.6: The control system shall provide the capability to identify and authenticate all users (humans, software processes or devices) engaged in wireless communication. RE 1: The control system shall provide the capability to uniquely identify and authenticate all users (humans, software processes or devices) engaged in wireless communication.
|  |  |  | SR 1.8 Public Key Infrastructure (PKI) certificates | 1 | SL-3 | No | Yes | Yes | Yes | n/a | n/a | For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | SR 1.8: Where PKI is utilized, the control system shall provide the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI.
|  |  |  | SR 1.9 Strength of Public Key Authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | n/a | For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | SR 1.9: For control systems utilizing public key authentication, the control system shall provide the capability to: a) validate certificates by checking the validity of the signature of a given certificate; b) validate certificates by constructing a certification path to an accepted CA or in the case of self-signed certificates by deploying leaf certificates to all hosts which communicate with the subject to which the certificate is issued; c) validate certificates by checking a given certificate’s revocation status; d) establish user (human, software process or device) control of the corresponding private key; and e) map the authenticated identity to a user (human, software process or device). RE 1: The control system shall provide the capability to protect the relevant private keys via hardware mechanisms according to commonly accepted security industry practices and recommendations.
|  |  |  | SR 1.10 Authenticator Feedback | 1 | SL-3 | Yes | Yes | Yes | Yes | n/a | n/a | Password is not shown in clear text on the login screen. Virtual keyboard does not highlight keystrokes if using a touchscreen. | Password is not shown in clear text on the login screen. Virtual keyboard does not highlight keystrokes if using a touchscreen. | SR 1.10: The control system shall provide the capability to obscure feedback of authentication information during the authentication process.
|  |  |  | SR 1.11 Unsuccessful Login Attempts | 1 | SL-3 | Yes | Yes | Yes | Yes | n/a | Navigation System is used in protected OR environment. Feature is not required. | n/a | n/a | SR 1.11: The control system shall provide the capability to enforce a limit of a configurable number of consecutive invalid access attempts by any user (human, software process or device) during a configurable time period. The control system shall provide the capability to deny access for a specified period of time or until unlocked by an administrator when this limit has been exceeded. For system accounts on behalf of which critical services or servers are run, the control system shall provide the capability to disallow interactive logons.
|  |  |  | SR 1.13 Access via Untrusted Networks | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1) | n/a | Firewall settings review [SDS_OS.TR31440: Ethernet port protection] [SDS_OS.TR31441: Ethernet port protection] | Firewall settings review [SDS_OS.TR31440: Ethernet port protection] [SDS_OS.TR31441: Ethernet port protection] | SR 1.13: The control system shall provide the capability to monitor and control all methods of access to the control system via untrusted networks. RE 1: The control system shall provide the capability to deny access requests via untrusted networks unless approved by an assigned role.
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes | n/a | n/a | Encryption of patient data at rest and patient data backups. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] | Encryption of patient data at rest and patient data backups. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] | SR 4.3: If cryptography is required, the control system shall use cryptographic algorithms, key sizes and mechanisms for key establishment and management according to commonly accepted security industry practices and recommendations.
120 | PAUT - The system shall provide the ability to authenticate users | Yes |  | SR 1.1 Human user identification and authentication | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3) | Stryker service technicians share a common admin account and therefore cannot be uniquly identified. This is required for maintenance of the Navigation Systems. | User can be uniquely identified, if the hospital IT has created one account for each user. Hospital IT can also use LDAP integtration. [SDS_APPM.TR32040: LDAP Configuration] | User can be uniquely identified, if the hospital IT has created one account for each user. Hospital IT can also use LDAP integtration. [SDS_APPM.TR32040: LDAP Configuration] | SR 1.1: The control system shall provide the capability to identify and authenticate all human users. This capability shall enforce such identification and authentication on all interfaces which provide human user access to the control system to support segregation of duties and least privilege in accordance with applicable security policies and procedures.
|  |  |  | SR 1.3 Account Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.3: The control system shall provide the capability to support the management of all accounts by authorized users, including adding, activating, modifying, disabling and removing accounts. RE 1: The control system shall provide the capability to support unified account management.
|  |  |  | SR 1.4 Identifier Management | 1 | SL-3 | Yes | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 1.4: The control system shall provide the capability to support the management of identifiers by user, group, role or control system interface.
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.5: The control system shall provide the capability to: h) initialize authenticator content; i) change all default authenticators upon control system installation; j) change/refresh all authenticators; and k) protect all authenticators from unauthorized disclosure and modification when stored and transmitted. RE1: For software process and device users, the control system shall provide the capability to protect the relevant authenticators via hardware mechanisms.
|  |  |  | SR 1.6 Wireless access management | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.6: The control system shall provide the capability to identify and authenticate all users (humans, software processes or devices) engaged in wireless communication. RE 1: The control system shall provide the capability to uniquely identify and authenticate all users (humans, software processes or devices) engaged in wireless communication.
|  |  |  | SR 1.7 Strength of Password-Based Authentication | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) (2) | SL-3 (1), SL-4 (1) (2) | Password policies have to be enforced manually by the hospital IT. Alternatively the LDAP integration can be used. | The Application Manager allows the Hospital's IT admin to manually set passwords according to the hospital's password policy. Hospital IT admin can also use LDAP integtration. [SDS_APPM.TR32040: LDAP Configuration] | The Application Manager allows the Hospital's IT admin to manually set passwords according to the hospital's password policy. Hospital IT admin can also use LDAP integtration. [SDS_APPM.TR32040: LDAP Configuration] | SR 1.7: For control systems utilizing password-based authentication, the control system shall provide the capability to enforce configurable password strength based on minimum length and variety of character types. RE1: The control system shall provide the capability to prevent any given human user account from reusing a password for a configurable number of generations. In addition, the control system shall provide the capability to enforce password minimum and maximum lifetime restrictions for human users. These capabilities shall conform with commonly accepted security industry practices.
|  |  |  | SR 1.8 Public Key Infrastructure (PKI) certificates | 1 | SL-3 | No | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 1.8: Where PKI is utilized, the control system shall provide the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI.
|  |  |  | SR 1.9 Strength of Public Key Authentication | 1 | SL-3 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.9: For control systems utilizing public key authentication, the control system shall provide the capability to: a) validate certificates by checking the validity of the signature of a given certificate; b) validate certificates by constructing a certification path to an accepted CA or in the case of self-signed certificates by deploying leaf certificates to all hosts which communicate with the subject to which the certificate is issued; c) validate certificates by checking a given certificate’s revocation status; d) establish user (human, software process or device) control of the corresponding private key; and e) map the authenticated identity to a user (human, software process or device). RE 1: The control system shall provide the capability to protect the relevant private keys via hardware mechanisms according to commonly accepted security industry practices and recommendations.
|  |  |  | SR 1.10 Authenticator Feedback | 1 | SL-3 | Yes | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 1.10: The control system shall provide the capability to obscure feedback of authentication information during the authentication process.
|  |  |  | SR 1.11 Unsuccessful Login Attempts | 1 | SL-3 | Yes | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 1.11: The control system shall provide the capability to enforce a limit of a configurable number of consecutive invalid access attempts by any user (human, software process or device) during a configurable time period. The control system shall provide the capability to deny access for a specified period of time or until unlocked by an administrator when this limit has been exceeded. For system accounts on behalf of which critical services or servers are run, the control system shall provide the capability to disallow interactive logons.
|  |  |  | SR 1.13 Access via Untrusted Networks | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) | Yes + (1) | SL-2 (1), SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.13: The control system shall provide the capability to monitor and control all methods of access to the control system via untrusted networks. RE 1: The control system shall provide the capability to deny access requests via untrusted networks unless approved by an assigned role.
|  |  |  | SR 2.3 Use Control for portable and mobile devices | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 2.3: The control system shall provide the capability to automatically enforce configurable usage restrictions that include: a) preventing the use of portable and mobile devices; b) requiring context specific authorization; and c) restricting code and data transfer to/from portable and mobile devices.
|  |  |  | SR 2.8 Auditable events | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 2.8 The control system shall provide the capability to generate audit records relevant to security for the following categories: access control, request errors, operating system events, control system events, backup and restore events, configuration changes, potential reconnaissance activity and audit log events. Individual audit records shall include the timestamp, source (originating device, software process or human user account), category, type, event ID and event result.
|  |  |  | SR 2.11 Timestamps | 2 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) (2) | n/a | Timestamps are included in operating system, Application Manager and Navigation Application log files. | Timestamps are included in operating system, Application Manager and Navigation Application log files. | SR 2.11: The control system shall provide timestamps for use in audit record generation.
|  |  |  | SR 2.12 Non-repudiation | 2 | SL-2 | No | No | Yes | Yes | SL-4 (1) | n/a | n/a | n/a | SR 2.12: The control system shall provide the capability to determine whether a given human user took a particular action.
|  |  |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2) | n/a | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. Use of Unique Identifier instead of Patient Name to protect confidentiality. [SRS_OPK.F2260: Patient folder] | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. Use of Unique Identifier instead of Patient Name to protect confidentiality. [SRS_OPK.F2260: Patient folder] | SR 4.1:The control system shall provide the capability to protect the confidentiality of information for which explicit read authorization is supported, whether at rest or in transit. RE1: The control system shall provide the capability to protect the confidentiality of information at rest and remote access sessions traversing an untrusted network.
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 4.3: If cryptography is required, the control system shall use cryptographic algorithms, key sizes and mechanisms for key establishment and management according to commonly accepted security industry practices and recommendations.
|  |  |  | SR 6.2 Continuous monitoring | 6 | SL-1 | No | Yes | Yes | Yes | n/a | n/a | n/a | n/a
|  |  |  | SR 1.12 – System use notification | 1 | SL-3 | Yes | Yes | Yes | Yes | n/a | Display of a disclaimer/message is not required. | n/a | n/a | SR 1.12: The control system shall provide the capability to display a system use notification message before authenticating. The system use notification message shall be configurable by authorized personnel.
130 | PLOK - The system components that maintain private data shall be physically secure | Yes |  | SR 1.1 Human user identification and authentication | 1 | SL-3 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3) | See above. | See above. | See above. | SR 1.1: The control system shall provide the capability to identify and authenticate all human users. This capability shall enforce such identification and authentication on all interfaces which provide human user access to the control system to support segregation of duties and least privilege in accordance with applicable security policies and procedures.
|  |  |  | SR 1.3 Account Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.3: The control system shall provide the capability to support the management of all accounts by authorized users, including adding, activating, modifying, disabling and removing accounts. RE 1: The control system shall provide the capability to support unified account management.
|  |  |  | SR 1.5 Authenticator Management | 1 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 1.5: The control system shall provide the capability to: h) initialize authenticator content; i) change all default authenticators upon control system installation; j) change/refresh all authenticators; and k) protect all authenticators from unauthorized disclosure and modification when stored and transmitted. RE1: For software process and device users, the control system shall provide the capability to protect the relevant authenticators via hardware mechanisms.
|  |  |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2) | See above. | See above. | See above. | SR 4.1:The control system shall provide the capability to protect the confidentiality of information for which explicit read authorization is supported, whether at rest or in transit. RE1: The control system shall provide the capability to protect the confidentiality of information at rest and remote access sessions traversing an untrusted network.
|  |  |  | SR 7.7 Least functionality | 7 | SL-1 | Yes | Yes | Yes | Yes | n/a | n/a | User has restricted access rights. User can only start application. I/O ports are secured. [DI.SR512021: I/O port usage is secured against unauthorised access] [SDS_OS.TR31200: USB ports with restricted access] | User has restricted access rights. User can only start application. I/O ports are secured. [DI_SPIDER.SR58010: I/O port usage is secured against unauthorised access] [SDS_OS.TR31200: USB ports with restricted access] | SR 7.7: The control system shall provide the capability to specifically prohibit and/or restrict the use of unnecessary functions, ports, protocols and/or services.
140 | RDMP - The system software development plan shall address security support of 3rd party components throughout system life cycle | Yes |  | SR 4.2 Information persistence | 4 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 4.2: The control system shall provide the capability to purge all information for which explicit read authorization is supported from components to be released from active service and/or decommissioned. RE1: The control system shall provide the capability to prevent unauthorized and unintended information transfer via volatile shared memory resources.
150 | SAHD - System shall provide hardening features to guard against cyber attacks and malware | Yes |  | SR 2.1 Authorization Enforcement | 2 | SL-2 | Yes | Yes + (1) (2) | Yes + (1) (2) (3) | Yes + (1) (2) (3) (4) | SL-2 (1) (2), SL-3 (1) (2) (3), SL-4 (1) (2) (3) (4) | See above. | See above. | See above. | SR 2.1 RE 1 : On all interfaces, the control system shall provide the capability to enforce authorizations assigned to all users (humans, software processes and devices) for controlling use of the control system to support segregation of duties and least privilege.
|  |  |  | SR 2.2 Wireless use control | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | n/a | WiFi can be deactivated by the hospital IT admin or service technician. [DI.SR512027: WiFi protection] | WiFi can be deactivated by the hospital IT admin or service technician. [DI_SPIDER.SR58070: WiFi protection] | SR 2.2: The control system shall provide the capability to authorize, monitor and enforce usage restrictions for wireless connectivity to the control system according to commonly accepted security industry practices.
|  |  |  | SR 2.3 Use Control for portable and mobile devices | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 2.3: The control system shall provide the capability to automatically enforce configurable usage restrictions that include: a) preventing the use of portable and mobile devices; b) requiring context specific authorization; and c) restricting code and data transfer to/from portable and mobile devices.
|  |  |  | SR 3.4 Software and information integrity | 3 | SL-2 | No | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 3.4: The control system shall provide the capability to detect, record, report and protect against unauthorized changes to software and information at rest.
|  |  |  | SR 5.1 Network Segmentation | 5 | SL-3 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3) | n/a | Segregated FP8000 ethernet network and Stryker device network from rest of the hospital network. Seperate ethernet card for FP8000 network. [DI.SR512040: Segregated FP8000 ethernet network.] | Segregated FP8000 ethernet network and Stryker device network from rest of the hospital network. Seperate ethernet card for FP8000 network. [DI_SPIDER.SR43010: FP8000 camera integration] | SR 5.1: The control system shall provide the capability to logically segment control system networks from non-control system networks and to logically segment critical control system networks from other control system networks. RE 1: The control system shall provide the capability to physically segment control system networks from non-control system networks and to physically segment critical control system networks from non-critical control system networks. RE 2: The control system shall have the capability to provide network services to control system networks, critical or otherwise, without a connection to non-control system networks.
|  |  |  | SR 5.2 Zone boundary protection | 5 | SL-3 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) (3) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) (3) | n/a | Segregated FP8000 ethernet network and Stryker device network from rest of the hospital network. Seperate ethernet card for FP8000 network. [DI.SR512040: Segregated FP8000 ethernet network.] Firewall settings [SDS_OS.TR31440: Ethernet port protection] [SDS_OS.TR31441: Ethernet port protection] | Segregated FP8000 ethernet network and Stryker device network from rest of the hospital network. Seperate ethernet card for FP8000 network. [DI_SPIDER.SR58090: Segregated FP8000 ethernet network.] Firewall settings [SDS_OS.TR31440: Ethernet port protection] [SDS_OS.TR31441: Ethernet port protection] | SR 5.2: The control system shall provide the capability to monitor and control communications at zone boundaries to enforce the compartmentalization defined in the risk-based zones and conduits model. RE1: The control system shall provide the capability to deny network traffic by default and allow network traffic by exception (also termed deny all, permit by exception). RE2: The control system shall provide the capability to prevent any communication through the control system boundary (also termed island mode).
|  |  |  | SR 5.3 General purpose person-to-person communication restrictions | 5 | SL-3 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 5.3: The control system shall provide the capability to prevent general purpose person-to-person messages from being received from users or systems external to the control system. RE1: The control system shall provide the capability to prevent both transmission and receipt of general purpose person-to-person messages.
|  |  |  | SR 2.4 – Mobile code | 2 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | n/a | Hardening of the OS. Usage of mobile code technologies like Java Script, ActiveX, Flash is restricted. Mobile code is prevented from execution: Users can only execute Stryker approved processes: Bootup ends in Application Manager, no other applications permitted. [DI.SR512026: Bootup is protected against manipulation] | Hardening of the OS. Usage of mobile code technologies like Java Script, ActiveX, Flash is restricted. Mobile code is prevented from execution: Users can only execute Stryker approved processes: Bootup ends in Application Manager, no other applications permitted. [DI_SPIDER.SR58060: Bootup is protected against manipulation] | SR 2.4: The control system shall provide the capability to enforce usage restrictions for mobile code technologies based on the potential to cause damage to the control system that include: a) preventing the execution of mobile code; b) requiring proper authentication and authorization for origin of the code; c) restricting mobile code transfer to/from the control system; and d) monitoring the use of mobile code.
|  |  |  | SR 5.4 Application partitioning | 5 | SL-3 | Yes | Yes | Yes | Yes | n/a | Application Manager and Navigation Application are seperated application. No further seperation required. | n/a | n/a | SR 5.4: The control system shall provide the capability to support partitioning of data, applications and services based on criticality to facilitate implementing a zoning model.
|  |  |  | SR 7.1 Denial of service protection | 7 | SL-1 | Yes | Yes + (1) | Yes + (1) (2) | Yes + (1) (2) | SL-2 (1), SL-3 (1) (2), SL-4 (1) (2) | Denial of Service is considered a low impact/severity risk. Protection against denial of service attacks are very complex. | It is recommended to the user, that surgeries should be performed without hospital network connection. Nevertheless, segregated FP8000 connection segregates the Camera from DoS attacks via hospital network. Degradation of camera performance in such an event is unlikely. [DI.SR512040: Segregated FP8000 ethernet network.] | It is recommended to the user, that surgeries should be performed without hospital network connection. Nevertheless, segregated FP8000 connection segregates the Camera from DoS attacks via hospital network. Degradation of camera performance in such an event is unlikely. [DI_SPIDER.SR58090: Segregated FP8000 ethernet network.] | SR 7.1: The control system shall provide the capability to operate in a degraded mode during a DoS event.
|  |  |  | SR 7.2 Resource management | 7 | SL-1 | Yes | Yes | Yes | Yes | n/a | n/a | Interference of security functions with essential performance requirements is regularly verified during system test. [SRS_OPK.F1700: Essential Performance] | Interference of security functions with essential performance requirements is regularly verified during system test. [SRS_OPK.F1700: Essential Performance] | SR 7.2: The control system shall provide the capability to limit the use of resources by security functions to prevent resource exhaustion.
|  |  |  | SR 7.5 Emergency power | 7 | SL-1 | Yes | Yes | Yes | Yes | n/a | In case of power breakdown it is considered acceptable that the Navigation System is no longer available. It has no life sustaining functions and the surgery would be interrupted/canceled anyways. UPS is not meant to serve as a backup in case of power outage. Navigation Applications are build to restore files critical for the surgery such as registration data, so that the surgery can be quickly resumed after a short term power outage. | n/a | n/a | SR 7.5: The control system shall provide the capability to switch to and from an emergency power supply without affecting the existing security state or a documented degraded mode.
|  |  |  | SR 7.7 Least functionality | 7 | SL-1 | Yes | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 7.7: The control system shall provide the capability to specifically prohibit and/or restrict the use of unnecessary functions, ports, protocols and/or services.
160 | SGUD - The system service/user manuals shall provide security guidance for operators and administrators | Yes |  | No applicable SECURITY CONTROLS | 0 | #N/A | N/A | N/A | N/A | N/A | n/a | n/a | n/a | n/a
170 | STCF - The system shall establish technical controls to mitigate the potential for compromise to the integrity and confidentiality of health data stored on product or removable media | No |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2) | See above. | See above. | See above. | SR 4.1:The control system shall provide the capability to protect the confidentiality of information for which explicit read authorization is supported, whether at rest or in transit. RE1: The control system shall provide the capability to protect the confidentiality of information at rest and remote access sessions traversing an untrusted network.
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 4.3: If cryptography is required, the control system shall use cryptographic algorithms, key sizes and mechanisms for key establishment and management according to commonly accepted security industry practices and recommendations.
180 | TXCF - The system shall ensure confidentiality of transmitted Health Data | Yes |  | SR 1.8 Public Key Infrastructure (PKI) certificates | 1 | SL-3 | No | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 1.8: Where PKI is utilized, the control system shall provide the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI.
|  |  |  | SR 4.1 Information confidentiality | 4 | SL-2 | Yes | Yes + (1) | Yes + (1) | Yes + (1) (2) | SL-2 (1), SL-3 (1), SL-4 (1) (2) | See above. | See above. | See above. | SR 4.1:The control system shall provide the capability to protect the confidentiality of information for which explicit read authorization is supported, whether at rest or in transit. RE1: The control system shall provide the capability to protect the confidentiality of information at rest and remote access sessions traversing an untrusted network.
|  |  |  | SR 4.3 Use of cryptography | 4 | SL-2 | Yes | Yes | Yes | Yes | n/a | See above. | See above. | See above. | SR 4.3: If cryptography is required, the control system shall use cryptographic algorithms, key sizes and mechanisms for key establishment and management according to commonly accepted security industry practices and recommendations.
190 | TXIG - The system shall ensure integrity of transmitted Health Data | Yes |  | SR 3.1 Communication Integrity | 3 | SL-2 | Yes | Yes | Yes + (1) | Yes + (1) | SL-3 (1), SL-4 (1) | See above. | See above. | See above. | SR 3.1: The control system shall provide the capability to protect the integrity of transmitted information.
|  |  |  | SR 3.8 Session Integrity | 3 | SL-2 | No | Yes | Yes + (1) (2) | Yes + (1) (2) (3) | SL-3 (1) (2), SL-4 (1) (2) (3) | n/a | Connection Care remote session is secured. No web interfaces or other session-based communications existent. | Connection Care remote session is secured. No web interfaces or other session-based communications existent. | SR 3.8: The control system shall provide the capability to protect the integrity of sessions. The control system shall reject any usage of invalid session IDs.
200 | PII | Yes |  | AP-1 Authority to Collect | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AP-2 Purpose Specification | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-1 Governance and Privacy Program | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-2 Privacy Impact and Risk Assessment | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-3 Privacy Requirements for Contractors and Service Providers | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-4 Privacy Monitoring and Auditing | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-5 Privacy Awareness and Training | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-6 Privacy Reporting | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-7 Privacy-enhanced system design and development | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AR-8 Accounting of Disclosures | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | DI-1 Data Quality | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | DI-2 Data Integrity and Data Integrity Board | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | DM-1 Minimization of personally identifiable information | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | Use of Unique Identifier instead of Patient Name to protect confidentiality. [SRS_OPK.F2260: Patient folder] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. | Use of Unique Identifier instead of Patient Name to protect confidentiality. [SRS_OPK.F2260: Patient folder] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files.
|  |  |  | DM-2 Data retention and disposal | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | DM-3 Minimization of PII Used in Testing, Training, and IP | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | IP-1 Consent | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | IP-2 Individual Access | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | IP-3 Redress | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | IP-4 Complaint Management | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | SE-1 Inventory of Personally Idntifiable Information | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | SE-2 Privacy Incident Response | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | TR-1 Privacy Notice | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | TR-2 System of Records Notices and Privacy Act | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | TR-3 Dissemination of Privacy Program Information | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | UL-1 Internal Use | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | UL-2 Information Sharing with Third Parties | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
210 | PHI | Yes |  | 164.312(a)(2)(i) Unique User Identification (R): Assign a unique name and/or number for identifying and tracking user identity. | [6 empty] | n/a | n/a | User can be uniquely identified, if the hospital IT has created one account for each user. Hospital IT can also use LDAP integtration. [SDS_APPM.TR32040: LDAP Configuration] | User can be uniquely identified, if the hospital IT has created one account for each user. Hospital IT can also use LDAP integtration. [SDS_APPM.TR32040: LDAP Configuration]
|  |  |  | AC-2 Account management | [6 empty] | n/a | n/a | [DI.SR61010: User management] | [SDS_APPM.TR31010: Creation of user] [SDS_APPM.TR31020: User attributes] [SDS_APPM.TR31030: Deletion of users] [SDS_APPM.TR31040: Edit user] | The organization: AC2-c. Establishes conditions for group and role membership; AC2-d. Specifies authorized users of the information system, group and role membership, and access authorizations (i.e., privileges) and other attributes (as required) for each account; AC2-f. Creates, enables, modifies, disables, and removes information system accounts in accordance with [Assignment : organization-defined procedures or conditions];
|  |  |  | AC-3: Access enforcement | [6 empty] | n/a | n/a | [SDS_APPM.TR32040: LDAP Configuration] | [SDS_APPM.TR32040: LDAP Configuration] | The information system enforces approved authorizations for logical access to information and system resources in accordance with applicable access control policies.
|  |  |  | IA-2 Identification and authentication (organizational users) | [6 empty] | n/a | n/a | [DI.SR61020: Network configuration] | [SDS_APPM.TR32030: Network Configuration] | The information system uniquely identifies and authenticates organizational users
|  |  |  | IA-3 Device Identification and Authentication | [6 empty] | n/a | Access of the Navigation System to the hospital network must be configured by hospital IT admin. Access limitations by means of device identification (eg. MAC Address of LAN Interface, IP Address) must be managed by the network. | n/a | n/a | The information system uniquely identifies and authenticates [Assignment : organization-defined specific and/or types of devices] before establishing a [Selection (one or more): local; remote; network] connection.
|  |  |  | IA-4 Identifier management | [6 empty] | n/a | Identifiers such as user account names and Passwords are managed by the hospital | n/a | n/a | The organization manages information system identifiers
|  |  |  | 164.312(a)(2)(ii) Emergency Access Procedure (R): Establish (and implement as needed) procedures for obtaining necessary electronic protected health information during an emergency. | [6 empty] | n/a | No emergency relevant data like e.g. blood type, medication is saved on the Navigation System. Patient scans can be accessed directly on the PACS system or the scanner. | n/a | n/a
|  |  |  | AC-2 Account management | [6 empty] | n/a | n/a | See AC-2 above. | See AC-2 above.
|  |  |  | AC-3: Access enforcement | [6 empty] | n/a | n/a | See AC-3 above. | See AC-3 above.
|  |  |  | CP-2 Contingency Plan | [6 empty] | n/a | All controls pertaining a contingency plan are Hospital/Stryker corp. Policy | n/a | n/a | The organization develops a contingency plan for the information system
|  |  |  | 164.312(a)(2)(iii) Automatic Logoff (A): Implement electronic procedures that terminate an electronic session after a predetermined time of inactivity. | [6 empty] | n/a | Session lock not applicable to user accounts. Navigation System needs to be available during surgery to not interupt the surgical workflow even after a longer time of inactivity. | [SDS_APPM.TR37010: Session Lock] | [SDS_APPM.TR37010: Session Lock]
|  |  |  | AC-11 Session lock | [6 empty] | n/a | Session lock not applicable to user accounts. Navigation System needs to be available during surgery to not interupt the surgical workflow even after a longer time of inactivity. | [SDS_APPM.TR37010: Session Lock] | [SDS_APPM.TR37010: Session Lock]
|  |  |  | AC-12 Session Termination | [6 empty] | n/a | Session termination not applicable to user accounts. Navigation System needs to be available during surgery to not interupt the surgical workflow even after a longer time of inactivity. | n/a | n/a
|  |  |  | 164.312(a)(2)(iv) Encryption and Decryption (A): Implement a mechanism to encrypt and decrypt electronic protected health information. | [6 empty] | n/a | n/a | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files.
|  |  |  | AC-3: Access enforcement | [6 empty] | n/a | n/a | See AC-3 above. | See AC-3 above.
|  |  |  | SC-13 Cryptographic protection | [6 empty] | n/a | n/a | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. | The information system implements [Assignment : organization-defined cryptographic uses and type of cryptography required for each use] in accordance with applicable federal laws, Executive Orders, directives, policies, regulations, and standards.
|  |  |  | 164.312(b) Audit Controls: Implement hardware, software, and/or procedural mechanisms that record and examine activity in information systems that contain or use electronic protected health information. | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AU-1 Audit and accountability policy and procedures | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | AU-2 Audit events | [6 empty] | n/a | No product requirement - Stryker corp. Policy | n/a | n/a
|  |  |  | AU-3 Content of audit records | [6 empty] | n/a | n/a | - Access control: Login of different users is logged using the Appman. - Operating system events: Yes - Control system events: Log4Cpp application events (Application Manager and Navigation Application). - Backup events: Logged by Navigation Application. - Restore events: Logged by Navigation Application. - Configuration changes: Not implemened for Navigation Application config files but for other parts of the system. [DI.SR512023: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] [SADD_OS] Chapter 9.15 Logging and Tracing | - Access control: Login of different users is logged using the Appman. - Operating system events: Yes - Control system events: Log4Cpp application events (Application Manager and Navigation Application). - Backup events: Logged by Navigation Application. - Restore events: Logged by Navigation Application. - Configuration changes: Not implemened for Navigation Application config files but for other parts of the system. [DI_SPIDER.SR58030: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] [SADD_OS] Chapter 9.15 Logging and Tracing | The information system generates audit records containing information that establishes what type of event occurred, when the event occurred, where the event occurred, the source of the event, the outcome of the event, and the identity of any individuals or subjects associated with the event.
|  |  |  | AU-4 Audit storage capacity | [6 empty] | n/a | n/a | System has enough disk space to store log files for a longer period of time. No capacity requirement needed. | System has enough disk space to store log files for a longer period of time. No capacity requirement needed.
|  |  |  | AU-6 Audit review, analysis and reporting | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a | The organization allocates audit record storage capacity in accordance with [Assignment : organization-defined audit record storage requirements].
|  |  |  | AU-7 Audit reduction and report generation | [6 empty] | n/a | Not required for Navigation Systems. Log files can be gathered, evaluated and summarized individually on request. | n/a | n/a | The information system provides an audit reduction and report generation capability that: a. Supports on-demand audit review, analysis, and reporting requirements and after-the-fact investigations of security incidents; and b. Does not alter the original content or time ordering of audit records.
|  |  |  | 164.312(c)(1) Integrity: Implement policies and procedures to protect electronic protected health information from improper alteration or destruction. | [6 empty] | n/a | n/a | Restrictive user permissions [DI.SR512023: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] | Restrictive user permissions [DI_SPIDER.SR58030: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point]
|  |  |  | CP-9 Information system backup | [6 empty] | n/a | Operating system backups are not necessary. The operating system can be installed from scratch if needed. | System recovery is done by Sales Rep with OS and application Installation media. PHI is not stored on the System partition. Patient data can be restored if regular backups are created. In worst case the loss of patient data is not considered being a critical event. [SRS_OPK.F2270: Patient Backup/export] [SRS_OPK.F2280: Patient Restore/import] Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. Predetermined state can be maintained by factory reset. [SDS_APPM.TR35030: Reset to factory default] | System recovery is done by Sales Rep with OS and application Installation media. PHI is not stored on the System partition. Patient data can be restored if regular backups are created. In worst case the loss of patient data is not considered being a critical event. [SRS_OPK.F2270: Patient Backup/export] [SRS_OPK.F2280: Patient Restore/import] Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. Predetermined state can be maintained by factory reset. [SDS_APPM.TR35030: Reset to factory default]
|  |  |  | MP-2 Media access | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | MP-5 Media Transport | [6 empty] | n/a | No product requirement - Hospital Policy | n/a | n/a
|  |  |  | SC-8 Transmission confidentiality and integrity | [6 empty] | n/a | n/a | Communication to ConnectedCare server is protected by encryption. For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | Communication to ConnectedCare server is protected by encryption. For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | The information system protects the [Selection (one or more): confidentiality; integrity] of transmitted information.
|  |  |  | SI-1 System and information integrity policy and procedures | [6 empty] | n/a | No product requirement - Hospital/Stryker corp. Policy | n/a | n/a
|  |  |  | SI-7 Software and information integrity | [6 empty] | n/a | n/a | Critical software configuration files and binaries cannot be accessed or modified by the user due to restricted permissions. [DI.SR512023: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] | Critical software configuration files and binaries cannot be accessed or modified by the user due to restricted permissions. [DI_SPIDER.SR58030: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] | The organization employs integrity verification tools to detect unauthorized changes to [Assignment : organization-defined software, firmware, and information].
|  |  |  | 164.312(c)(2) Mechanism to Authenticate Electronic Protected Health Information (A): Implement electronic mechanisms to corroborate that electronic protected health information has not been altered or destroyed in an unauthorized manner. | [6 empty] | n/a | n/a | Restrictive user permissions [DI.SR512023: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point] | Restrictive user permissions [DI_SPIDER.SR58030: Principle of least privileges] [SDS_OS.TR31480: Application Manager is single entry point]
|  |  |  | SC-8 Transmission confidentiality and integrity | [6 empty] | n/a | n/a | See SC-8 above. | See SC-8 above.
|  |  |  | SI-7 Software and information integrity | [6 empty] | n/a | n/a | See SI-7 above. | See SI-7 above.
|  |  |  | 164.312(d) Person or Entity Authentication: Implement procedures to verify that a person or entity seeking access to electronic protected health information is the one claimed. | [6 empty] | n/a | n/a | [SDS_APPM.TR31020: User attributes] | [SDS_APPM.TR31020: User attributes]
|  |  |  | IA-2 Identification and authentication (organizational users) | [6 empty] | n/a | n/a | See IA-2 above. | See IA-2 above.
|  |  |  | IA-3 Device Identification and Authentication | [6 empty] | n/a | n/a | See IA-3 above. | See IA-3 above.
|  |  |  | IA-4 Identifier management | [6 empty] | n/a | n/a | See IA-4 above. | See IA-4 above.
|  |  |  | 164.312 (e)(1) Transmission Security: Implement technical security measures to guard against unauthorized access to electronic protected health information that is being transmitted over an electronic communications network. | [6 empty] | n/a | No external input / data is imported for intended use of the navigation system. | Communication to ConnectedCare server is protected by encryption. For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | Communication to ConnectedCare server is protected by encryption. For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI.
|  |  |  | SC-9 Transmission Confidentiality | [6 empty] | n/a | n/a | N/A - NIST 800-53 Rev. 4: Withdrawn: Incorporated into SC-8 | N/A - NIST 800-53 Rev. 4: Withdrawn: Incorporated into SC-8
|  |  |  | 164.312(e)(2)(i) Integrity Controls (A): Implement security measures to ensure that electronically transmitted electronic protected health information is not improperly modified without detection until disposed of. | [6 empty] | n/a | No external input / data is imported for intended use of the navigation system. | Communication to ConnectedCare server is protected by encryption. For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI. | Communication to ConnectedCare server is protected by encryption. For ConnectedCare the Navigation System provides the capability to operate a PKI according to commonly accepted best practices or obtain public key certificates from an existing PKI.
|  |  |  | SC-8 Transmission confidentiality and integrity | [6 empty] | n/a | n/a | See SC-8 above. | See SC-8 above. | The information system protects the [Selection (one or more): confidentiality; integrity] of transmitted information.
|  |  |  | SI-7 Software and information integrity | [6 empty] | n/a | n/a | See SI-7 above. | See SI-7 above. | The organization employs integrity verification tools to detect unauthorized changes to [Assignment : organization-defined software, firmware, and information].
|  |  |  | 164.312(e)(2)(ii) Encryption (A): Implement a mechanism to encrypt electronic protected health information whenever deemed appropriate. | [6 empty] | n/a | n/a | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files.
|  |  |  | SC-9 Transmission Confidentiality | [6 empty] | n/a | n/a | N/A - NIST 800-53 Rev. 4: Withdrawn: Incorporated into SC-8 | N/A - NIST 800-53 Rev. 4: Withdrawn: Incorporated into SC-8
|  |  |  | SC-12 Cryptographic key establishment and management | [6 empty] | n/a | No product requirement - Stryker corp. Policy | n/a | n/a | The organization establishes and manages cryptographic keys for required cryptography employed within the information system in accordance with [Assignment : organization-defined requirements for key generation, distribution, storage, access, and destruction].
|  |  |  | SC-13 Cryptographic protection | [6 empty] | n/a | n/a | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. | Patient data at rest and patient data backups are stored in an encrypted manner. [SRS_OPK.F1230: Encryption of patient data at rest] [SRS_OPK.F1240: Encryption of patient data backup] [SRS_OPK.F2230: Log file] Doesn't contain pateint name in log files. | The information system implements [Assignment : organization-defined cryptographic uses and type of cryptography required for each use] in accordance with applicable federal laws, Executive Orders, directives, policies, regulations, and standards.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.
Note on P162: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Seems this requirement is for previous row - 161

## System Security Level
REFERENCE FOR SYSTEM SECURITY LEVEL
NIST SP 800-53:
SC information type = {(confidentiality, impact), (integrity, impact), (availability, impact)},
The potential impact is LOW if—
− The loss of confidentiality, integrity, or availability could be expected to have a limited adverse effect on organizational operations, organizational assets, or individuals.2
AMPLIFICATION: A limited adverse effect means that, for example, the loss of confidentiality, integrity, or availability might: (i) cause a degradation in mission capability to an extent and duration that the organization is able to perform its primary functions, but the effectiveness of the functions is noticeably reduced; (ii) result in minor damage to organizational assets; (iii) result in minor financial loss; or (iv) result in minor harm to individuals.
The potential impact is MODERATE if—
− The loss of confidentiality, integrity, or availability could be expected to have a serious adverse effect on organizational operations, organizational assets, or individuals.
AMPLIFICATION: A serious adverse effect means that, for example, the loss of confidentiality, integrity, or availability might: (i) cause a significant degradation in mission capability to an extent and duration that the organization is able to perform its primary functions, but the effectiveness of the functions is significantly reduced; (ii) result in significant damage to organizational assets; (iii) result in significant financial loss; or (iv) result in significant harm to individuals that does not involve loss of life or serious life threatening injuries.
The potential impact is HIGH if—
− The loss of confidentiality, integrity, or availability could be expected to have a severe or catastrophic adverse effect on organizational operations, organizational assets, or individuals.
AMPLIFICATION: A severe or catastrophic adverse effect means that, for example, the loss of confidentiality, integrity, or availability might: (i) cause a severe degradation in or loss of mission capability to an extent and duration that the organization is not able to perform one or more of its primary functions; (ii) result in major damage to organizational assets; (iii) result in major financial loss; or (iv) result in severe or catastrophic harm to individuals involving loss of life or serious life threatening injuries.
A low-impact system is an information system in which all three of the security objectives are low.
A moderate-impact system is an information system in which at least one of the security objectives is moderate and no security objective is greater than moderate.
A high-impact system is an information system in which at least one security objective is high.
IEC 62443-3-3
Security Level Vector: SL-?([FR,]domain) = { IAC UC SI DC RDF TRE RA } Example: SL-C ("Product") = {3 2 2 1 1 3 2}
SL= Security Level
FR = Foundational Requirement
IAC = Identification and authentication control
UC = Use control
SI = System integrity
DC = Data confidentiality
RDF = Restricted data flow
TRE = Timely response to events
RA = Resource availability
RE = Requirement Enhancement
SL-? = (Required) The SL type. SL-C = Capabilities SL [FR,] = (Optional) Field indicating the FR that the SL value applies. The FRs are written out in abbreviated form instead of numerical form to aid in readability. domain = (Required) The applicable domain that the SL applies. Domains can refer to zones, control systems, subsystems or components.In this particular standard, all requirements refer to a control system, so the domain term is not used as it would be by other documents in the IEC 62443 series.
General description and distinctions between Security Levels
SL 0: No specific requirements or security protection necessary SL 0 has multiple meanings depending on the situation in which it is applied. In defining SL-C it would mean that the component or system fails to meet some of the SL 1 requirements for that particular FR. This would most likely be for components or systems that would be part of a larger zone where other components or systems would provide compensating countermeasures. In defining SL-T for a particular zone it means that the asset owner has determined that the results of their risk analysis indicate that less than the full SL 1 specific requirements are necessary for that particular FR on that component or system. This would more likely happen for individual components within a system or zone that do not contribute in any way to the FR-specific requirements. In defining SL-A it would mean that the particular zone fails to meet some of the SL 1 requirements for that particular FR.
SL 1: Protection against casual or coincidental violation Casual or coincidental violations of security are usually through the lax application of security policies. These can be caused by well-meaning employees just as easily as they can be by an outsider threat. Many of these violations will be security program related and will be handled by enforcing policies and procedures.A simple example would be an operator able to change a set point to a value outside certain conditions determined by the engineering staff. The system did not enforce the proper authentication and use control restrictions to disallow the change by the operator. Another example would be a password being sent in clear text, allowing a network engineer to view the password while troubleshooting the system. The system did not enforce proper data confidentiality to protect the password.
SL 2: Protection against intentional violation using simple means with low resources, generic skills and low motivation Simple means do not require much knowledge on the part of the attacker. The attacker does not need detailed knowledge of security, the domain or the particular system under attack. These attack vectors are well known and there may be automated tools for aiding the attacker. They are also designed to attack a wide range of systems instead of targeting a specific system, so an attacker does not need a significant level of motivation or resources at hand. An example would be a virus that infects a workstation spreading to another engineering workstation since they both use the same general purpose operating system. Another example would be an attacker compromising a web server in the enterprise network by an exploit downloaded from the Internet for a publicly known vulnerability in the general purpose operating system of the web server. The attacker uses the web server as a pivot point in an attack against other systems in the enterprise network as well as the industrial network. A third example would be an operator that views a website which downloads a Trojan that opens a hole in the routers and firewalls to the Internet.
SL 3: Protection against intentional violation using sophisticated means with moderate resources, IACS specific skills and moderate motivation Sophisticated means require advanced security knowledge, advanced domain knowledge, advanced knowledge of the target system or any combination of these. An attacker going after a SL 3 system will likely be using attack vectors that have been customized for the specific target system. The attacker may use exploits in operating systems that are not well known, weaknesses in industrial protocols, specific information about a particular target to violate the security of the system or other means that require a greater motivation as well as skill and knowledge set than are required for SL 1 or 2. An example of sophisticated means could be password or key cracking tools based on hash tables. These tools are available for download, but applying them takes knowledge of the system (such as the hash of a password to crack). Using Figure A.1, another example would be an attacker that gains access to the FS-PLC through the serial conduit after gaining access to the control PLC through a vulnerability in the Ethernet controller. Using Figure A.2, a third example would be an attacker that gains access to the data historian by using a brute-force attack through the industrial/enterprise DMZ firewall initiated from the enterprise wireless network.
SL 4: Protection against intentional violation using sophisticated means with extended resources, IACS specific skills and high motivation SL 3 and SL 4 are very similar in that they both involve sophisticated means used to violate the security requirements of the system. The difference comes from the attacker being even more motivated and having extended resources at their disposal. These may involve high-performance computing resources, large numbers of computers or extended periods of time. An example of sophisticated means with extended resources would be using super computers or computer clusters to conduct brute-force password cracking using large hash tables. Another example would be a botnet used to attack a system using multiple attack vectors at once. A third example would be an organized crime organization that has the motivation and resources to spend weeks attempting to analyze a system and develop custom “zero-day” exploits.
Printed copies for reference only
Stryker Confidential - This document contains information that is confidential and proprietary. Neither this document nor the information herein may be reproduced, used, or disclosed to or for the benefit of any third party without the prior written consent of Stryker.